## ORDER INFORMATION (hidden)
Bangladesh
This order is for how many months : | 9 |  |  |  |  | Belarus
Cameroun
DATA FOR THIS ORDER | [5 empty] | Cote d'Ivoire
Number of samples for processing: | [5 empty] | Djibouti
Numbre of LPA: | [5 empty] | Ethiopia
Number of TB Culture on solid media: | [5 empty] | Georgia
Number of TB Culture on liquid media: | [5 empty] | Haiti
Number of DST on liquid media (SIRE) | [5 empty] | India
Indonesia
Kazakhstan
Kenya
Kyrgyzstan
Budget Line | Sum of Total Cost |  |  |  |  | Lesotho
BD | 0 |  |  |  |  | Mozambique
Generic | 0 |  |  |  |  | Myanmar
HAIN | 0 |  |  |  |  | Peru
SD | 0 |  |  |  |  | Republic of Moldova
TTM | 0 |  |  |  |  | Rwanda
Grand Total | 0
Senegal
Swaziland
Tajikistan
Sum of Total Volume | Column Labels |  |  |  |  | Uganda
Supplier | Keep cool 2-8°C | Dangerous goods | Keep cool - 20°C | (blank) | Grand Total | U.R Tanzania
BD | 0 |  |  | 0 | 0 | Uzbekistan
HAIN | 0 |  |  | 0 | 0 | Vietnam
SD |  |  |  | 0 | 0
TTM | 0 | 0 | 0 | 0 | 0
Grand Total | 0 | 0 | 0 | 0 | 0

## GDF TB Diagnostics Catalogue
|  | TB Diagnostics Catalogue ­ Product specifications | [5 empty] | Ordering |  | Product packaging | [5 empty] | Product price Concessional prices are valid only for some countries https://www.finddx.org/find-negotiated-product-pricing/ |  |  | Product information
GDF code | Product Name | Supplier | Unit | Supplier technical specifications | Shelf Life | Supplier Lead Time EXW | Type of Transport | GDF recommendations | Quantity to order | Weight Kg | Depth cm | Width cm | Height cm | Volume m3 | Total Volume | Indicative FCA Price | Currency | Total Cost | Manufacturer product code | QA of product | Dangerous goods | Custom Tariff #
1. GDF microscopy kits
106522 | Diagnostic consumables kit - ZN light microscopy | Svizera | 1 | Ziehl Neelsen stain solution, 1 L (Strong Carbol Fuchsin 1% solution for ZN staining): 6 , Ziehl Neelsen decolourisation solution, 1 L (Ready-made decolourisation solution of ZN stained slides – 3,0% HCl/alcohol): 6, Ziehl Neelsen counterstaining solution, 1 L (Methylene blue 0,1%, counterstaining solution for ZN stain): 6 , 50 x 20 Microscope slides: 1, 100x Filter papers: 1, 10x Pencils: 1, 2x Waterproof marker pens: 1, 500x Plastic bags W 200 mm x L 300 mm: 1, 1000x Wooden applicator sticks: 2, 400x Nitrile pairs of gloves (100 small, 200 medium, 100 large): 1, 30x150 Carton of paper towels: 1, lens cleaning tissue -100 tissues: 1, immersion oil 100 ml: 1, denatured alcohol 96% 0.5 L: 5 and stable chlorine disinfectant 50 tabs: 5. | 35 months | 1 - 3 weeks | Dangerous goods | · Shelf lives of products in the kit vary, but no product has a shelf life shorter than 35 months. · Tip: 72 kits (11 pallets) will fill up a 20 foot container · For more information on transport container specifications, visit: http://cargow.com/equipment |  | 27.67 | 60 | 40 | 40 | 0.10152 | 0 | 244 | USD | 0 | Various (Kit assembled by Svizera) |  | Dangerous goods Carbol fuchsin : UN 2810 III 8 Ethanolic solutions : UN 1993 II 3 (for alcohol lamp) Ethanol solution UN 1170 II 3 (Decolorizer) Chlorine tablets : UN3077 III 9 | DGR: 38220000 Non DGR: 90189084
106523 | Diagnostic consumables kit - LED/Auramine microscopy | Svizera | 1 | Auramine stain solution, reagents (Set of 3 components: 200 g - Colourless Phenol crystals C6H6O, 1 litter - Ethanol for stain solutions C2H6O and 5 g - Auramine O C17H22ClN3 to produce Auramine stain solution): 1 set, Auramine decolourisation solution, 1 L (Ready-made decolourisation solution 0,5% HCl/alcohol): 6, Auramine counterstaining solution, 1 L (Methylene blue 0,3%, counterstaining solution for auramine O stain) : 6, 50 x 20 Microscope slides: 1, 100x Filter papers: 1, 10x Pencils: 1, 2x Waterproof marker pens: 1, 500x Plastic bags W 200 mm x L 300 mm: 1, 1000x Wooden applicator sticks: 1, 400x Nitrile pairs of gloves (100 small, 200 medium, 100 large): 1, 30x150 Carton of paper towels: 1, lens cleaning tissue -100 tissues: 1, immersion oil 100 ml: 1, denatured alcohol 96% 0.5 L: 1 and stable chlorine disinfectant 50 tabs: 5. | 35 months | 1 - 3 weeks | Dangerous goods | · Shelf lives of products in the kit vary, but no product has a shelf life shorter than 35 months; · Tip: 72 kits (11 pallets) will fill up a 20 foot container; · For more information on transport container specifications, visit: http://cargow.com/equipment |  | 37.43 | 60 | 40 | 40 | 0.1148 | 0 | 182.5 | USD | 0 | Various (Kit assembled by Svizera) |  | Dangerous goods Phenol, solid UN 1671 II 6.1 Ethanolic solutions) UN 1993 II 3 (for alcohol lamp) Ethanol solution UN 1170 II 3 (Decolorizer) Chlorine tablets : UN3077 III 9 | DGR: 38220000 Non DGR: 90189084
106526 | Equipment and starter kit for LED or ZN light microscopy | Svizera | 1 | Glass alcohol lamp with 1 meter pre-cut wicks: 2, timer: 1, staining rack: 2, slide drying racks: 2, slide storage boxes: 6, slide-holding forceps: 2, reagent bottle: 4, staining bottle 250 ml: 3, staining bottle 500 ml: 3, beaker 250ml: 1, beaker 1000ml: 1, measuring cylinder 10 ml: 1, goggles: 2, 4x funnels (1 large 155 mm, 1 medium 120 mm and 2 small 45 mm) with 100 filter papers: 1, set of pencil sharpener with 10 pencils: 1, 3x stands for waste bags with 500 plastic bags: 1, paper towel dispenser with 1 carton of paper: 1, GLI Handbook on Lab Diagnosis of TB by Sputum Microscopy: 5, SOPs: 1 and a WHO Policy statement. | N/A | 1 week | General Cargo | · Tip: 104 kits (11 pallets) will fill up a 20 foot container; · For more information on transport container specifications, visit: http://cargow.com/equipment · GLI Handbook on Lab Diagnosis of TB by Sputum Microscopy also available here: http://www.stoptb.org/wg/gli/assets/documents/TB%20MICROSCOPY%20HANDBOOK_FINAL.pdf |  | 14.5 | 60 | 40 | 40 | 0.084 | 0 | 117.5 | USD | 0 | Various (Kit assembled by Svizera) |  |  | 90189084
2. LED / Light microscopy: Instruments, reagents and consumables
106535 | Binocular bright field microscope - ZN | MEG | 1 | Binocular microscope Olympus CX23LEDRFS1-set with: - LED illumination; Binocular tube - Stand with quadruple revolving nose-piece - Eyepieces 10x with Helicoid adjustment (FN 20, anti fungus); right-handle- double plate stage - Abbe condenser; power cord - Plan objectives (anti-fungus) 4x, 10x, 40x, 100x - With AC adapter and fixing belt for transportation - Power supply: 100-240VAC, 50/60Hz. | N/A | 11 weeks | General Cargo | · Product linked to 106629; 106630; · 1 year warranty; · Olympus CX23 features: https://www.olympus-lifescience.com/en/microscopes/upright/cx23/ |  | 6 | 40 | 50 | 50 | 0.02 | 0 | 837.36 | USD | 0 | Olympus CX23 | CE
106629 | Binocular bright field microscope - reflecting mirror | MEG | 1 | Mirror unit for microscope Olympus CH20/CX21/CX22 and CX23 | N/A | 11 weeks | General Cargo | · Product linked to 106535 |  | 0.1 | 75 | 67 | 75 | 0.002 | 0 | 12.21 | USD | 0 | Olympus | CE
106630 | Binocular bright field microscope - transport case | MEG | 1 | Wooden storage/transport box for Olympus microscopes models CH20/CX21/CX22 and CX/23, with carrying handle, size 26.4x30.5x47.4cm | N/A | 11 weeks | General Cargo | · Product linked to 106535; · If the microscope will stay in the lab, a transport case may not be necessary and may be excluded from an order |  | 6.25 | 26.4 | 30.5 | 47.4 | 0.0383 | 0 | 86.13 | USD | 0 | Olympus | N/A
106531 | LED microscope (basic configuration) | Svizera | 1 | ZEISS fluorescence microscope “Primo Star iLED” (fixed Koehler version) with binocular tube 30°, equipped with quadruple revolving nosepiece (tilted backwards) and objectives “Planachromatic” 10x, 20x, 40x and 100x (oil), with white LED illumination for transmitted-bright field and blue LED illumination for epi-fluorescence. | N/A | 3 - 4 weeks | General Cargo | · Product linked to 106532; 106533; 106534; · 2 years warranty |  | 15 | 30 | 61 | 60 | 0.14 | 0 | 1339.45 | USD | 0 | Primo Star iLED | CE
106532 | LED microscope - light mirror | Svizera | 1 | Attachable transmitted-light mirror for Primo Star fixed Köhler | N/A | 3 - 4 weeks | General Cargo | · Product linked to 106531 |  | 0.3 |  |  |  | 0.002 | 0 | 24.71 | USD | 0 | Primo Star iLED | CE
106533 | LED microscope - battery supply unit | Svizera | 1 | Battery supply unit for Primo Star Fluorescence iLED or LED consisting of battery and charger in one unit, for power supply via the line or a battery. With country-specific adapters, for reflected-light FL illumination iLED or/and LED transmitted-light illumination (not suitable for HAL), includes 5x NiCd battery, 5000mAh. | N/A | 1 week | General Cargo | Product can be used for a light microscope (106535) or for a LED microscope (106531) |  | 5.6 | 27 | 22 | 14 | 0.02 | 0 | 94.11 | USD | 0 | Primo Star iLED | CE
106534 | LED microscope - transport case | Svizera | 1 | Case for transport and storage for Primo Star FL iLED (D); also suitable for Primo Star without iLED. With trolley, 2 wheels, lateral handle and lockable; dustproof; outside dimensions of the case: 615x540x325 mm, weight 7.5 kg | N/A | 3 - 4 weeks | General Cargo | · Product linked to 106531; · If the microscope will stay in the lab, a transport case may not be necessary and may be excluded from an order |  | 10 | 43 | 28 | 50 | 0.29 | 0 | 243.35 | USD | 0 | Primo Star iLED | N/A
106525 | Sputum containers | Svizera | 1000 | Break/leak-resistant (waterproof) container, made of transparent plastic pure polyethylene or polypropylene, completely combustible without generating toxic compounds; wide mouth to easily collect sputum, opening angle of 270°; IEC 60529 certified IP67; a multi-thread closing screw cap of double thread; with identification through a frosted writing panel that accepts pencil; dimensions: Volume in the range 80 ml; height in the range 66 mm; mouth diameter in the range 52 mm. | N/A | 1 - 3 weeks | General Cargo | · Tip: 176 boxes (11 pallets) will fill up a 20' container; · For more information on container specifications, visit: http://cargow.com/equipment |  | 14.63 | 60 | 50 | 40 | 0.11765 | 0 | 83.3 | USD | 0 | Suzhou Hengxiang Co. Ltd. |  |  | 90189084
106610 | AFB quality control slides | Svizera | 80 | Quality control slides to validate routine staining procedures Individually packed in sealed foil pouches to preserve control sample integrity, 3 slides for moderate organism count (3+) per high-powered field, 3 slides for low organism count (2+) per high-powered field, each slide includes positive test sample and negative control per high-powered field, stain control log sheets shall be included to comply with regulatory requirements. | 18 months | 3 weeks | General Cargo | · Product has slide with low organism count |  | 0.072 | 9 | 13 | 11.5 | 0.0006 | 0 | 14.5 | USD | 0 | Alphatec Systems | CE |  | 382200
106564 | Spirit lamp | Svizera | 1 | Glass alcohol lamp with screw top lid, complete cotton wick, 200ml | N/A | 1 - 3 weeks | General Cargo |  |  | 0.23 | 1.45 | 48 | 33 | 0.0015 | 0 | 1.88 | USD | 0 | Atlas Surgical |  |  | 90189084
106565 | Roll of cotton wick for spirit lamp | Svizera | 1 | Spare wick for spirit lamp, approx. 1 meter. resalable packaging | N/A | 1 - 3 weeks | General Cargo |  |  | 0.02 | 1.24 | 22 | 8 | 6.4e-05 | 0 | 0.14 | USD | 0 | Atlas Surgical |  |  | 90189084
106566 | Staining rack | Svizera | 1 | Constructed in the laboratory from glass rods and rubber or silicon tubes, or from wire (preferably stainless steel wire) of appropriate size. Racks to have a bowl (made of glass or stainless steel) to collect staining solutions. | N/A | 1 - 3 weeks | General Cargo |  |  | 0.77 | 48 | 10 | 2.5 | 0.0095 | 0 | 4.38 | USD | 0 | Yongfeng outdoors product Chao'an Caitang NOBO |  |  | 90189084
106567 | Slide drying rack | Svizera | 1 | Hold standard slides (3 x 1" or 25 x 75mm) during drying and cleaning. Constructed of strong, chemical-resistant Stainless steal Capacity: 40 slides; measures rack 11.75 x 4.13 x 1". | N/A | 1 - 3 weeks | General Cargo |  |  | 0.14 | 3.1 | 22 | 8 | 0.00048 | 0 | 1.7 | USD | 0 | Chao'an Caitang NOBO Hardware Factory |  |  | 90189084
106568 | Slide storage box | Svizera | 1 | Boxes able to hold 100 slides (slides 76 mm ~ 26 mm), with a sturdy lock, a cork-lined base and a flap cover with an index-card holder, made of plastic. | N/A | 1 - 3 weeks | General Cargo |  |  | 0.23 | 0.72 | 42 | 42 | 0.001184 | 0 | 1.94 | USD | 0 | Suzhou Hengxiang |  |  | 90189084
106569 | SIide-holding forceps | Svizera | 1 | Slide holding forceps - 11,5cm in length, coating made of PFTE, stainless steel, packed in a plastic pouch, CE marked or equivalent | N/A | 1 - 3 weeks | General Cargo |  |  | 0.02 | 0.3 | 11 | 6 | 2.5e-06 | 0 | 0.9 | USD | 0 | Atlas Surgical |  |  | 90189084
106570 | Timer | Svizera | 1 | Robust mechanical interval timer, for periods up to 60 min, the end of the period signalled by loud, long bell ring. With indicator cum winding mechanism, CE marked or equivalent | N/A | 1 - 3 weeks | General Cargo |  |  | 0.08 | 0.37 | 28 | 32 | 0.000224 | 0 | 1.11 | USD | 0 | Suzhou Hengxiang |  |  | 90189084
106571 | Reagent bottle | Svizera | 1 | Bottles 1 L, made of amber borosilicate glass, with a GL45 thread and cap, compliant with ISO 4746 or equivalent (Oxygen-free copper – scale adhesion test). | N/A | 1 - 3 weeks | General Cargo |  |  | 0.57 | 2.31 | 40 | 40 | 0.0021 | 0 | 4.72 | USD | 0 | Nantong Hailun Bio-medical Apparatuses Manufacturing |  |  | 90189084
106572 | Staining bottle - 250ml | Svizera | 1 | Polyethylene flasks with a narrow neck with a screw cap, a swan neck or a jet dispenser; volume 250 ml | N/A | 1 - 3 weeks | General Cargo |  |  | 0.03 | 0.29 | 44 | 44 | 0.00045 | 0 | 0.46 | USD | 0 | Suzhou Hengxiang |  |  | 90189084
106573 | Staining bottle - 500 ml | Svizera | 1 | Polyethylene flasks with a narrow neck with a screw cap, a swan neck or a jet dispenser; volume 500 ml | N/A | 1 - 3 weeks | General Cargo |  |  | 0.05 | 0.56 | 43 | 43 | 0.00074 | 0 | 0.53 | USD | 0 | Suzhou Hengxiang |  |  | 90189084
106574 | Beaker - 250 ml | Svizera | 1 | 250ml, borosilicate, graduated, heavy duty, for general laboratory use. The measures (50, 100, 150, 200, 250) printed in white on the glass | N/A | 1 - 3 weeks | General Cargo |  |  | 0.22 | 0.61 | 40 | 40 | 0.00056 | 0 | 1.08 | USD | 0 | Ningbo Joan Lab Equipment |  |  | 90189084
106575 | Beaker - 1000 ml | Svizera | 1 | 1000ml, borosilicate, graduated, heavy duty, for general laboratory use. The measures (50, 100, 150, 200, 250) printed in white on the glass | N/A | 1 - 3 weeks | General Cargo |  |  | 0.6 | 2 | 40 | 40 | 0.0019 | 0 | 3.47 | USD | 0 | Ningbo Joan Lab Equipment |  |  | 90189084
106576 | Measuring cylinder | Svizera | 1 | 10 ml, acid resistant. Incremental measures (1 - 10) imprinted. | N/A | 1 - 3 weeks | General Cargo |  |  | 0.06 | 1.3 | 18 | 11 | 0.00024 | 0 | 1.11 | USD | 0 | Ningbo Joan Lab Equipment |  |  | 90189084
106609 | Microscope slides | Svizera | 1000 | Made of soda-lime glass; 76 mm × 24 mm; 1.0–1.2 mm thick; compliant with ISO 8037–1 (Optics and optical instruments --Microscopes --Slides --Part 1: Dimensions, optical properties and marking); cleaned and degreased; straight edges and corners; cellophane wrapped and tropical packing; frosted ends for marking. | N/A | 1 - 3 weeks | General Cargo |  |  | 5.6 |  |  |  | 0.01 | 0 | 10.4 | USD | 0 | Hebei Medicines & Health Products Imp. & Exp. Group Corp |  |  | 90189084
106627 | Immersion oil, 100 ml | Svizera | 1 | Immersion oil, 100 ml | 47 months | 1 - 3 weeks | General Cargo |  |  | 0.12 |  |  |  | 0.000306 | 0 | 9.24 | USD | 0 | GCC Diagnostics |  |  | 38220000
106631 | Set of 4 funnels with 100 filters | Svizera | 1 | Set (1 large, 1 medium and 2 small) of 4 funnels with 100 filter papers | N/A | 1 - 3 weeks | General Cargo |  |  | 0.39 |  |  |  | 0.006 | 0 | 6.6 | USD | 0 | Suzhou Hengxiang Hangzhou Ocome Technology |  |  | 90189084
106597 | Denatured Alcohol | TTM | 1L | Ethanol, approx. 99 %, 1 l | 12 months | 3 - 4 weeks | Dangerous goods |  |  | 0.92 | 12 | 12 | 22 | 0.003168 | 0 | 12.89 | USD | 0 | GCC Diagnostics |  | UN 1170 Class 3 Packing group II | 38220000
106598 | Ethanol for stain solutions and decolourization | TTM | 1L | Ethanol, approx. 99 %, 1 l | 12 months | 3 - 4 weeks | Dangerous goods |  |  | 0.92 | 12 | 12 | 22 | 0.003168 | 0 | 12.89 | USD | 0 | VWR Chemicals 84835.290 |  | UN 1170 Class 3 Packing group II
106599 | Colourless phenol crystals | TTM | 500g | Phenol, 2 x 250 g | 12 months | 3 - 4 weeks | Dangerous goods |  |  | 0.8 | 20 | 20 | 14 | 0.0056 | 0 | 96.99 | USD | 0 | Merck Millipore 2 x 1.00206.0250 |  | UN 1671 Class 6.1 Packing group II | 29071100
106276 | Auramine O | TTM | 25g | Auramine O, 25 g | 12 months | 3 - 4 weeks | Dangerous goods |  |  | 0.045 | 4 | 4 | 9 | 0.000144 | 0 | 48.71 | USD | 0 | Sigma Aldrich 861030-25g |  | UN 2811 Class 6.1 Packing group III | 32041300
106601 | Basic Fuchsin | TTM | 100g | Basic Fuchsin, 4 x 25 G | 12 months | 3 - 4 weeks | General Cargo |  |  | 0.4 | 6 | 6 | 13 | 0.000468 | 0 | 331.04 | USD | 0 | Sigma Aldrich 4 x 215597-25g |  |  | 38220000
106602 | Methylene blue | TTM | 100g | Methylene blue, 10 x 10 g | 12 months | 3 - 4 weeks | General Cargo |  |  | 0.4 | 6 | 6 | 13 | 0.000468 | 0 | 3781.1 | USD | 0 | Sigma Aldrich 10 x 556416-10g |  |  | 32041300
106603 | Hydrochloric acid | TTM | 1L | Hydrochloric acid, 37 %, ACS reagent, 2 x 500 g | 12 months | 3 - 4 weeks | Dangerous goods |  |  | 1.15 | 11 | 11 | 22 | 0.002662 | 0 | 154.06 | USD | 0 | Sigma Aldrich 2 x 320331-500ml |  | UN 1789 class 8 Packing group II
106604 | Sulphuric Acid | TTM | 1L | Sulphuric Acid, 1000 ml | 12 months | 3 - 4 weeks | Dangerous goods |  |  | 1.6 | 11 | 11 | 22 | 0.002662 | 0 | 368.23 | USD | 0 | Sigma Aldrich 84716-1l |  | UN 1830 Class 8 Packing group II | 28070000
106605 | Suction piston graduated pipette 50 ml | TTM | 1 | Graduated pipette with piston, capacity: 50 ml, graduated: 0,2 ml | N/A | 3 - 4 weeks | General Cargo |  |  | 0.02 | 3 | 3 | 50 | 0.00045 | 0 | 44.24 | USD | 0 | Hirschmann 612-1010
106606 | Weighing boats | TTM | 1 pack of 10 | Weighing boats, rectangular, glazed porcelain, 55 x 25 x 10 mm, pack of 12 pieces | N/A | 3 - 4 weeks | General Cargo |  |  | 1.3 | 6 | 3 | 15 | 0.00027 | 0 | 67.03 | USD | 0 | Avignon 459-0022
106607 | Reagent bottle, 2 litters | TTM | 1 | Laboratory bottle, round, amber, borosilicate glass, with screw cap, 2000 ml | N/A | 3 - 4 weeks | General Cargo |  |  | 1.05 | 18 | 18 | 28 | 0.009072 | 0 | 20.02 | USD | 0 | VWR 215-2329
3. Molecular testing: GeneXpert system
GXIV-4-D | GeneXpert GX-IV 4-Module System with Desktop | Cepheid | 1 | GeneXpert 4-module instrument with desktop computer | N/A | 2 weeks | General Cargo | · Most commonly used configuration globally; · GeneXpert IV features: http://www.cepheid.com/us/cepheid-solutions/systems/genexpert-systems/genexpert-iv |  | 32 | 70 | 51 | 70 | 0.39 | 0 | 17000 | USD | 0 | GXIV-4-D | CE
GXIV-4-L | GeneXpert GX-IV 4-Module System with Laptop | Cepheid | 1 | GeneXpert 4-module instrument with laptop computer | N/A | 2 weeks | General Cargo | · If the GeneXpert will stay in a laboratory and will not be travelling, the GXIV-4-D desktop may be a better option |  | 24 | 70 | 65 | 55 | 0.37 | 0 | 17500 | USD | 0 | GXIV-4-L | CE
GXIV-2-D | GeneXpert GX-IV 2-Module System with Desktop | Cepheid | 1 | GeneXpert 4-module instrument shell with only 2 modules, with desktop computer | N/A | 2 weeks | General Cargo | · This system may be a good option for a site with low expected utilization, but with the possibility to buy additional modules in the future when utilization increases |  | 29 | 70 | 51 | 70 | 0.39 | 0 | 11780 | USD | 0 | GXIV-2-D | CE
GXIV-2-L | GeneXpert GX-IV 2-Module System with Laptop | Cepheid | 1 | GeneXpert 4-module instrument shell with only 2 modules, with laptop computer | N/A | 2 weeks | General Cargo | · If the GeneXpert will stay in a laboratory and will not be travelling, the GXIV-2-D desktop may be a better option; · This system may be a good option for a site with low expected utilization, but with the possibility to buy additional modules in the future when utilization increases |  | 21 | 70 | 65 | 55 | 0.37 | 0 | 12280 | USD | 0 | GXIV-2-L | CE
GXXVI-16-D | GeneXpert GX-XVI 16-Module System with Desktop | Cepheid | 1 | GeneXpert 16-module instrument with desktop computer | N/A | 2 weeks | General Cargo |  |  | 112 | 90 | 75 | 150 | 0.9 | 0 | 71000 | USD | 0 | GXXVI-16-D | CE
GXXVI-16-L | GeneXpert GX-XVI 16-Module System with Laptop | Cepheid | 1 | GeneXpert 16-module instrument with laptop computer | N/A | 2 weeks | General Cargo |  |  | 102 | 90 | 75 | 120 | 0.9 | 0 | 71500 | USD | 0 | GXXVI-16-L | CE
700-4510 | Dust Filter for GeneXpert System | Cepheid | 1 | KIT,GX4 R2 SVC & AIR FILTRATION FOR HBDC (Dust Filter) | N/A | 2 weeks | General Cargo | · Need additional filter for dusty environments; · Product linked to GXIV-2-D; GXIV-2-L; GXIV-4-D; GXIV-4-L; GXXVI-16-D; GXXVI-16-L |  | 0.1 |  |  |  | 0.001 | 0 | 350 | USD | 0 | 700-4510 | CE
CGXMTB-RIF-50 | Xpert MTB/RIF kit of 50 tests | Cepheid | 50 | Xpert MTB/RIF kit of 50 tests | 24 months | 2 weeks | General Cargo |  |  | 3.6 | 31 | 27 | 19 | 0.0159 | 0 | 499 | USD | 0 | CGXMTB-RIF-50 | CE | General Cargo UN 2920 – CLASS 8(3)) in excepted quantity (EQ)
GXMTB/RIF-ULTRA-50 | Xpert MTB/RIF ULTRA kit of 50 tests | Cepheid | 50 | Xpert MTB/RIF ULTRA kit of 50 tests | 9 - 10 months | 2 weeks | General Cargo |  |  | 3.6 | 31 | 27 | 19 | 0.0159 | 0 |  |  | 0 | GXMTB/RIF-ULTRA-50 | CE | General Cargo UN 2920 – CLASS 8(3)) in excepted quantity (EQ)
GX4-4 VAL | Validation Pack for new GeneXpert GX-IV 4-Module System | Cepheid | 1 | Validation Pack for new GeneXpert GX-IV 4-module systems | N/A | 2 weeks | General Cargo | · Validation packs (also known as verification kits) are free and should be sent with all new GeneXperts, to verify initial performance after installation |  | 0.1 |  |  |  | 0.0001 | 0 | 0 | USD | 0 | GX4-4 VAL | CE
XPERTCHECKCE- 5 | XpertCheck Calibration Pack for GeneXpert GX-IV 4-Module System | Cepheid | 1 | GeneXpert XpertCheck Calibration Packs for GeneXpert System with 4 Modules | N/A | 2 weeks | General Cargo | · An XpertCheck pack should be ordered when ordering a new GeneXpert system without upfront 3 years of warranty extension (see WX04RG36), for use after year 1 during the initial free 2-year warranty; · Also, an XpertCheck pack is provided for free when a 1-year warranty extension (see WX04RG12) is ordered for a GeneXpert. |  | 0.4 | 23 | 14 | 17 | 0.005474 | 0 | 450 | USD | 0 | XPERTCHECKCE- 5 | CE
GXIV-Module | Additional/Replacement GeneXpert Module | Cepheid | 1 | GeneXpert additional module (new) | N/A | 2 weeks | General Cargo |  |  | 2.5 | 36 | 36 | 20 | 0.02592 | 0 | 3360 | USD | 0 | GXIV-Module | CE
WX04RG12 | 1-year warranty extension- GX-IV (4-Module System) | Cepheid | 1 | 1-year warranty extension- GX-IV (4-Module System) | N/A | On demand | Service |  |  | N/A | N/A | N/A | N/A | N/A | N/A | 2898 | USD | 0 | WX04RG12 | N/A |  | N/A
WX04RG36 | 3-years warranty extension- GX-IV (4-Module System) | Cepheid | 1 | 3-years warranty extension- GX-IV (4-Module System), paid before the end of the two first year of warranty | N/A | On demand | Service |  |  | N/A | N/A | N/A | N/A | N/A | N/A | 7902 | USD | 0 | WX04RG36 | N/A |  | N/A
WX04UP36 | 3-years warranty extension- GX-IV (4-Module System), ordered with System | Cepheid | 1 | 3-year warranty extension for GeneXpert GX-IV 4-module system, paid upfront together with the system purchase | N/A | On demand | Service | · No need to order GeneXpert Calibration Pack for GeneXpert System as the 5 year total warranty includes 5 sets of calibration packs |  | N/A | N/A | N/A | N/A | N/A | N/A | 6840 | USD | 0 | WX04UP36 | N/A |  | N/A
WX02RG12 | 1-year warranty extension- GX-IV (2-Module System) | Cepheid | 1 | 1-year warranty extension- GX-IV (2-Module System) | N/A | On demand | Service |  |  | N/A | N/A | N/A | N/A | N/A | N/A | 1896 | USD | 0 | WX02RG12 | N/A |  | N/A
WX02RG36 | 3-years warranty extension- GX-IV (2-Module System) | Cepheid | 1 | 3-years warranty extension- GX-IV (2-Module System), paid before the end of the two first years of warranty | N/A | On demand | Service |  |  | N/A | N/A | N/A | N/A | N/A | N/A | 5184 | USD | 0 | WX02RG36 | N/A |  | N/A
WX02UP36 | 3-years warranty extension- GX-IV (2-Module System), ordered with System | Cepheid | 1 | 3-year warranty extension for GeneXpert GX-IV 2-module system, paid upfront together with the system purchase | N/A | On demand | Service | · No need to order GeneXpert Calibration Pack for GeneXpert System as the 5 year total warranty includes 5 sets of calibration packs |  | N/A | N/A | N/A | N/A | N/A | N/A | 4500 | USD | 0 | WX02UP36 | N/A |  | N/A
WV16RG12 | 1-year warranty extension- GX-XVI (16-Module System) | Cepheid | 1 | 1-year warranty extension- GX-XVI (16-Module System) | N/A | On demand | Service |  |  | N/A | N/A | N/A | N/A | N/A | N/A | 7800 | USD | 0 | WV16RG12 | N/A |  | N/A
WX16RG36 | 3-years warranty extension- GX-XVI (16-Module System) | Cepheid | 1 | 3-years warranty extension- GX-XVI (16-Module System), paid before the end of the two first year of warranty | N/A | On demand | Service |  |  | N/A | N/A | N/A | N/A | N/A | N/A | 20898 | USD | 0 | WX16RG36 | N/A |  | N/A
WX16UP36 | 3-years warranty extension- GX-XVI (16-Module System), ordered with System | Cepheid | 1 | 3-year warranty extension for GeneXpert GX-XVI 16-module system, paid upfront together with the system purchase | N/A | On demand | Service | · No need to order GeneXpert Calibration Pack for GeneXpert System as the 5 year total warranty includes 5 sets of calibration packs |  | N/A | N/A | N/A | N/A | N/A | N/A | 18504 | USD | 0 | WX16UP36 | N/A |  | N/A
106493 | Solar panel and battery system for GeneXpert System | TTM | 1 | Solar panel system 24 V for min 8 hour operation time with GeneXpert 4-module system. Consisting of: Solar panel type STP085s-12/Bb, specs according to EN 60904-3 (STC), DIN EN 50380, STECA PR charge controller, model PR 3030 with integrated display built in Ah meter, automatic load reconnection, SOC regulation, automatic voltage selection (12V, 24V), data logger, simple filed adjustable parameter by 2 button, ext. temp. sensor, day and night alarm CE certified, DIN EN ISO 14001, invertor type Phonix 24/ 750 with remote control panel DIP switch 50/ 60 Hz and power safe mode, 8 batteries 6v 330Ah, installation kit for 36 m, with 1x multi plug for 4 connections (example: equipment, computer, monitor, printer), set of 10 spare fuses | N/A | 3 weeks | General Cargo | · Solar panel system 24 V for minimum 8 hour operation time using GeneXpert 4-module system · 2 years warranty · Voltage / Amp / Frequency: 12/24V invertor |  | 519 | 120 | 80 | 160 | 1.536 | 0 | 9714.06 | USD | 0 | TTM-Elektrowerkstatt KALKU, 24 V Solar 8 h | CE/ DIN |  | 85023980
4. Molecular testing: Loop-mediated isothermal amplification (TB-LAMP)
106632 | HumaLoop T | Human | 1 | Preprogrammed incubator with built-in heating block, reaction block, timer and fluorescence detection unit for lysis, amplification, enzyme inactivation and visual result interpretation. | N/A | 11 weeks | General Cargo | · The operator can run a maximum of 16 reactions in one run of HumaLoop T. Of these 16 tests, 14 are patient samples and 2 are controls. Each cycle takes 1.5 hours, therefore 5 cycles may be run per day in a well-functioning laboratory (up to 70 patient samples per day); · A UPS (106443) should be ordered with this product; · 1 year warranty; · WHO Guidance for LAMP: http://www.who.int/tb/publications/lamp-diagnosis-molecular/en/ |  | 10.5 | 37.5 | 39 | 35 | 0.0512 | 0 | 2450 | EURO | 0 | 961000 | CE |  | 90275000
106443 | UPS - 700 VA - 500 W | TTM | 1 | On-Line UPS Vanguard VGD-700: UPS capacity 700 VA 230V sine wave output, 50/60 Hz outlets 4x Iec 320, surge protection RJ11/RJ45, dimensions W152xd420xH237 | N/A | 3 weeks | General Cargo special provision A67 | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 15.3 | 55 | 28 | 34 | 0.05236 | 0 | 467.03 | USD | 0 | PCM VGD-700 | CE | General Cargo Valve Regulated Lead Acid Battery UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 90275000
106634 | Loopamp MTBC Detection Kit | Human | 96 tests (2 sets of 48 tests) | Qualitative in vitro diagnostic test for the detection of MTB complex. The amplification under isothermal conditions enables within a short time the fluorescence or turbidity based detection of MTB complex. | 9 months | 11 weeks | General Cargo | · A Detection kit of 96 tests and an Extraction kit of 90 tests will provide the appropriate number of reagents for 6 cycles comprising 84 patient samples plus 6 negative controls and 6 positive controls; · Need to order separately pipette of 60 ul (106055) and pipette tips (106388) |  | 0.19 | 19.5 | 17.5 | 7.5 | 0.0026 | 0 | 352.5 | EURO | 0 | 972000 | CE |  | 90275000
106635 | Loopamp PURE DNA Extraction Kit | Human | 90 tests | DNA extraction and isolation method for sputum samples through lysis/heating, removal of inhibitors and elution of DNA directly into the amplification tube. | 21 months | 11 weeks | General Cargo | · A Detection kit of 96 tests and an Extraction kit of 90 tests will provide the appropriate number of reagents for 6 cycles comprising 84 patient samples plus 6 negative controls and 6 positive controls |  | 1.9 | 23 | 21.5 | 32 | 0.0158 | 0 | 298.2 | EURO | 0 | 970000 | CE |  | 38220000
106633 | HumaTurb C+ HumaTurb A | Human | 1 C & 1 A | Programmable system with preinstalled settings, amplification and control units for isothermal amplification, enzyme inactivation and real-time turbidity detection. | N/A | 11 weeks | General Cargo | · High throughput system, for reference laboratories (up to 96 tests per run), as an alternative to HumaLoop T |  | 5,4 | 52 | 35 | 23.5 | 0.0445 | 0 | 15900 | EURO | 0 | 963200 | CE |  | 38220000
106636 | HumaHeat | Human | 1 | Heater designed for the sample lysis with Loopamp PURE DNA Extraction Kit before performing the LAMP reaction with HumaTurb system. | N/A | 11 weeks | General Cargo | · Accessory product for reference laboratories using HumaTurb C + HumaTurbA |  | 2.6 | 31 | 19.5 | 23.5 | 0.0142 | 0 | 960 | EURO | 0 | 964000 | CE |  | 84198998
106637 | Installation for TB-LAMP | Human | 1 | • On-site inspection according to site preparation check list, and installation of system(s) • Testing incubator of HumaHeat and turbidimetric functionality of HumaTurb C + A • Testing incubators and fluorescence module of HumaLoop according to Operation Qualification (OQ) • All equipment and test kits needed for OQ • Issue of Installation and Operation Qualification certificates, and start of basic warranty • Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. | N/A | N/A | Service | · Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. |  | N/A | N/A | N/A | N/A | N/A | N/A | 1295 | EURO | 0 | 210 | N/A |  | N/A
106638 | Training for TB-LAMP | Human | 1 | End user training for 2 days: • Introduction to Loopamp technology • Familiarization with products and procedures • Preparation of materials and workspace • Performing TB assay • Results interpretation & troubleshooting • Examination of trainees. Each trainee receives a certificate after passing the exam. • The training includes needed training documentation. • Test kits to be used for training purposes are not included in the pricing and should be provided by the trainee or can be ordered via the assigned partner of HUMAN • Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. | N/A | N/A | Service | · Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. |  | N/A | N/A | N/A | N/A | N/A | N/A | 250 | EURO | 0 | 210 | N/A |  | N/A
106639 | Maintenance for TB-LAMP | Human | 1 | • 1 year checking on-site conditions • Extensive cleaning of working area, instrument surface, heating blocks and fluorescence reading unit according to maintenance check list • Inspection and validation of UPS and power connection • Temperature check and verifying the incubation and reaction block are reaching the target temperature according to maintenance check list • Performance of maintenance test kit to assure the instrument is working properly according to Operation Qualification (OQ) • Needed equipment and test kit to perform the OQ • Issue of OQ certificate • Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. | N/A | N/A | Service | · Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. |  | N/A | N/A | N/A | N/A | N/A | N/A | 1235 | EURO | 0 | 150 | N/A |  | N/A
106640 | Maintenance with warranty extension for TB-LAMP | Human | 1 | • Warranty extension for 1 year (on-site) if the maintenance is done before the first warranty period ends and a copy of the filled maintenance check list is returned to HUMAN. • Extensive cleaning of working area, instrument surface, heating blocks and fluorescence reading unit according to maintenance check list • Inspection and validation of UPS and power connection • Temperature check and verifying the incubation and reaction block are reaching the target temperature according to maintenance check list • Performance of maintenance test kit to assure the instrument is working properly according to Operation Qualification (OQ) • Needed equipment and test kit to perform the OQ • Issue of OQ certificate • Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. | N/A | N/A | Service | · Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. |  | N/A | N/A | N/A | N/A | N/A | N/A | 1675 | EURO | 0 | 150 | N/A |  | N/A
5. Processing and transport: Equipment, reagents and consumables
106221 | Safety gas burner | TTM | 1 | Gasprofi 1 SCS, safety gas burner, with foot switch, IR sensor, with rest heat display and splash guard | N/A | 3 weeks | General Cargo | This item needs to be ordered with the tubing (106222) |  | 2 | 83 | 83 | 45 | 0.31 | 0 | 515.71 | USD | 0 | Wartewig Fuego SCS basic | CE |  | 84198998
106222 | Safety gas tubing | TTM | 1 | Safety gas tube 750mm long, with clamps inner diameter 9mm | N/A | 3 weeks | General Cargo | Product linked to 106221 |  | 0.1 | 23 | 23 | 45 | 0.002645 | 0 | 6.97 | USD | 0 | Deutsch & Neumann + Landefeld 370 0001+ SS16 | CE |  | 84198998
106220 | Overhead stirrer with universal adapter | TTM | 1 | Overhead stirrer and universal adaptor, with individual adjustment from 20-100 U/min, for 2 vessels, max weight 1 kg, universal adaptor for vessels from 50-160mm height | N/A | 3 weeks | General Cargo |  |  | 9.4 | 67 | 49 | 31 | 0.10177 | 0 | 1492.26 | USD | 0 | Heidolph Reax 2 (541-21001-00) |  |  | 84798200
106122 | Compact orbital shaker (for BSC) | TTM | 1 | Compact, universal small shaker suitable for shaking tasks with all small vessels and microtiter plates, adjustable speed of up to 3.000 rpm | N/A | 3 weeks | General Cargo |  |  | 3.6 | 33 | 22 | 14 | 0.010164 | 0 | 509.79 | USD | 0 | IKA MS 3 basic vortex mixer |  |  | 84798200
106231 | Test tube rack 18/8 stainless steel 35 mn | TTM | 1 | Test tube rack stainless steel, 3x7 places, 35mm | N/A | 3 weeks | General Cargo |  |  | 0.75 | 27 | 13 | 12 | 0.004212 | 0 | 85.64 | USD | 0 | Bochem 8254 | CE |  | 73269098
106240 | Alarm clock | TTM | 1 | Alarm clock, with stop function, digital, programmable for mx. 99 min, memory function, with alarm function at 0, | N/A | 3 weeks | General Cargo |  |  | 0.1 | 7 | 7 | 3 | 0.000147 | 0 | 4.12 | USD | 0 | VWR Collection 609-0128 | CE |  | 91069010
106242 | Transport box for specimen transport | TTM | 1 | Transport box for specimen, isolated, PE formed material, with 2 cooling elements | N/A | 3 weeks | General Cargo |  |  | 0.8 | 43 | 25 | 28 | 0.0301 | 0 | 30.93 | USD | 0 | Campinggaz 1 x 30060, 2 x 09377 | CE |  | 39269097
106243 | Cool box for specimen transport | TTM | 1 | Electric cooler Powerbox Plus 28 L, operates of 12V car cigarette lighter and 230V, cooling performance 20°C below ambient temperature | N/A | 3 weeks | General Cargo |  |  | 5.6 | 37 | 24 | 50 | 0.0444 | 0 | 122.43 | USD | 0 | Campinggaz 69180, 203 164 | CE |  | 39269097
106244 | Sturdy container | TTM | 1 | Container, Wide Mouth, with Screw Cap HDPE, white Can be used for foodstuffs. A unit includes the container, a transparent inner seal and a lid with tamper-resistant seal and predetermined breaking point, 2 lt, 1pc | N/A | 3 weeks | General Cargo | Product linked to 106245 |  | 0.1 | 13 | 13 | 21.5 | 0.0036335 | 0 | 7.59 | USD | 0 | CurTec 4320-50-048 | CE |  | 39269097
106245 | Material for packing specimen | TTM | 1 | Soft sorbant granulate material for the transport of chemicals and specimens - 1 unit = 8 kg | N/A | 3 weeks | General Cargo | Product linked to 106244 |  | 10 | 10 | 40 | 30 | 0.12 | 0 | 8.7 | USD | 0 | Alex Breuer 65000000.27 | CE |  | 68062090
106237 | Universal test tube rack 16mm | TTM | 1 | Stands for test tubes, made of PP, autoclavable, stackable, white, approx., 246x104x64mm, Ø 16mm | N/A | 3 weeks | General Cargo |  |  | 0.1 | 26 | 12 | 3 | 0.000936 | 0 | 12.99 | USD | 0 | Kleinfeld Labortechnik 7460001 | CE |  | 39269097
106238 | Universal test tube rack 18mm | TTM | 1 | Test tube rack, 50 holes, with alphanumeric markings, Ø 18mm | N/A | 3 weeks | General Cargo |  |  | 0.1 | 24 | 12 | 8 | 0.002304 | 0 | 6.16 | USD | 0 | Herenz 1131963 | CE |  | 39269097
106239 | Universal test tube rack 28mm | TTM | 1 | Test tube rack, 21 holes, with alphanumeric markings, Ø 30mm | N/A | 3 weeks | General Cargo |  |  | 0.2 | 20 | 20 | 10 | 0.004 | 0 | 8.6 | USD | 0 | Brand 43 400 41 | CE |  | 39269097
106286 | Cetylpyridinium chloride monohydrate | TTM | 100g | Cetylpyridinium chloride monohydrate | 47 months | 3 weeks | Dangerous goods |  |  | 0.5 | 28 | 21 | 32 | 0.018816 | 0 | 111.12 | USD | 0 | Sigma-Aldrich C0732-100G | SDS/ CoA | UN 2811 Class 6.1 Packing group II | 38220000
106287 | Sodium chloride | TTM | 1 kg | Sodium chloride for analyses EMSURE® ACS, ISO, Reag. Ph Eur>99% | 35 months | 3 weeks | General Cargo |  |  | 1.1 | 9 | 9 | 18 | 0.001458 | 0 | 21.7 | USD | 0 | Merck 1064041000 | SDS/ CoA |  | 25010091
106283 | N-acetyl-L-cysteine (NALC) | TTM | 100g | N-Acetyl-L-cysteine for biological usage in bottles of 100g | 47 months | 3 weeks | General Cargo |  |  | 0.15 | 5 | 5 | 11 | 0.000275 | 0 | 133.53 | USD | 0 | Merck Millipore 1124220100 | SDS/ CoA |  | 29309016
106284 | tri-Sodium citrate dihydrate | TTM | 1 kg | Tri-Sodium citrate dihydrate, for analyses EMSURE® ACS, ISO, Reag. Ph Eur>99% | 23 months | 3 weeks | General Cargo |  |  | 1.1 | 9 | 9 | 21 | 0.001701 | 0 | 72.6 | USD | 0 | Merck 1064481000 | SDS/ CoA |  | 28520000
106281 | Sodium hydroxide | TTM | 1 | Sodium hydroxide, NaOH, MW 40.00, purum ≥ 98%, pellets, in plastic bottle, tight screw cap (hygroscopic and CO2 binding from the air) for analyses EMSURE® ISO › 99% | 23 months | 3 weeks | Dangerous goods |  |  | 1.1 | 9 | 9 | 21 | 0.001701 | 0 | 34.62 | USD | 0 | Merck 1064981000 | SDS/ CoA | UN 1823 Class: 8 Packing group II | 28151100
106288 | 5-Sulfosalicylic acid dihydrate | TTM | 100 g | 5-Sulfosalicylic acid dihydrate, for synthesis ≥ 99%, 100g plastic bottle | 35 months | 3 weeks | General Cargo |  |  | 0.3 | 10 | 10 | 15 | 0.0015 | 0 | 36.67 | USD | 0 | Merck 8006910100 | SDS/ CoA |  | 29182900
106282 | Tri-Sodium phosphate dodecahydrate | TTM | 1 kg | Tri-sodium phosphate dodecahydrate , NA3PO4·12H2O, MW:380.12 in plastic bottles | 47 months | 3 weeks | General Cargo |  |  | 1.1 | 11.5 | 11.5 | 18 | 0.0023805 | 0 | 34.22 | USD | 0 | Merck 1065721000 | SDS/ CoA |  | 28352930
106280 | Ethanol, 96% | TTM | 1L | Ethanol, 96%, technical grade, C2H6O, MW: 46,07, boiling point: 75 -78 °C | N/A | 3 - 4 weeks | Dangerous goods |  |  | 1.1 | 28.5 | 20.5 | 32 | 0.019 | 0 | 17.5 | USD | 0 | Carl Roth T171.5 |  | UN 1170 Class 3 Packing Group II
106355 | Marker pen | TTM | 1 | Marker pen, black, for all surfaces, resistant against alcohol/ water, writes even on wet surface | N/A | 3 weeks | General Cargo |  |  | 0.1 | 15 | 8 | 4 | 0.00048 | 0 | 1.51 | USD | 0 | Edding 04 400 001 | N/A |  | 39269097
106506 | Marker pen for PCR tubes | TTM | 1 | 0.6mm marker pen, specially for laboratory usage, permanent with Dry safe technology to allow the pen to stay open (Test ISO 554), rechargeable | N/A | 3 weeks | General Cargo |  |  | 0.1 | 10 | 2 | 2 | 4e-05 | 0 | 1.48 | USD | 0 | Staedtler Lumocolor | N/A |  | 39269097
106356 | Adhesive labels | TTM | 1 | Labels, adhesive 50x19mm, packaging 486 pcs per box (10boxes) | N/A | 3 weeks | General Cargo |  |  | 1 | 18.5 | 13.5 | 7 | 0.0017483 | 0 | 50.97 | USD | 0 | Avery Zweckform 3327 | TÜV |  | 48219010
106509 | Special pH indicator strips | TTM | 100 | pH- indication sticks in box of 100 tests, ph. from 6.5-10,0, graduation 0,2-0,5, with colour indication, can not blend as it is in tubes | N/A | 3 weeks | General Cargo |  |  | 0.3 | 10 | 9 | 3 | 0.00027 | 0 | 11.65 | USD | 0 | Merck 1095430001 | SDS |  | 28530090
6. Solid culture: Instruments, reagents and consumables
106247 | Precision Incubator | TTM | 1 | Precision Incubator, ventilation and Control, natural convection, continuous adjustment of pre-heated fresh air admixture, vent connection with restrictor flap, adaptive multifunctional digital PID microprocessor | N/A | 3 weeks | General Cargo | · This item needs to be ordered with a perforated stainless steel shelf (106248); · To be transported upright, cargo aircraft only |  | 290 | 133 | 105 | 192 | 2.6813 | 0 | 5976.24 | USD | 0 | Memmert IN750 inkl. 2 E20182 |  |  | 84198998
106248 | Perforated stainless steel shelf | TTM | 1 | Perforated stainless steel shelf | N/A | 3 weeks | General Cargo | · Product linked to 106247 |  | 15 | 80 | 80 | 4 | 0.0256 | 0 | 188.41 | USD | 0 | Memmert B00328 |  |  | 84198998
106223 | Electric micro incinerator for loops | TTM | 1 | Bacti-Cinerator IV Sterilizer Safely sterilizes loops and needles by infrared heat, Complete sterilization occurs in only 5 to 7 seconds at optimum temperature of 1500°F. Heat is contained in deep ceramic tube to safeguard laboratory personnel—no open flame. No need to wait for instruments to turn red. | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 2 | 27 | 13 | 15 | 0.005265 | 0 | 690.69 | USD | 0 | WLD-TEC 5101000 | CE |  | 84198998
106249 | Hand lens | TTM | 1 | Hand lens model 2642100, lens diameter 100mm, 2X | N/A | 3 weeks | General Cargo |  |  | 0.2 | 25 | 14 | 3 | 0.00105 | 0 | 53.28 | USD | 0 | Eschenbach Optik GmbH 2642100 | CE |  | 90138090
106232 | Loop holder | TTM | 1 | Loop holder, upper part aluminium or stainless steel, with plastic handle, length 240mm, according to Kolle Schwarz | N/A | 3 weeks | General Cargo |  |  | 0.1 | 2 | 2 | 25 | 0.0001 | 0 | 8.05 | USD | 0 | KöhlerW-1402 | CE |  | 73269098
106233 | Rack for loop holder | TTM | 1 | Rack for loop holder, PP, disinfect able, 6x holders | N/A | 3 weeks | General Cargo | · Product linked to 106232 |  | 0.3 | 10 | 10 | 5 | 0.0005 | 0 | 23.84 | USD | 0 | Köhler A-1750 | CE |  | 73269098
106234 | Inoculation loop 4mm | TTM | 1 | These NiCr 8020 loop, lengths 5cm, with internal diameters (Ø) 4 mm, They are heat resistant and can be flamed to sterilise the loop. | N/A | 3 weeks | General Cargo | · Product linked to 106232 |  | 0.01 | 6 | 1.5 | 1.5 | 1.35e-05 | 0 | 0.82 | USD | 0 | RSG Solingen 631-7133 VWR | CE |  | 73269098
106235 | Inoculation loop 3mm | TTM | 1 | Inoculation loops, Platinum/Iridium wire, length 5 cm, wire diameter 0.6, loop diameter 3mm, | N/A | 3 weeks | General Cargo | · Product linked to 106232 |  | 0.005 | 4 | 0.3 | 0.3 | 3.6e-07 | 0 | 38.95 | USD | 0 | Fischer Labortechnik 631-7122 VWR | CE |  | 71101910
106236 | Disposable loops 10 µl | TTM | 500 | VWR disposable loop, flexible 10µl, packaging in bags of 10 pcs, in box of 500 pcs | 24 months | 3 weeks | General Cargo | · Product linked to 106232 |  | 1 | 31 | 27 | 11 | 0.009207 | 0 | 16.28 | USD | 0 | VWR Collection 612-9357 VWR | CE |  | 39269097
106096 | Glass beads, massive glass, 3 mm diameter | TTM | 1 kg | Glass beads 3mm, solid, spherical, clear glass, pack of 1 kg | N/A | 3 weeks | General Cargo |  |  | 1 | 15 | 15 | 10 | 0.00225 | 0 | 6.19 | USD | 0 | VWR 632645103000 | CE |  | 70179000
106097 | Glass beads, massive glass, 5 mm diameter | TTM | 1 kg | Glass beads 5mm, solid, spherical, clear glass, pack of 1 kg | N/A | 3 weeks | General Cargo |  |  | 2.2 | 36 | 27 | 7 | 0.006804 | 0 | 6.7 | USD | 0 | VWR 632645105000 | CE |  | 70179000
106098 | Caped glass tubes for preparing suspension | TTM | 100 | Thick walled pyrex tube, Duran, GL 16 thread, A-DMRXL, packaging 100pcs | N/A | 3 weeks | General Cargo |  |  | 1.6 | 34 | 11 | 14 | 0.005236 | 0 | 46.21 | USD | 0 | Duran Group 391-0145 VWR | CE |  | 70172000
106269 | Sodium-L-Glutamate-monohydrate | TTM | 1 kg | Sodium-L-Glutamate-monohydrate, C5H8NNaO4 • H2O, MW: 187.13, purum, ≥ 98%, 1 kg plastic bottle | 24 months | 3 weeks | General Cargo |  |  | 1.2 | 20 | 20 | 25 | 0.01 | 0 | 127.77 | USD | 0 | Merck 1.06445.1000 | SDS/ CoA |  | 29224200
106263 | Potassium dihydrogen phosphate | TTM | 1 kg | Potassium dihydrogen phosphate, ≥99.5 %, p.a., ACS, Reagent for phosphate, buffer, molecular biology, KH2PO4 , M 136,09 g/mol, D 2,34, mp ~253 °C, WGK1, EG-No. 2319134, CAS-No. [7778-77-0] | 24 months | 3 weeks | General Cargo |  |  | 1.1 | 9.5 | 9.5 | 22 | 0.0019855 | 0 | 37.24 | USD | 0 | Carl Roth 3904.1 | SDS/ CoA |  | 28352400
106264 | Magnesium sulphate-heptahydrate | TTM | 500 g | Magnesium sulphate- heptahydrate for analyses, EMSURE® ACS reagent Ph Eur | 35 months | 3 weeks | General Cargo |  |  | 0.6 | 10 | 10 | 18 | 0.0018 | 0 | 35.78 | USD | 0 | Merck 1058860500 | SDS/ CoA |  | 28332100
106265 | L-Asparagine-monohydrate | TTM | 100g | L-Asparagin-Monohydrate for biochemical usage › 99% | 24 months | 3 weeks | General Cargo |  |  | 0.3 | 10 | 10 | 12 | 0.0012 | 0 | 25.31 | USD | 0 | AppliChem A3721,0100 | SDS/ CoA |  | 29241200
106266 | tri-Magnesium di-citrate nonahydrate | TTM | 100g | Tri-Magnesium di-citrate nonahydrate, Mg3(C6H5O7)2 • 9 H2O, MW: 613,25, ≥ 95% | 35 months | 3 weeks | General Cargo |  |  | 0.3 | 10 | 10 | 12 | 0.0012 | 0 | 41.09 | USD | 0 | Sigma-Aldrich 63067-100G-F | SDS/ CoA |  | 29181500
106267 | Malachite green oxalate | TTM | 100 g | Malachite green oxalate C.I. 42000) for microscopy Certistain® | 35 months | 3 weeks | Dangerous goods |  |  | 0.5 | 10 | 10 | 18 | 0.0018 | 0 | 68.34 | USD | 0 | Merck 1.15942.0100 | SDS/ CoA | UN 2811 Class 6.1 Packing group III | 32041300
106268 | Glycerol | TTM | 1kg | Glycerol for analyses EMSURE®ACS, Reag. Ph Eur | 35 months | 3 weeks | General Cargo |  |  | 1.2 | 20 | 20 | 25 | 0.01 | 0 | 175.25 | USD | 0 | Merck 1040921000 | SDS/ CoA |  | 29054500
106259 | Ready mixed TB / LJ Base | TTM | 500 g | TB-medium according to Löwenstein- Jensen (basis) | 35 months | 3 weeks | General Cargo |  |  | 0.6 | 12.5 | 7 | 19.5 | 0.0017063 | 0 | 59.69 | USD | 0 | BD 244420 | SDS/ CoA |  | 38210000
106261 | Glycerol ≥ 99% | TTM | 1 L | Glycerol for analyses, 99% Reag. Ph Eur, 1 lt | 12 months | 3 weeks | General Cargo |  |  | 1.3 | 9.5 | 9.5 | 23 | 0.0020757 | 0 | 17.02 | USD | 0 | Caelo 7245-1L | SDS/ CoA |  | 29054500
106262 | Sodium pyruvate, 50g; bottle | TTM | 25g | Sodium Pyruvate reagent plus ›= 99% | 24 months | 3 weeks | Keep cool 2-8°C |  |  | 0.025 | 6 | 6 | 11 | 0.000396 | 0 | 30.01 | USD | 0 | Sigma-Aldrich P2256-25g | SDS/ CoA |  | 38220000
106278 | di-Sodium hydrogen phosphate, anhydrous | TTM | 1 kg | di-Sodium hydrogen phosphate anhydrous (offered also with crystal water) + 2water) Na2HPO4, MW: 141.96, p.a. for buffer | 47 months | 3 weeks | General Cargo |  |  | 1.1 | 12 | 12 | 18 | 0.002592 | 0 | 63.26 | USD | 0 | Merck 1065861000 | SDS/ CoA |  | 28352200
106279 | di-Sodium hydrogen phosphate dodecahydrate | TTM | 1 kg | di-Sodium hydrogen phosphate dodecahydrate for analyses EMSURE ® ISO, Reagent Ph Eur | 24 months | 3 weeks | General Cargo |  |  | 1.2 | 20 | 20 | 25 | 0.01 | 0 | 41.22 | USD | 0 | Merck 1065791000 | SDS/ CoA |  | 28352200
106303 | Skim milk powder | TTM | 500 g | Skim milk powder for microbiology | 17 months | 3 weeks | General Cargo |  |  | 0.7 | 20 | 20 | 10 | 0.004 | 0 | 41.17 | USD | 0
106407 | Brain Heart Infusion agar | TTM | 500 g | Brain Heart infusion Agar, for cultivation of fastidious microorganisms 500g | 35 months | 3 weeks | General Cargo |  |  | 0.6 | 9 | 9 | 22 | 0.001782 | 0 | 83.86 | USD | 0 | Carl Roth X915.1 | CE |  | 38210000
106339 | Universal glass bottle | TTM | 1 | Marienfeld, pathological vials, white flint glass, washed with aluminium screw cap, with black rubber seal, 28ml, wide neck. | N/A | 3 weeks | General Cargo |  |  | 0.05 | 3 | 3 | 8 | 7.2e-05 | 0 | 1.39 | USD | 0 | Marienfeld 3615002 | CE |  | 70172000
106581 | Screw-cap for universal bottle | TTM | 4000 | Screw-cap (for pathological vials with number 106339), made of aluminium, with rubber seal, 4000 pieces. | N/A | 3 - 4 weeks | General Cargo |  |  | 12 | 48 | 33 | 28 | 0.044 | 0 | 1120 | USD | 0 | Marienfeld 66999992
106583 | Universal bottle rack | TTM | 1 | Tube / Brush rack, epoxy coated steel wire racks, 5 x 10 places, for tubes I bottles with o 33 mm. | N/A | 3 - 4 weeks | General Cargo |  |  | 0.002 | 3.5 | 18.9 | 6 | 0.004 | 0 | 28.78 | USD | 0 | Heathrow Scientific HS120088
106338 | Culture tubes with screw cap | TTM | 100 | Culture tubes, diameter 16 mm, length 125 mm, wall strength 1mm, thread outside; inner diameter min.13 mm; borosilicate glass with ISO winding, autoclavable screw caps with gasket out of PP or Teflon. Some prefer 18 mm diameter | N/A | 3 weeks | General Cargo |  |  | 6.5 | 46 | 23 | 18 | 0.019044 | 0 | 25.98 | USD | 0 | Max Wiegand | CE |  | 70172000
106432 | Slanted rack | TTM | 1 | Nalgene, PC, white/transparent, inclined rack, angle of either 5° or 20°, 4x10 places, heat resistant up to 135°C, autoclavable, alphanumeric labelling | N/A | 3 weeks | General Cargo | Product linked to 106338 |  | 0.4 | 20 | 10 | 8 | 0.0016 | 0 | 34.4 | USD | 0 | Nalgene 5935-0016 | CE |  | 39269097
106496 | Slanted rack for BD LJ tubes | TTM | 1 | Nalgene, PC, white/transparent, inclined rack, angle of either 5° or 20°, 4x10 places, heat resistant up to 135°C, autoclavable, alphanumeric labelling | N/A | 3 weeks | General Cargo |  |  | 0.4 | 20 | 10 | 8 | 0.0016 | 0 | 34.4 | USD | 0 | Nalgene 59350020 | CE |  | 39269097
106038 | BBL LJ slants w. Glyc. & PACT | BD | 100 | Lowenstein-Jensen Medium + PACT as an egg-based medium containing antimicrobials for the cultivation of Mycobacterium tuberculosis and other mycobacterial species. | 7 months | 8 weeks | Keep cool 2-8°C |  |  | 4.04 | 35.2 | 19.8 | 16 | 0.011151 | 0 | 125 | USD | 0 | BD 220502 | IVD Directive 98/79/EC |  | 3821000000
106039 | BD BBL™ Lowenstein-Jensen Medium, slants (100/sp) | BD | 100 | Lowenstein-Jensen Medium is an egg-based medium used for the isolation and cultivation of mycobacteria. The medium tubed as deeps is used for the semi quantitative catalase test as an aid to the classification of mycobacteria. Tube size A | 11 months | 8 weeks | Keep cool 2-8°C |  |  | 4.1 | 35.2 | 19.8 | 16 | 0.011151 | 0 | 90 | USD | 0 | BD 220909 | IVD Directive 98/79/EC |  | 3821000000
106273 | BD BBL™ Middlebrook 7H9 Broth with Glycerol, 5 mL (10/sp) | BD | 1 | Middlebrook 7H9 Broth with Glycerol is a non-selective liquid culture medium for the cultivation of mycobacteria, including M. Tuberculosis. It is used primarily for growth of pure cultures of mycobacteria for use in laboratory studies. Prepared tubes, 5 ml. | 7 months | 8 weeks | Keep cool 2-8°C |  |  | 0.2 | 11.6 | 5.2 | 10.8 | 0.00065 | 0 | 25.5 | USD | 0 | BD 221832 | IVD Directive 98/79/EC
106558 | Bottle Middlebrook 7H9 Broth 500 G - 271310 | BD | 500g | BD Difco™ Middlebrook 7H9 Broth, Use with Tween™ 80 or Glycerol and Middlebrook ADC Enrichment for cultivation of mycobacteria and preparation of tubercle emulsion for susceptibility testing, 500 g (1 ea) | 38 months | 8 weeks | General Cargo |  |  | 0.54 | 12.1 | 7 | 15.9 | 0.00134 | 0 | 42.04 | USD | 0 | BD 271310
106559 | BD BBL™ Middlebrook OADC Enrichment, 100 mL per bottle (6/sp) - 212240 | BD | 6x100ml | BD BBL™ Middlebrook OADC Enrichment, 100 mL per bottle (6/sp) - 212240 | 7 months | 8 weeks | Keep cool 2-8°C |  |  | 1.36 | 20.2 | 14 | 14.8 | 0.004185 | 0 | 234.15 | USD | 0 | BD 212240
106614 | Middlebrook 7H11 Agar Base - 212203 | BD | 500g | BD BBL™ Prepared Plated Media 7H11 Agar Base, 500 g | 31 months | 8 weeks | General Cargo |  |  | 0.6 | 12 | 7 | 20 | 0.001621 | 0 | 93.33 | USD | 0 | BD 212203
106615 | Middlebrook 7H10 Agar - 262710 | BD | 500g | BD Difco™ Middlebrook 7H10 Agar, Use with Glycerol and Middlebrook OADC Enrichment for isolation and cultivation of mycobacteria, 500 g | 38 months | 8 weeks | General Cargo |  |  | 0.59 | 12 | 6.5 | 19.5 | 0.00152 | 0 | 83.87 | USD | 0 | BD 262710
106620 | Lowenstein-Jensen Medium - 10 slants - large-size slant - 220908 | BD | 10 | 220908 - BD BBL™ Lowenstein-Jensen Medium, Slants (10/sp) | 11 months | 8 weeks | Keep cool 2-8°C |  |  | 0.28 | 14 | 6.6 | 15.6 | 0.001441 | 0 | 15.7 | USD | 0 | BD 220908 | IVD Directive 98/79/EC
106621 | Seven H11 Agar (Deep Fill) Plates - 221870 | BD | 10 | 221870 - BD BBL™ Prepared Plated Media 7H11 Agar (Deep Fill) (10/sp) | 4 months | 8 weeks | Keep cool 2-8°C | This item has a very short shelf life. Need to order according to consumption. |  | 0.34 | 10.4 | 10.4 | 17.2 | 0.00186 | 0 | 13.6 | USD | 0 | BD 221870 | IVD Directive 98/79/EC
7. Liquid culture: Instruments, reagents and consumables
106013 | BACTEC MGIT 960 System | BD | 1 | BD BACTEC™ MGIT™ 960 System: Non-radiometric, fully automated system (annual capacity of 8,300 tubes) for the rapid detection of mycobacteria in clinical specimens, other than blood, as well as antimicrobial susceptibility testing of anti-tuberculosis drugs known as S.I.R.E., I.R., and PZA. This product includes: Starter Kit (440107), and AST Starter Kit (445941) | N/A | 8 weeks | General Cargo Large item | · 1 year warranty; · Printer (106016) not included with BACTEC system, to be procure separately. |  | 421.84 | 100.5 | 160 | 157.5 | 2.5326 | 0 | 38950 | USD | 0 | BD 445870 | IVD Directive 98/79/EC |  | 9027809100
106408 | Installation and Training - BACTEC 960 | BD | 1 | Installation and Training - BACTEC 960 | N/A | On demand | Service |  |  | N/A | N/A | N/A | N/A | N/A | N/A | 3500 | USD | 0 | N/A | N/A |  | N/A
106596 | BACTEC MGIT 320 System | BD |  | BD BACTEC™ MGIT™ 320 System: Non-radiometric, fully automated system (annual capacity of 2,700 tubes) for the rapid detection of mycobacteria in clinical specimens, other than blood, as well as antimicrobial susceptibility testing of anti-tuberculosis drugs known as S.I.R.E., I.R., and PZA. This product includes: Starter Kit (440107), and AST Starter Kit (445941). | N/A | 8 weeks | General Cargo Large item | · The BACTEC MGIT320 contains only one drawer holding 320 MGIT tubes, for a throughput of approximately 2700 specimens per year; · Printer (106016) not included with BACTEC system, to be procure separately. · 1 year warranty; · Leaflet: http://www.bd.com/resource.aspx?IDX=11570 |  | 203.6 | 118.5 | 92.5 | 86.5 | 0.94815 | 0 | 19500 | USD | 0 | BD 441743 | IVD Directive 98/79/EC |  | 9027809100
106081 | UPS - 3000 VA - 2100W | TTM | 1 | On-Line UPS Vanguard VGD-3000: UPS capacity 3000 VA 230V sine wave output, 50/60 Hz outlets 4x Iec 320, surge protection RJ11/RJ45, 20-30 min buffer, dimensions 35x22x40 cm | N/A | 3 weeks | General Cargo special provision A67 | · 2 years warranty · Product linked to 106013, 106596; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 36.5 | 52 | 35 | 47 | 0.08554 | 0 | 1318.88 | USD | 0 | PCM VGD-3000 | N/A | General Cargo Valve Regulated Lead Acid Battery UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 85359000
106015 | Starter kit: BACTEC | BD | 1 | Product includes: Manual Users BACTEC MGIT960; Quick Ref Guide MGIT960 RAW; Label Barcode BBL MGIT Tube; Station Plug Detail; Temperature QC Device; Kit Calibrator MGIT960 51 Pk; Cord 17255 Belden; Cordset 3 Cond Detachable; Cordset UK 6A/250V Black; Filter Air 5.5in X 5.5in; Filter Intake; Software Bactec MGIT USB Packaged; Carton 14.25in x 11.12 x 7.5in; Bubble Bag 4in x 7-1/2in; Label Drawer Needs Service; Label All Instrument (Label MGIT960 Tube Brkg Warn, Label MGIT960 Tube Brkg Jap, Label MGIT960 Tube Brkg Span, Label MGIT960 Tube Brkg Ital, Label MGIT960 Tube Brkg Germ, Label MGIT960 Tube Brkg French) | 15 months | 8 weeks | General Cargo | Included with the BACTEC |  | 4.663 | 19 | 10.1 | 1.5 | 0.00028 | 0 | 0 | USD | 0 | BD 440107 |  |  | 9027908000
106016 | Printer | BD | 1 | PRINTER BROTHER L5000D | N/A | 8 weeks | General Cargo |  |  | 11.96 | 45 | 48 | 38 | 0.10819 | 0 | 532 | USD | 0 | BD 257613 |  |  | 8443321090
106473 | Filters, Air MGIT 960 (square) | BD | 1 | Filters, Air MGIT 960 (square) - 444374 | N/A | 8 weeks | General Cargo |  |  | 0.13 | 26 | 21 | 6 | 0.00327 | 0 | 52.2 | USD | 0 | BD 444374
106474 | Filters, Air MGIT 960 (rectangular) | BD | 1 | Filters, Air MGIT 960 (rectangular) - 445888 | N/A | 8 weeks | General Cargo |  |  | 0.1 | 30.4 | 18.4 | 5 | 0.00279 | 0 | 70.43 | USD | 0 | BD 445888
106272 | BD BBL Calibrators Kit (for 1 drawer) (1 ea) | BD | 1 | BD BBL™ Calibrators Kit (for 1 drawer) (1 ea) | 15 months | 8 weeks | General Cargo | For replenishment: Need to order 3 units per BACTEC 960 system: 1 unit is for 1 drawer |  | 0.36 | 15 | 10.5 | 11.7 | 0.00184 | 0 | 155.75 | USD | 0 | BD 445871 | IVD Directive 98/79/EC
106018 | AST Carrier Set (5-tube) - 3 carriers | BD | 3 | AST Carrier Set (5-tube) - 3 carriers | N/A | 8 weeks | General Cargo |  |  | 0.0375 | 14.6 | 3.8 | 4.4 | 0.000244 | 0 | 36 | USD | 0 | BD 445943 |  |  | 3923909000
106452 | BACTEC MGIT 960 AST transport rack | BD | 1 | BACTEC MGIT 960 AST transport rack – 445942. | N/A | 8 weeks | General Cargo |  |  | 0.945 | 46.8 | 22.6 | 12.8 | 0.01353 | 0 | 135 | USD | 0 | BD 445942
106026 | BBL MGIT tubes for use in Bactec MGIT 960 (7ml) - 100 tubes/pkg | BD | 100 | The BACTEC™ MGIT™ Tube is intended for the detection and recovery of mycobacteria using the BD BACTEC™ MGIT™ 960 system. Each plastic tube has a screw-top cap and contains 7 ml of modified Middlebrook 7H9 Broth base, which is supplemented with BD BACTEC™ MGIT™ Growth Supplement (Cat. No. 245124). Acceptable specimen types are digested and decontaminated clinical specimens (except urine) and sterile body fluids (except blood). additionally, these tubes can also be used for antimicrobial susceptibility testing of mycobacteria, including SIRE and PZA testing. | 11 months | 8 weeks | General Cargo |  |  | 1.86 | 30.4 | 16.8 | 12.2 | 0.006231 | 0 | 195 | USD | 0 | BD 245122 | IVD Directive 98/79/EC |  | 3822000000
106027 | BACTEC MGIT 960 Supplement Kit (100 tests, PANTA and OADC combined) | BD | 100 | 1 pack is 6 tubes of 15 ml OADC and 6 tubes of PANTA The kit includes 6 vials of lyophilised BBL™ MGIT™ PANTA™ Antibiotic Mixture and 6 vials of 15 ml BD BACTEC™ MGIT™ 960 Growth Supplement. The BD BACTEC™ MGIT™ Growth Supplement contains Middlebrook OADC enrichment (Oleic acid, Bovine albumin, Dextrose, Catalase and Polyoxyethylene stearate [POES]). The BBL™ MGIT™ PANTA™ Antibiotic Mixture contains Polymyxin B, Amphotericin B, Nalidixic acid, Trimethoprim and Azlocillin. | 11 months | 8 weeks | Keep cool 2-8°C | Need to order one unit of MGIT tubes (106026) |  | 0.54 | 23.2 | 16.4 | 7 | 0.002663 | 0 | 70.8 | USD | 0 | BD 245124 | IVD Directive 98/79/EC |  | 3822000000
106028 | BACTEC MGIT 960 SIRE kit, One kit is sufficient for 40 test | BD | 40 | The BD BACTEC™ MGIT™ 960 SIRE Kit is a rapid qualitative procedure for susceptibility testing of Mycobacterium tuberculosis, from culture, to streptomycin (STR), isoniazid (INH), rifampicin (RIF) and ethambutol (EMB). | 11 months | 8 weeks | Keep cool 2-8°C | · For each box of SIRE kits, you need 2 boxes of MGIT Tubes (106026) |  | 0.74 | 22.4 | 15 | 7.4 | 0.002486 | 0 | 72.7 | USD | 0 | BD 245123 | IVD Directive 98/79/EC |  | 3822000000
106037 | MycoPrep Specimen Digestion / Decontamination Kit - 10 bottles of 150ml | BD | 1 | This kit contains reagents necessary for the preparation of respiratory and no respiratory (blood excluded) specimens for identification of mycobacteria. The kit includes bottles of NaOH-citrate solution (containing a plastic ampule with NALC) and powdered phosphate buffer. After activation of the NALC ampule, the NALC-NaOH mixture is stable for 24 hours. | 7 months | 8 weeks | Dangerous goods | This item has a very short shelf life. Need to order according to consumption. |  | 2.35 | 21 | 21 | 18 | 0.007938 | 0 | 81.5 | USD | 0 | BD 240863 | IVD Directive 98/79/EC | UN 1824 Class 8 Packing group III | 3821000000
106623 | BBL MGIT OADC Enrichment | BD | 1 | OADC to be ordered with lyophilized SLDs for DST on liquid culture. 6 vials of 15 ml (90ml)- 245116 | 11 months | 8 weeks | Keep cool 2-8°C |  |  | 0.06 | 17.6 | 10.6 | 8.2 | 0.00152 | 0 | 55 | USD | 0 | BD 245116 | CE |  | 3821000000
106582 | BD BBL Middlebrook OADC Enrichment, 20 mL per tube, ''A'' size tube (10/sp) | BD | 10 | BD BBL™ Middlebrook OADC Enrichment, 20 mL per tube, ''A'' size tube (10/sp) - 211886 | 7 months | 8 weeks | Keep cool 2-8°C | This item has a very short shelf life. Need to order according to consumption. |  | 0.5 | 14 | 6.6 | 15.6 | 0.001441 | 0 | 65.8 | USD | 0 | BD 211886 |  |  | 3821000000
106017 | BD’s MGIT Fluorescence Manual Tube Reader | BD | 1 | The BD BACTEC™ MicroMGIT Fluorescence Reader is used for the qualitative reading of fluorescence in Manual MGIT™ Tubes (4 ml). The reader has a standard 9 volt battery, a low voltage indicator, and dimensions of 9.2 cm x 14.5 cm x 12.0 cm (W x D x H). A calibration tube is included. | 15 months | 8 weeks | General Cargo |  |  | 0.76 | 27.8 | 20.6 | 18.2 | 0.010423 | 0 | 850 | USD | 0 | BD 445923 | IVD Directive 98/79/EC |  | 9027809100
106019 | AST Carrier Set (4-tube) - 3 carriers | BD | 3 | AST Carrier Set (4-tube) - 3 carriers | N/A | 8 weeks | General Cargo |  |  | 0.0261 | 16.7 | 10.1 | 3 | 0.000506 | 0 | 36 | USD | 0 | BD 445944 |  |  | 3923909000
106020 | AST Carrier Set (3-tube) - 3 carriers | BD | 3 | AST Carrier Set (3-tube) - 3 carriers | N/A | 8 weeks | General Cargo |  |  | 0.0187 | 16.7 | 10.1 | 3 | 0.000506 | 0 | 36 | USD | 0 | BD 445945 |  |  | 3923909000
106021 | AST Carrier Set (2-tube) - 3 carriers | BD | 3 | AST Carrier Set (2-tube) - 3 carriers | N/A | 8 weeks | General Cargo |  |  | 0.0314 | 30.4 | 18.4 | 4.5 | 0.00279 | 0 | 36 | USD | 0 | BD 445946 |  |  | 3923909000
106022 | AST Carrier Set (8-tube) - 3 carriers | BD | 3 | AST Carrier Set (8-tube) - 3 carriers | N/A | 8 weeks | General Cargo |  |  | 0.0575 | 25.4 | 15.2 | 3.1 | 0.001197 | 0 | 36 | USD | 0 | BD 445993 |  |  | 3923909000
106274 | BD BBL AST Starter Group (1 ea.) | BD | 1 | This item includes: • User Manual; Quick reference guide • 2 x Transport rack • 16 x 5-tube AST Set carrier • 16 x 2-tube AST Set carrier • 3 x 3-tube AST Set carrier • 3 x 4-tube AST Set carrier and • 3 x 8-tube AST Set carrier | N/A | 8 weeks | General Cargo | This item is included with the BACTEC system |  | 2.8200 | 44.1 | 39.5 | 28.9 | 0.05033 | 0 | 0 | USD | 0 | BD 445941
106616 | BD BACTEC MGIT 960 - Temperature QC Tube - 445872 | BD | 1 | BD BBL™ QC Temperature Tube (1 ea.) | N/A | 8 weeks | General Cargo |  |  | 0.1545 | 20 | 1.8 | 8.6 | 0.00237 | 0 | 95.62 | USD | 0 | BD 445872
106617 | BD BACTEC MGIT 960 - Spare bar code Set carrier - 445959 | BD | 1 | BD BBL™ AST Carrier Set, Spare Barcodes (1/ca) | N/A | 8 weeks | General Cargo |  |  | 0.0111 | 19.6 | 13.3 | 0.5 | 0.00013 | 0 | 58.2 | USD | 0 | BD 445959
106618 | BD BACTEC MGIT 960 - Plug for Bad Station- 445873 | BD | 10 pack | BBL™ Plug, Bad Station (10/sp) | N/A | 8 weeks | General Cargo |  |  | 0.016 | 1.5 | 12.5 | 19 | 0.00035 | 0 | 640 | USD | 0 | BD 445873
106619 | MicroMGIT Calibrator - 441049 | BD | 1 | TUBE CALIBRATOR MICRO MGIT (1 ea.) | 15 months | 8 weeks | General Cargo |  |  | 0.1247 | 20 | 13.8 | 8.6 | 0.00237 | 0 | 150.95 | USD | 0 | BD 441049 | IVD Directive 98/79/EC
106033 | BACTEC MGIT PZA Kit | BD | 50 | The BD BACTEC™ MGIT™ 960 PZA Kit is used for susceptibility testing of Mycobacterium tuberculosis in culture to pyrazinamide (PZA) in a qualitative test lasting 4-17 days. The Kit contains 2 vials of lyophilised antimicrobic and 6 vials of SIRE Supplement. PZA must be reconstituted with 2.5 ml of sterile/deionized water before addition to BBL™ MGIT™ tubes. Carton of 2 lyophilised vials and 6 PZA Supplements, sufficient for 50 tests. | 11 months | 8 weeks | Keep cool 2-8°C |  |  | 0.46 |  |  |  | 0.00153 | 0 | 90.6 | USD | 0 | BD 245128 | IVD Directive 98/79/EC |  | 3822000000
106036 | BACTEC MGIT PZA Tubes - 25 | BD | 25 | The BD BACTEC™ MGIT™ 960 PZA Medium is a tube containing modified Middlebrook 7H9 Broth with a reduced pH of 5.9 and is to be used for susceptibility testing of Mycobacterium tuberculosis. | 11 months | 8 weeks | General Cargo |  |  | 0.42 | 12.4 | 11.4 | 12.6 | 0.001781 | 0 | 53.55 | USD | 0 | BD 245115 | IVD Directive 98/79/EC |  | 3821000000
106205 | UV-light illuminator - Liquid culture - 365 nm | TTM | 1 | Spectroline, ETX-20L, Trans illuminator long life UV filter glass , minimal UV damage 312nm and 365 nm virtually eliminate transmission of 254nm radiation, variable intensity control from 100% down to 20% for maximum sample preparation time and sensitivity 15 Watt unit | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 6 | 30 | 38 | 62 | 0.07068 | 0 | 1685.61 | USD | 0 | PEQLAB 60-ETX-20L | CE |  | 85394900
106451 | Densitometer | TTM | 1 set | Densitometer, MiniDen, designed to measure cell suspensions turbidity in the range of 0.3-0.5 McFarland units (100x109 cells/ml - 150x109 cells/ml). With adaptor 16mm and set of McFarland stds | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 1 | 24.5 | 24.5 | 15.5 | 0.0093039 | 0 | 680 | USD | 0 | BioSan: 2001, Z0026 Remel: R20421 | CE |  | 90275000
106468 | Mc Farland Standards kits | TTM | 1 | 1 set of Mc Farland standards (0.5, 1.0, 2.0, 3.0,4.0) | 18 months | 3 weeks | General Cargo |  |  | 0.2 | 15 | 10 | 5 | 0.00075 | 0 | 228.91 | USD | 0 | Remel R20421 | SDS |  | 38220000
106072 | 1-channel pipette, variable from 20 -200 ul - LC | TTM | 1 | Gilson pipetman NEO P200 N (20-200 μl) ISO 8655, Chemical and UV resistant, resistant to disinfectant, autoclavable | N/A | 3 weeks | General Cargo |  |  | 0.2 | 32 | 11 | 4 | 0.001408 | 0 | 226.94 | USD | 0 | Gilson F144565 | CE |  | 84798997
106075 | 1-channel pipette, variable from 200 - 1000 µl - LC | TTM | 1 | Gilson pipetman NEO P1000 N (200-1000 μl) ISO 8655, Chemical and UV resistant, resistant to disinfectant, autoclavable | N/A | 3 weeks | General Cargo |  |  | 0.2 | 32 | 11 | 4 | 0.001408 | 0 | 226.94 | USD | 0 | Gilson F144566 | CE |  | 84798997
106419 | Tips, PP, 20-200 μl, sterile, autoclavable with Filter | TTM | 960 | Pipette filter tips PP, capacity 1-200µl, free from DNA/RNAse and ATP, autoclavable, sterile, 10 racks of 96 pcs | 35 months | 3 weeks | General Cargo | Product linked to 106072 |  | 1.5 | 33 | 20 | 13.5 | 0.00891 | 0 | 52.66 | USD | 0 | AHN Biotechnologie 2-113-96-0 | CE |  | 39269097
106420 | Tips, PP, 100-1000 μl, sterile, autoclavable with Filter | TTM | 960 | Pipette filter tips PP, capacity 100-1000µl, free from DNA/RNAse and ATP, length 89mm, autoclavable, sterile, 10 racks of 96 pcs | 35 months | 3 weeks | General Cargo | Product linked to 106075 |  | 4 | 80 | 60 | 32 | 0.1536 | 0 | 58.97 | USD | 0 | AHN Biotechnologie 2-201-96-0 | CE |  | 39269097
106115 | Hand-dispenser pipette, 5 dosage steps, suited for 10ml Tips / Eppendorf | TTM | 1 | Eppendorf Hand-dispenser pipette, Multipipette ® Plus | N/A | 3 weeks | General Cargo |  |  | 1.2 | 43 | 17 | 15 | 0.010965 | 0 | 440.72 | USD | 0 | Eppendorf 613-2890 VWR |  |  | 84798997
106180 | Multipette® stream, includes charging adapter | TTM | 1 | Multi pipette ®stream: Automatic Combitip ®plus tip recognition and volume recognition, volume range from 1µ-50ml, illuminated easy to - read display, 1 button tip rejector, rapid, ergonomic dispensing with motorized piston, high precision while handling | N/A | 3 weeks | General Cargo |  |  | 2 | 20 | 20 | 35 | 0.014 | 0 | 781.13 | USD | 0 | Eppendorf |  |  | 84798997
106181 | Dispenser-tips Universal 10ml sterile | TTM | 100 | Combitips plus 10ml Biopur sterile | 35 months | 3 weeks | General Cargo |  |  | 1.2 | 30 | 25 | 25 | 0.01875 | 0 | 197.34 | USD | 0 | Eppendorf 0030089677 | SDS |  | 39269097
106085 | Electricity net independent automatic pipetting aid | TTM | 1 | Accu automatic pipetting assistance. Adjustable 5 levels, incl. Charger. | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 1 | 24 | 23 | 8 | 0.004416 | 0 | 462.92 | USD | 0 | Hirschmann 9907200 | CE |  | 84223000
106209 | Serological pipette - 5 ml | TTM | 200 | Serological pipette -5,0 ml, graduation 0.1 ml, packaging single packed, 200 pcs per box | 24 months | 3 weeks | General Cargo |  |  | 2.3 | 40 | 21 | 19 | 0.01596 | 0 | 57.55 | USD | 0 | Corning 734-1737 VWR | CE |  | 39269097
106550 | Serological pipette - 10 ml | TTM | 200 | Serological pipette - 10 ml, graduation 0.1 ml, packaging single packed, 200 pcs per box | 24 months | 3 weeks | General Cargo |  |  | 9 | 35 | 25 | 20 | 0.0175 | 0 | 57.56 | USD | 0 | Corning 734-1738 VWR | CE |  | 39269097
106394 | Single use plastic Pasteur-pipettes sterile individually packed | TTM | 500 | Pipette Pasteur Graduation 0,5ml, Ster/10, individually packed in packaging of 500 pcs. Length 150 mm | 35 months | 3 weeks | General Cargo |  |  | 1.5 | 36 | 26 | 24 | 0.022464 | 0 | 44.64 | USD | 0 | Ratiolab 26 55 111 | CE |  | 39269097
106343 | Single use plastic Pasteur-pipettes sterile pack by 10 | TTM | 500 | Pipette Pasteur Graduation 0,5ml, Ster/10, packaging in bags of 10pcs in box of 500 pcs. Length 150 mm | N/A | 3 weeks | General Cargo |  |  | 1.5 | 36 | 36 | 24 | 0.031104 | 0 | 44.64 | USD | 0 | Ratiolab 26 55 111 | CE |  | 39269097
106351 | Single use syringes, sterile | TTM | 50 | Single use syringes, sterile 20ml individual packaging in box 50 | 35 months | 3 weeks | General Cargo |  |  | 0.8 | 43 | 17 | 11 | 0.008041 | 0 | 3.61 | USD | 0 | Dispomed 20020 | CE |  | 90183110
106352 | Syringe filter for single use, sterile | TTM | 50 | ACRODISC LC PVDF 0.2µm 25mm | 35 months | 3 weeks | General Cargo |  |  | 0.3 | 15.5 | 10.5 | 20.5 | 0.0033364 | 0 | 120.87 | USD | 0 | Pall Life Sciences 514-4126 VWR | CE |  | 84213920
106241 | Measuring pipette, short to work inside BSC, sterile, 1 ml | TTM | 1 | Measuring pipette, 1ml, subdivision 0.1ml | N/A | 3 weeks | General Cargo |  |  | 0.2 | 26 | 8 | 7 | 0.001456 | 0 | 4.02 | USD | 0 | Hirschmann 1170135 612-1087 VWR | CE |  | 70179000
106342 | PP-tubes for centrifuge, sterile, 15 ml | TTM | 500 | PP-tubes for centrifuge, 15 ml with screw cap, tight, High quality, crystal-clear polypropylene centrifuge tubes. Centrifuge tubes can be safely centrifuged to approximately 3,200 × g (50 mL size) and 3,500 × g (15 mL size) depending on the specific gravity of the solution being centrifuged and the type of rotor. | 35 months | 3 weeks | General Cargo |  |  | 2.5 | 25 | 15 | 10 | 0.00375 | 0 | 85.34 | USD | 0 | Greiner Bio-One 188271 | CE |  | 39269097
106341 | PP-tubes for centrifuge, sterile, 50 ml | TTM | 500 | Centrifuge tubes 50ml, super clear™, PP aor Ps, conical, printer graduation, extra thick walls to allow centrifugation up to 15000g, with extra large, solvent resistant, white labelling area and black graduations flat PE caps can close with quick 3/4 turn, allowing one handed operation, patented earth friendly® racks can be labelled with lab markers, sterile tubes are radiation sterilised to ensure non/cytotoxicity, 1 rack contains 25 tubes | 35 months | 3 weeks | General Cargo |  |  | 8.05 | 60 | 40 | 37 | 0.0888 | 0 | 74.1 | USD | 0 | Greiner Bio-One 227261 | CE |  | 39269097
106230 | Test tube rack 18/8 stainless steel 18 mn | TTM | 1 | Test tube rack stainless steel, 3x12 places, 18mm | N/A | 3 weeks | General Cargo | Product linked to 106338 |  | 0.2 | 25 | 8 | 7 | 0.0014 | 0 | 25.98 | USD | 0 | Hupfer Metallwerke 212-6475 VWR | CE |  | 73269098
106114 | Petri-dishes glass for reuse | TTM | 72 | Petri dish absolutely plain, heat resisting glass, two parts incl. cover, superior quality, art no. 1070506. 94 x 16 mm | N/A | 3 weeks | General Cargo |  |  | 4.3 | 32 | 50 | 40 | 0.064 | 0 | 95.31 | USD | 0 | Hecht 1042/10 | CE |  | 70179000
106113 | Petri-dishes plastic | TTM | 480 | Petri dishes with vents, PS, crystal clear packed machine sterile, packed in 20pcs/bag in box of 480pcs. 94 x 16 mm | N/A | 3 weeks | General Cargo |  |  | 7.5 | 59 | 40 | 38 | 0.08968 | 0 | 37.41 | USD | 0 | Greiner Bio-One 633180 Kobe | CE |  | 39269097
106475 | Slide warmer | TTM | 1 | Slide warmer with black anodised working plate, temp. Up to 100°C, regulation in 1°C steps, digital display, overheating protection, size WxDxH 160x290x30mm | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 3.2 | 20 | 32 | 30 | 0.0192 | 0 | 1493.206 | USD | 0 | Medax 18801 | CE |  | 84798200
106289 | Para-nitrobenzoic acid | TTM | 1 | 4-Nitrobenzoic acid for synthesis, 100 g | 12 months | 3 - 4 weeks | General Cargo |  |  | 0.2 | 10 | 10 | 16 | 0.0016 | 0 | 28.31 | USD | 0 | Merck 8.00646.0000
8. Solid and liquid culture DST: First-line drug pure substances
106307 | Dimethyl sulfoxide | TTM | 250ml | Dimethyl sulfoxide, for analyses > 99,5%, packaging 250ml | 35 months | 3 weeks | General Cargo |  |  | 18 | 10 | 10 | 18 | 0.0018 | 0 | 15.46 | USD | 0 | Carl Roth A994.2 | SDS/ CoA |  | 29309085
106308 | Isoniazid (INH) - for laboratory use | TTM | 5 g | 5g Isoniazid (INH) Free Base | 24 months | 3 weeks | General Cargo | · Powder presentation; needs to be prepared manually. Follow manufacturers instructions |  | 12 | 10 | 10 | 12 | 0.0012 | 0 | 26.55 | USD | 0 | Sigma-Aldrich I3377-5G | SDS/ CoA |  | 38220000
106309 | Rifampicin (RMP) - for laboratory use | TTM | 1 g | RIFAMPICIN R 3501 | 71 months | 3 weeks | Keep cool 2-8°C | · About 1000 test done with 1 gr; · Powder presentation; needs to be prepared manually. Follow manufacturers instructions |  | 12 | 10 | 10 | 12 | 0.0012 | 0 | 152.28 | USD | 0 | Sigma-Aldrich R3501-1G | SDS/ CoA |  | 38220000
106310 | Ethambutol - for laboratory use | TTM | 25 g | Ethambutol | 47 months | 3 weeks | General Cargo | · Powder presentation; needs to be prepared manually. Follow manufacturers instructions |  | 25 | 15 | 20 | 25 | 0.0075 | 0 | 158.03 | USD | 0 | Sigma-Aldrich E4630-25G | SDS/ CoA |  | 29221985
106311 | Dihydro-streptomycin (DHSM) - for laboratory use | TTM | 5g | Dihydro-streptomycin | 24 months | 3 weeks | Keep cool 2-8°C | · Powder presentation; needs to be prepared manually. Follow manufacturers instructions |  | 12 | 10 | 10 | 12 | 0.0012 | 0 | 65.61 | USD | 0 | Sigma-Aldrich D7253-5G | SDS/ CoA |  | 29412030
106312 | Nicotinic acid amide - for laboratory use | TTM | 100g | Nicotinic acid amide | 35 months | 3 weeks | General Cargo | · Powder presentation; needs to be prepared manually. Follow manufacturers instructions |  | 13 | 7 | 7 | 13 | 0.000637 | 0 | 45.3 | USD | 0 | Calbiochem -Merck 481907-100 VWR | SDS/ CoA |  | 29362990
9. Solid and liquid culture DST: Second-line drug pure substances
106313 | Ofloxacin (OFL) - for laboratory use | TTM | 10g | Ofloxacin | 12 months | 3 weeks | Keep cool 2-8°C | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.015 | 4 | 4 | 8.5 | 0.000136 | 0 | 315.39 | USD | 0 | Sigma-Aldrich O8757-10G
106314 | Moxifloxacin (MOX) - for laboratory use | TTM | 1g | Moxifloxacin | 24 months | 3 weeks | General Cargo | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.032 | 5 | 4.5 | 6 | 0.000135 | 0 | 368.69 | USD | 0 | Molekula M47682403 | SDS/ CoA |  | 38220000
106315 | Potionamid (PTH) - for laboratory use | TTM | 2 g | prothionamide | 24 months | 3 weeks | General Cargo | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.03 | 2.5 | 2.5 | 6 | 3.75e-05 | 0 | 128.27 | USD | 0 | Chemos 121876 | SDS/ CoA |  | 29333999
106316 | Ethionamide (ETH) - for laboratory use | TTM | 5 g | Ethionamide | 24 months | 3 weeks | Keep cool 2-8°C | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.7 | 4 | 4 | 8 | 0.000128 | 0 | 112.97 | USD | 0 | Sigma-Aldrich E6005-5G | SDS/ CoA |  | 29333999
106317 | Kanamycin (KAN) - for laboratory use | TTM | 5 g | Kanamycin, for biochemical usage | 47 months | 3 weeks | Keep cool 2-8°C | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.05 | 3 | 3 | 6.5 | 5.85e-05 | 0 | 67.22 | USD | 0 | Merck Millipore 1.05177.0005 | SDS/ CoA |  | 29419000
106318 | Amikacin (AMI) - for laboratory use | TTM | 5 g | Amikacin, free base | 17 months | 3 weeks | Keep cool 2-8°C | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.05 | 4 | 4 | 8 | 0.000128 | 0 | 689.73 | USD | 0 | Sigma-Aldrich A3650-5G | SDS/ CoA |  | 29419000
106319 | Capreomycin (CM) - for laboratory use | TTM | 1 g | Capreomycin sulphate from Streptomy CES*CAPREOLUS | 17 months | 3 weeks | Keep cool 2-8°C | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.1 | 3.5 | 3.5 | 6 | 7.35e-05 | 0 | 300.12 | USD | 0 | Sigma-Aldrich C4142-1G | SDS/ CoA |  | 29419000
106320 | Para-aminosalicylic acid (PAS) - for laboratory use | TTM | 1 | 4-Aminosalicylic acid, 98+% - 5G | 24 months | 3 weeks | General Cargo | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.8 | 15 | 15 | 20 | 0.0045 | 0 | 22 | USD | 0 | Sigma-Aldrich A79604-5g | SDS/ CoA |  | 29049050
106560 | Levofloxacin - for laboratory use | TTM | 1 g | Levofloxacin ≥98.0% (HPLC) CAS Number 100986-85-4 - Empirical Formula (Hill Notation) C18H20FN3O4 - Molecular Weight 361.37 | N/A | 3 - 4 weeks | General Cargo | · The quantities are essential: with 1 g about a 100 tests can be done in solid or liquid culture; · Powder presentation, needs to be prepared manually. Follow manufacturers instructions | [6 empty] | 0 | 48.23 | USD | 0 | Sigma-Aldrich 28266-1G-F |  |  | 29349990900
106561 | D-Cycloserine | TTM | 250 mg | D-Cycloserine - synthetic CAS Number 68-41-7 - Empirical Formula (Hill Notation) C3H6N2O2 - Molecular Weight 102.09 | N/A | 3 - 4 weeks | Keep cool 2-8°C | · Powder presentation, needs to be prepared manually. Follow manufacturers instructions |  | 0.01 | 3 | 3 | 5 | 4.5e-05 | 0 | 806.63 | USD | 0 | Sigma-Aldrich C3909-250MG |  |  | 29419000000
106584 | Kanamycin – lyophilized drug for laboratory use | BD | 6 | 6-vial pack, 830 µg / vial - to perform 240 tests | 7 months | 8 weeks | Keep cool 2-8°C | · MGIT tubes (106026) need to be ordered to perform liquid culture; · Item with very short shelf life (9 months after manufacturing) so about 4-5 months of use at arrival in country; · Ready to use vials. Need to order according to consumption |  | 0.16 | 9.8 | 6.6 | 6.6 | 0.00042083 | 0 | 90 | USD | 0 | BD 215348
106585 | Moxifloxacin - lyophilized drug for laboratory use | BD | 6 | 6-vial pack, 498 µg / vial - - to perform 240 tests | 7 months | 8 weeks | Keep cool 2-8°C | · MGIT tubes (106026) need to be ordered to perform liquid culture; · Item with very short shelf life (9 months after manufacturing) so about 4-5 months of use at arrival in country; · Ready to use vials. Need to order according to consumption |  | 0.16 | 9.8 | 6.6 | 6.6 | 0.00042083 | 0 | 90 | USD | 0 | BD 215349
106586 | Amikacin - lyophilized drug for laboratory use | BD | 6 | 6-vial pack, 332 µg / vial - to perform 240 tests | 7 months | 8 weeks | Keep cool 2-8°C | · MGIT tubes (106026) need to be ordered to perform liquid culture; · Item with very short shelf life (9 months after manufacturing) so about 4-5 months of use at arrival in country; · Ready to use vials. Need to order according to consumption |  | 0.16 | 9.8 | 6.6 | 6.6 | 0.00042083 | 0 | 90 | USD | 0 | BD 215350
106587 | Capreomycin - lyophilized drug for laboratory use | BD | 6 | 6-vial pack, 830 µg / vial - to perform 240 tests | 7 months | 8 weeks | Keep cool 2-8°C | · MGIT tubes (106026) need to be ordered to perform liquid culture; · Item with very short shelf life (9 months after manufacturing) so about 4-5 months of use at arrival in country; · Ready to use vials. Need to order according to consumption |  | 0.16 | 9.8 | 6.6 | 6.6 | 0.00042083 | 0 | 90 | USD | 0 | BD 215351
106588 | Ofloxacin - lyophilized drug for laboratory use | BD | 6 | 6-vial pack, 664 µg / vial - to perform 240 tests | 7 months | 8 weeks | Keep cool 2-8°C | · MGIT tubes (106026) need to be ordered to perform liquid culture; · Item with very short shelf life (9 months after manufacturing) so about 4-5 months of use at arrival in country; · Ready to use vials. Need to order according to consumption |  | 0.16 | 9.8 | 6.6 | 6.6 | 0.00042083 | 0 | 90 | USD | 0 | BD 215352
10. Solid and liquid culture: Rapid tests for MTB identification
106414 | Rapid test for detection of MPT 64 Antigen - SD Bioline TB Ag MPT64 | SD | 25 | SD BIOLINE TB Ag MPT64 Rapid is a rapid immunochromatographic identification test for the M. tuberculosis complex. M. tuberculosis has been known to secrete more than 33 different proteins. One of the predominant/Ant proteins, MPT64 was found in the culture fluid of only strains of the M. tuberculosis complex. | 17 months | 4 weeks | General Cargo | Product features: http://www.standardia.com/en/home/product/Rapid_Diagnostic_Test/TB_Ag_MPT64.html |  | 0.25 | 16.7 | 12.5 | 7 | 0.0014613 | 0 | 35 | USD | 0 | 08FK50 | CE
106611 | Capilia TB-Neo - Rapid test for detection of MPT 64 Antigen | TAUNS | 100 | To detect M. tuberculosis complex antigens in a suspension of bacteria cultured on the medium for acid-fast bacteria (AFB) or in a portion of the liquid medium for AFB in which bacteria have been cultured (to assist in the diagnosis of tuberculosis). | 26 months | 8 weeks | General Cargo | · Product linked to 106612; · Product features: http://capilia.jp/english/capilia_tb-neo.html |  | 0.79 | 17.4 | 130 | 134 | 3.03 | 0 | 100 | USD | 0 | CATB0870 | CE
106612 | Capilia TB-Neo Extraction buffer | TAUNS | 1 bottle of 20 ml | This product is an extraction buffer using Capilia TB-Neo to prepare a sample that is employed for the identification test of M. Tuberculosis when using a solid medium for acid-fast bacteria (AFB). | 29 months | 8 weeks | General Cargo | · Product linked to 106611; · Mandatory with Capilia TB-Neo - Rapid test for Detection of MPT 64 Antigen |  | 0.08 | 80 | 40 | 103 | 0.33 | 0 | 8 | USD | 0 | CATB0877 | CE
11. Solid and liquid culture: Chemicals for MTB identification
106292 | Tween® 80 | TTM | 1 kg | Tween® (Polyoxyethene) Ph Eur, NF, JP | 17 months | 3 weeks | General Cargo |  |  | 1.2 | 220 | 20 | 25 | 0.01 | 0 | 116.82 | USD | 0 | Merck 8170611000 | SDS/ CoA |  | 34021300
106293 | Hydrogen peroxide | TTM | 250 ml | Hydrogen peroxide (Perhydrol®) stabilized for higher temperature | 47 months | 3 weeks | Dangerous goods |  |  | 1 | 28 | 21 | 32 | 0.018816 | 0 | 32.46 | USD | 0 | Merck 1072100250 | SDS/ CoA | UN 2811 Class 5.1 Packing group III | 28470000
106295 | di-Sodium hydrogen phosphate | TTM | 1 kg | Di-Sodium hydrogen phosphate | 35 months | 3 weeks | General Cargo |  |  | 1.2 | 20 | 20 | 25 | 0.01 | 0 | 51.58 | USD | 0 | Merck 1065751000 | SDS/ CoA |  | 28352200
106297 | Sulphanilamide | TTM | 100 g | Sulphanilamide, AnalaR NORMPUR ≥ 99%, melting point > 164 °C, 100g plastic bottle | 35 months | 3 weeks | General Cargo |  |  | 0.2 | 10 | 10 | 15 | 0.0015 | 0 | 87.6 | USD | 0 | VWR Chemicals 21159.181 | SDS/ CoA |  | 29350090
106298 | N-(1-Naphthyl)-ethylendiamine dihydrochloride in glass bottle | TTM | 25 g | N-(1-Naphthyl)-ethylendiamine dihydrochloride, for analyses ACS, 25g glass bottle | 35 months | 3 weeks | General Cargo |  |  | 0.8 | 10 | 10 | 15 | 0.0015 | 0 | 143.89 | USD | 0 | VWR Chemicals 25792.130 | SDS/ CoA |  | 29215990
106301 | Cyanogen bromide | TTM | 50g | Cyanogen bromide for synthesis | 5 months | 3 weeks | Dangerous goods | Dangerous item that can ONLY be transported by road |  | 0.8 | 28 | 21 | 32 | 0.018816 | 0 | 94.12 | USD | 0 | Merck 820193.0050 | SDS/ CoA | UN 1889 Class 6.1 Packing Group - | 28530090
106302 | Analine / o-toluidine | TTM | 100ml | Analine/ o-toluidine for synthesis | 24 months | 3 weeks | Dangerous goods |  |  | 0.8 | 28 | 21 | 32 | 0.018816 | 0 | 30.33 | USD | 0 | Merck 8083120100 | SDS/ CoA | UN 1078 Class 6.1 Packing group II | 29214300
12. Molecular testing: LPA equipment and reagents
106512 | GenoType MTBDR plus V2 Test kit contains: 1) GenoType MTBDRplus Version 2.0 (30496A); 2) GenoLyse including service and support pack (51610). | HAIN | 96 | Molecular line probe assay for identification of M. tuberculosis complex and its resistance to rifampicin and isoniazid from smear-positive specimens and cultures. With Master mix | 12- 15 months | 4 weeks | Keep cool 2-8°C | WHO Guidance for 1st-line LPA: http://www.who.int/tb/publications/molecular-test-resistance/en/ |  | 1.5 | 28 | 15 | 13 | 0.0054 | 0 | 720 | EURO | 0 | 30496AM | CE |  | 38220000
106513 | GenoType MTBDRsl 2.0 (without Genolyse) | HAIN | 96 | Molecular line probe assay for identification of M. tuberculosis complex and its resistance to fluoroquinolones and aminoglycosides/cyclic peptides from rifampicin-resistant or MDR sputum specimens or cultures. With Master mix | 12- 15 months | 4 weeks | Keep cool 2-8°C | · Validated with GenoLyse only; · It is recommended to use MTBDRplus extracts; if not, Genolyse must be ordered separated; · WHO Guidance for 2nd-line LPA: http://www.who.int/tb/publications/lpa-mdr-diagnostics/en/ |  | 1.5 | 28 | 15 | 13 | 0.0054 | 0 | 720 | EURO | 0 | 31796A | CE |  | 38220000
106590 | GenoType Mycobacterium CMV2 (without GenoLyse) | HAIN | 96 | Molecular line probe assay for identification of clinically relevant mycobacterial species from culture | 12- 15 months | 4 weeks | Keep cool 2-8°C | Validated with GenoLyse only; Genolyse must be ordered separately |  | 1.5 | 28 | 15 | 13 | 0.0054 | 0 | 720 | EURO | 0 | 29996 | CE |  | 38220000
106591 | GenoType Mycobacterium AS (without GenoLyse) | HAIN | 96 | Molecular line probe assay for identification of nontuberculous mycobacteria from cultured material | 12- 15 months | 4 weeks | Keep cool 2-8°C | Validated with GenoLyse only; Genolyse must be ordered separately |  | 1.5 | 28 | 15 | 13 | 0.0054 | 0 | 720 | EURO | 0 | 29896 | CE |  | 38220000
106622 | GenoType MTBC | HAIN | 96 | Molecular line probe assay for differentiation of the mycobacterium tuberculosis complex from culture | 12- 15 months | 4 weeks | Keep cool 2-8°C | Product features: http://www.hain-lifescience.de/en/products/microbiology/mycobacteria/tuberculosis/genotype-mtbc.html |  | 1.5 | 28 | 15 | 13 | 0.0054 | 0 | 720 | EURO | 0 | 30196 | CE |  | 38220000
106254 | Molecular grade water | TTM | 10x1.7ml | DNAase, RNAase free water - not chemically treated molecular grade water - 1.7mlx10 or 10ml | N/A | 3 weeks | General Cargo |  |  | 0.01 | 11 | 9 | 4 | 0.000396 | 0 | 195.44 | USD | 0 | AppliChem A8510,1017 | SDS/ CoA |  | 28530010
106050 | Thermoshaker / Twincubator | HAIN | 1 | The Thermoshaker /TwinCubator is for use with reverse line blot strip assays and for manual hybridization and washing steps | N/A | 4 weeks | General Cargo |  |  | 9 | 50 / 54 | 44 / 37 | 40 / 28 | 0.04788 | 0 | 2200 | EURO | 0 | 11621 | CE |  | 84798997
106477 | GenoLyse (96 tests), Version 1.0 | HAIN | 96 | Kit for extraction of bacterial DNA | 12- 15 months | 4 weeks | Keep cool 2-8°C | · If the lab performs only Genotype MTBDRsl, need to add Genolyse in the order · If Genotype MTBDRsl is performed after GenoType MTBDRplus, there is no need to order Genolyse. |  | 1 | 13 | 5 | 7 | 0.002 | 0 | 50 | EURO | 0 | 51610 | CE
106073 | Special marker pens for Hain strips | HAIN | 6 | Special marker pens for Hain strips (better than pencil, soft grade, which could be used) | N/A | 4 weeks | General Cargo | · This product cannot be provided alone without any HAIN kits. · If the lab needs only a marker, order item: 106355 marker pen | [6 empty] | 0 | 0 | EURO | 0
106064 | GT blot - 48 | HAIN | 1 | The GT-Blot 48 is for use with reverse line blot strip assays and provides full automation for the washing and hybridization steps | N/A | 4 weeks | General Cargo | For customers using the GT-Blot 48 only; approximately 1 Reagent set is suitable for 4 GT kits |  | 75 | 100 | 85 | 90 | 0.7476 | 0 | 14352 | EURO | 0 | 1003/1 | CE
106438 | GT-Blot 48 Tray for 96 strips (black) | HAIN | 1 | Tray for use in combination with the GT-Blot 48 | N/A | 4 weeks | General Cargo |  |  | 0.1 | 45 | 16 | 2 | 0.00702 | 0 | 2.9 | EURO | 0 | 1003/5 | CE
106439 | GT-Blot 48 reagent kits | HAIN | 1 | The Reagent Set for DNA•STRIP Kits is a reagent set for use with test kits based on the DNA•STRIP technology. The reagents can replace or replenish the solutions provided with DNA•STRIP kits. The procedures described in the instructions for use of the DNA•STRIP test kit will be unaffected. | N/A | 4 weeks | General Cargo |  |  | 1.5 | 25 | 15 | 13 | 0.00057 | 0 | 0 | EURO | 0 | R003 | CE
106066 | Genoscan reader (GS-001) + PC + Screen + software | HAIN | 1 | The GenoScan® system is used for analysis of line probe assays (GenoType, GenoQuick®) from Hain Lifescience. The GenoScan® software reads strips from digital images generated by the GenoScan® reader, analyses the strip patterns, and supports the user in evaluation. | N/A | 4 weeks | General Cargo |  |  | 50 | 70 | 80 | 30 | 0.3 | 0 | 12180 | EURO | 0 | GS-001 | CE
106613 | GTQ-Cycler 96 - 7024007 | HAIN | 1 | High-capacity thermal cycler assisting you processing with the GenoType and GenoQuick® test systems | N/A | 4 weeks | General Cargo | [8 empty] | 4595 | EURO | 0 | 7024007 | CE
106074 | Microfuge / minispin for 0.2-2ml tubes for short-rapid spin down of reagents | TTM | 1 | Mini Centrifuge Sprout: ideal for quick spin-downs of reaction-tubes. Easy to use function. Rotor 6000 rpm (2000 x g), circular for: 6 x 1.5 and 2.0 ml tubes, including adaptor for 6x0.4 and 0.5 ml tubes, 3x 0.4 ml and 3x 1.5 and 20 ml tubes. Rectangular for two 8 x 0.2ml PCR tube-strips or 16 individual 0.2 ml PCR tubes. | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 3.5 | 31 | 16 | 13 | 0.006448 | 0 | 134.43 | USD | 0 | Heathrow Scientific 6120301 | CE |  | 84211920
106048 | Microliter Centrifuge | TTM | 1 | Hettich Centrifuge type Micro 200: capacity: 24x 2.0ml, rpm 15.000min-1, voltage 220-240V, 50-60Hz, with rotor type 2424-A, angle rotor 24x02-2.0ml | N/A | 3 weeks | General Cargo |  |  | 18 | 49 | 41 | 44 | 0.088396 | 0 | 2026.94 | USD | 0 | HETTICH 2400, 2424-B, 2423 |  |  | 84211920
106045 | PCR-Workstation, UV light | TTM | 1 | PCR Workstation: glass sides, which are resistant to UV-cracking and can be cleaned with the medium of choice, equipped with 1 x 25 watt UV-lamp (254 nm), 1x 15 watt white light lamp and 1 x UV/Air cleaner (254 nm).Dimensions: 69x51x57 cm | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz; · Item to be upright transported, fragile |  | 70 | 87 | 70 | 72 | 0.43848 | 0 | 2208.36 | USD | 0 | Kleinfeld Labortechnik UVC/T-M-AR | CE |  | 90189084
106411 | UV Crosslinker | TTM | 1 | BLX-254: UC crosslinking of DANN and RNA after Northern, Southern slot or dot blots and colony or plaque lifts, Programmable Smart Microprocessor Controller Operation modes, optimal Crosslink mode automatically provides a present UV energy dosage of 120 mJ/cm², energy set mode (0-999.990 µJ/cm²), Time mode 80-9,900 seconds) , intensity mode shows steady state UV intensity output, with irradiance display resolution of ± 5µ W/cm² over the entire range | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 7.9 | 50 | 28 | 25 | 0.035 | 0 | 1907.62 | USD | 0 | LTF Labortechnik LTF00061 | CE |  | 85394900
106219 | Minishaker / Vortex | TTM | 1 | Genie 2 Vortexer is designed with a solid-state electronic control that provides adjustable speed from 40 to 3200 rpm. Units can be operated in touch or continuous mode using the three-way switch. | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 4.7 | 27 | 20 | 25 | 0.0135 | 0 | 256.82 | USD | 0 | LMS Consult VTX-3000L, VTX-3000MT, VTX-3000MP | CE |  | 84798200
106510 | Water bath 7 Litters | TTM | 1 | Waterbath, with microprocessor PID-temp. Controller with integrated autodiagnostic system with fault indicator, solid state switching unit, one Pt100 sensor class A in 4-wire-circuit, integrated digital timer from 1min. To 99,59 hours, continuous operation | N/A | 3 weeks | General Cargo |  |  | 14 | 57 | 45 | 42 | 0.10773 | 0 | 677.61 | USD | 0 | Memmert WNB7 incl. 1 x B02686 |  |  | 84198998
106131 | Water bath 22 Litters | TTM | 1 | Water bath WNB 22: microprocessor PID temperature controller with integrated autodiagnostic system with fault indicator, one Pt 100 sensor class a in 4 wire circuit, integrated digital timer from 1 min to 99.59 hours for ON continuous operation | N/A | 3 weeks | General Cargo |  |  | 21.5 | 67 | 53 | 45 | 0.15979 | 0 | 925.52 | USD | 0 | Memmert WNB22 inkl. 1 B02688 |  |  | 84198998
106216 | Floating rack | TTM | 5 | Floating rack 24 places for tubes 1.5-2ml, yellow, for ultrasonic bath, Set of 5 pcs | N/A | 3 weeks | General Cargo |  |  | 0.04 | 17 | 11 | 16 | 0.001122 | 0 | 11 | USD | 0 | Heathrow Scientific HEA2165A 212-9125 VWR | CE |  | 39239090
106257 | Universal oven 32 Litters | TTM | 1 | Memmert UNB200 is been faced out the replacement is UN 30: Ventilation and Control: natural convection, vent connection with restrictor flap, adaptive multifunctional digital microprocessor PID-temperature controller with TFT colour display, 1 Pt100 sensor | N/A | 3 weeks | General Cargo |  |  | 43.5 | 65 | 59 | 85 | 0.32598 | 0 | 1031.29 | USD | 0 | Memmert UN30 inkl. 1 E28884 |  |  | 84198998
106455 | Rack: Heat-resistant for Cryo vial 1.5ml - Modular Heating Blocks | TTM | 1 | Constructed from a solid anodised aluminium block - Single block WxDxH: 95x76x51 mm, Close contact of tubes to block walls allows for maximum heat retention, Each block has a thermometer well for measuring block temperature | N/A | 3 weeks | General Cargo | Product linked to 106350 |  | 0.5 | 100 | 80 | 6 | 0.048 | 0 | 88.55 | USD | 0 | VWR Collection 460-3212 VWR | CE |  | 84798200
106453 | Thermometer for Oven | TTM | 1 | Oven thermometer aluminium 70x75mm 50°~ +300°C 42g | N/A | 3 weeks | General Cargo |  |  | 0.1 | 10 | 9 | 5 | 0.00045 | 0 | 6.57 | USD | 0 | Möller-Therm 106668 | CE |  | 90251980
106046 | Thermocycler, 96 positions for 0.2 ml tubes | TTM | 1 | Thermocycler, 96 positions for 0.2 ml tubes; Static temperature uniformity • ±0.5ºC, 30 seconds after clock-start at 95ºC; Temperature accuracy • ±0.25ºC (range: 35–100ºC)Temperature calibration • Calibrated to standards traceable to the National Institute of Standards and Technology (NIST). Heated cover • Maintains constant temperature of 105ºC for oil-free operation. Ramp time reproducibility • Reaches thermal set points within ±5 seconds. | N/A | 3 weeks | General Cargo |  |  | 9 | 48 | 33 | 33 | 0.052272 | 0 | 7646.92 | USD | 0 | PEQLAB peqSTAR 96X |  |  | 90328900
106213 | UV-light illuminator - Ethidium bromide - 312 nm | TTM | 1 | Spectroline, UV - light illuminator, ETX-20M, 15 Watt 312nm, 20x20cm Filter 15W | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 15 |  |  |  | 0.044 | 0 | 1844.47 | USD | 0 | Biotec-Fischer 151412-105 | CE |  | 90182000
106063 | Flexible thermometer for thermocycler and Twincubator validation: (e.g. GTH 1170 with GTF300) | TTM | 1 | Greisinger GTH 1170 Precision quick response thermometer: specs; measuring range -65 to + 199.9°C, resolution 0.1°C or 1 °C, accuracy -65 to +199.9°C ±0.5°C ; digital offset and scale adjustment for optimum precision, 9V battery type IEC 6F22 included. | N/A | 3 weeks | General Cargo |  |  | 0.2 | 13 | 9 | 4 | 0.000468 | 0 | 116.29 | USD | 0 | Greisinger 600000 + 600072 |  |  | 90251920
106224 | Refrigerated bench top centrifuge | TTM | 1 | Capacity 4x 290ml, rpm 15000 min-1/ RCF 24.400, DESIGN - Metal housing - Stainless steel centrifuging chamber - Viewing port in the lid - Powered lid locking - Ergonomic control and information panel - Easy rotor changing - For 50 ml tubes | N/A | 3 weeks | General Cargo | · 2 years warranty; · This item needs to be ordered with: 1. 106225 2. 106226 3. 106227 4. 106228 5. 106229 |  | 107 | 92 | 62 | 63 | 0.35935 | 0 | 5794.32 | USD | 0 | Hettich 1706 |  |  | 84211920
106225 | 4-place swing out rotor | TTM | 1 | 4-place swing out 90˚rotor for bench top centrifuge capacity 4x290ml | N/A | 3 weeks | General Cargo | Product linked to 106224 |  | 3 | 26 | 21 | 11 | 0.006006 | 0 | 922.72 | USD | 0 | Hettich 1754 |  |  | 84219100
106226 | Rotor buckets | TTM | set of 4 | 1752 Rotor bucket set of 4 pcs | N/A | 3 weeks | General Cargo | Product linked to 106224 |  | 2 | 29 | 21 | 10 | 0.00609 | 0 | 558.73 | USD | 0 | Hettich 1752 |  |  | 84219100
106227 | Aerosol tight lids for rotor buckets | TTM | set of 4 | 1751 Aerosol tight lids for rotor bucket, set of 4 pcs | N/A | 3 weeks | General Cargo | Product linked to 106224 |  | 0.1 | 22 | 22 | 26 | 0.002904 | 0 | 167.21 | USD | 0 | Hettich 1751 |  |  | 84219100
106228 | Reducing adapter for 50 ml tubes | TTM | set of 4 | 1772 Reducing adaptor size 13, 4 times (50ml) Falcon | N/A | 3 weeks | General Cargo | Product linked to 106224 |  | 0.4 | 20 | 17 | 9 | 0.00306 | 0 | 314.03 | USD | 0 | Hettich 1772-A |  |  | 84219100
106229 | Reducing adapter for 15 ml/12 ml tubes | TTM | set of 4 | 1771 Reducing adapter for 15 ml/12 ml tubes Falcon, set of 4 pcs | N/A | 3 weeks | General Cargo | Product linked to 106224 |  | 2 | 10 | 10 | 5 | 0.0005 | 0 | 314.03 | USD | 0 | Hettich 1771-A |  |  | 84219100
106211 | Electrophoresis chamber | TTM | 1 | Electrophoresis, horizontal mini gel system for separation of nucleic acids fragments PHERO-sub 0710-E complete system incl. 1x gel tray 70x100mm, UV transparent, 1x casting dams (2pcs.), 2x comb, 8 slots, 1mm, loading guides, platform, Cable set | N/A | 3 weeks | General Cargo | · This item needs to be ordered with the power supply (106212); · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 1.5 | 30 | 20 | 15 | 0.009 | 0 | 368.77 | USD | 0 | Biotec-Fischer 111100-002 | CE |  | 90272000
106212 | Power supply for electrophoresis | TTM | 1 | PHEROstab EV245, 400V, 500mA, 50W (230VAC)/ 131200-001 | N/A | 3 weeks | General Cargo | · Product linked to 106211; · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 7 | 30 | 25 | 20 | 0.015 | 0 | 647.93 | USD | 0 | Biotec-Fischer 131200-001 | CE |  | 90272000
106425 | Thermoblock / Thermo heater | TTM | 1 | Thermo heater, exchangeable heating blocks for different standard reaction tubes, range + 5 ºC above ambient to >100 ºC, heating time 25 ºC to 100 ºC ≤ 20 min., uniformity within the block +/- 0,2°C (at 37°C) | N/A | 3 weeks | General Cargo |  |  | 0.8 | 25 | 26.5 | 8.3 | 0.0054988 | 0 | 1052.48 | USD | 0 | Peqlab 91-D1100, 91-D120B, 91-D115B |  |  | 84199085
106443 | UPS - 700 VA - 500 W | TTM | 1 | On-Line UPS Vanguard VGD-700: UPS capacity 700 VA 230V sine wave output, 50/60 Hz outlets 4x Iec 320, surge protection RJ11/RJ45, dimensions W152xd420xH237 | N/A | 3 weeks | General Cargo special provision A67 | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 15.3 | 55 | 28 | 34 | 0.05236 | 0 | 467.03 | USD | 0 | PCM VGD-700 | CE | General Cargo Valve Regulated Lead Acid Battery UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 85359000
106492 | Battery Pack for UPS 700VA | TTM | 1 | Battery pack for UPS 3000 VA, to amplify the existing back up time of the UPS up to 150 min, auto select 50/60Hz 220/230/240v +/- 2 % Sinus, with UPS software | N/A | 3 weeks | General Cargo special provision A67 | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 14 | 24 | 16 | 12 | 0.016128 | 0 | 393.03 | USD | 0 | Multimatic BEM 1/4 (S) | CE | General Cargo Valve Regulated Lead Acid Battery UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 85359000
106081 | UPS - 3000 VA - 2100W | TTM | 1 | On-Line UPS Vanguard VGD-3000: UPS capacity 3000 VA 230V sine wave output, 50/60 Hz outlets 4x Iec 320, surge protection RJ11/RJ45, 20-30 min buffer, dimensions 35x22x40 cm | N/A | 3 weeks | General Cargo special provision A67 | · Product linked to 106013, 106596; · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 36.5 | 52 | 35 | 47 | 0.08554 | 0 | 1318.88 | USD | 0 | PCM VGD-3000 | CE | General Cargo Valve Regulated Lead Acid Battery UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 85359000
106491 | Battery Pack for UPS 3000 VA | TTM | 1 | Battery pack for UPS 3000 VA, to amplify the existing back up time of the UPS up to 150 min, auto select 50/60Hz 220/230/240v +/- 2 % Sinus, with UPS software | N/A | 3 weeks | General Cargo special provision A67 | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 36.5 | 52 | 35 | 47 | 0.08554 | 0 | 907.76 | USD | 0 | Multimatic BEM 2/12 (S) | CE | General Cargo Valve Regulated Lead Acid Battery UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 85359000
106595 | Voltage Stabilizer 2000W | TTM | 1 | Voltage Stabilizer 2000W | N/A | 3 - 4 weeks | General Cargo | · Product linked to 106592; · 2 years warranty; · Voltage / Amp / Frequency: min. 171 - 272 V, 8 Amp |  | 6.5 | 27.5 | 38.5 | 16 | 0.01694 | 0 | 286 | USD | 0 | Sollatek SVS08-22 | CE |  | 85363010
13. Molecular testing: Consumables
106053 | 1-channel pipette, variable from 1 - 10 µl - LPA | TTM | 1 | Gilson NEO P10N, 1-channel pipette, variable from 1-10µl, accuracy< 0.10 to 0.20, precision < 0.03 to 0.06 | N/A | 3 weeks | General Cargo | Product linked to 106554 |  | 0.2 | 32 | 11 | 4 | 0.001408 | 0 | 245.03 | USD | 0 | Gilson F144562 | CE |  | 84798997
106054 | 1-channel pipette, variable from 2,0 - 20 µl - LPA | TTM | 1 | Gilson NEO P20N, 1-channel pipette, variable from 2.0-20µl, accuracy< 0.10 to 0.20, precision < 0.03 to 0.06 | N/A | 3 weeks | General Cargo | Product linked to 106555 |  | 0.2 | 32 | 11 | 4 | 0.001408 | 0 | 226.94 | USD | 0 | Gilson F144563 | CE |  | 84798997
106055 | 1-channel pipette, variable from 10 - 100 µl - LPA | TTM | 1 | Gilson pipetman NEO P100 N (10-100 μl) ISO 8655, Chemical and UV resistant, resistant to disinfectant, autoclavable | N/A | 3 weeks | General Cargo | Product linked to 106388 |  | 0.2 | 32 | 11 | 4 | 0.001408 | 0 | 226.94 | USD | 0 | Gilson F144564 | CE |  | 84798997
106056 | 1-chanel pipette, variable from 20 -200 ul - LPA | TTM | 1 | Gilson pipetman NEO P200 N (20-200 μl) ISO 8655, Chemical and UV resistant, resistant to disinfectant, autoclavable | N/A | 3 weeks | General Cargo | Product linked to 106389 |  | 0.2 | 32 | 11 | 4 | 0.001408 | 0 | 226.94 | USD | 0 | Gilson F144565 | CE |  | 84798997
106057 | 1-channel pipette, variable from 200 - 1000 µl - LPA | TTM | 1 | Gilson pipetman NEO P1000 N (200-1000 μl) ISO 8655, Chemical and UV resistant, resistant to disinfectant, autoclavable | N/A | 3 weeks | General Cargo | Product linked to 106390 |  | 0.2 | 32 | 11 | 4 | 0.001408 | 0 | 226.94 | USD | 0 | Gilson F144566 | CE |  | 84798997
106058 | Racks for 1 channel pipette - 6 positions | TTM | 1 | Pipette rack, 6 positions, clear acrilic stand for 6x pipettes | N/A | 3 weeks | General Cargo |  |  | 0.8 | 36 | 21 | 23 | 0.017388 | 0 | 36.49 | USD | 0 | VWR Collection 631-0725 | CE |  | 39269097
106456 | Racks for 1 channel pipette - 3 positions | TTM | 1 | Stand for 3 micro pipettes multi function | N/A | 3 weeks | General Cargo |  |  | 0.4 | 26 | 21 | 18 | 0.009828 | 0 | 27.63 | USD | 0 | VWR Collection 631-0723 VWR | CE |  | 39269097
106270 | Racks for 1 channel pipette - 4 positions | TTM | 1 | Pipettor stand, Ergonomical: slightly curved to the user, universal fitting for all pipettors, material: easy to clean, 3 mm thick acryl, outer dimensions: 20x12 x10 cm (wxhxd) | N/A | 3 weeks | General Cargo |  |  | 0.5 | 26 | 21 | 23 | 0.012558 | 0 | 30.83 | USD | 0 | VWR Collection 631-0724 VWR | CE |  | 39269097
106554 | Sterile, DNA-RNAse-free tips, 0.1 - 10 µl | TTM | 960 | Pipette filter tips PP, capacity 0.1-10µl, free from DNA/RNAse and ATP, autoclavable, sterile, 10 racks of 96 pcs | 35 months | 3 weeks | General Cargo | Product linked to 106053 |  | 3 | 30 | 18 | 24 | 0.01296 | 0 | 129.72 | USD | 0 | Neptune BT10XL | CE |  | 39269097
106555 | Sterile, DNA-/RNAse-free tips, 1.0 - 20 µl | TTM | 960 | Pipette filter tipsPP, capacity 0.2-20µl, free from DNA/RNAse and ATP, autoclavable, sterile, 10 racks of 96 pcs | 35 months | 3 weeks | General Cargo | Product linked to 106054 |  | 3 | 30 | 18 | 24 | 0.01296 | 0 | 129.72 | USD | 0 | Neptune BT20 | CE |  | 39269097
106388 | Sterile, DNA-/RNAse-free tips, 10 - 100 µl | TTM | 960 | Pipette filter tipsPP, capacity 5-100µl, free from DNA/RNAse and ATP, length 89mm, autoclavable, sterile, 10 racks of 96 pcs | 35 months | 3 weeks | General Cargo | Product linked to 106055 |  | 1.5 | 33 | 20 | 13.5 | 0.00891 | 0 | 72.75 | USD | 0 | AHN Biotechnologie 2-112-96-0 | CE |  | 39269097
106389 | Sterile, DNA-/RNAse-free tips, 20 - 200 µl | TTM | 960 | Pipette filter tipsPP, capacity 1-200µl, free from DNA/RNAse and ATP, autoclavable, sterile, 10 racks of 96 pcs | 35 months | 3 weeks | General Cargo | Product linked to 106056 |  | 1.5 | 33 | 20 | 13.5 | 0.00891 | 0 | 53.59 | USD | 0 | AHN Biotechnologie 2-113-96-0 | CE |  | 39269097
106390 | Sterile, DNA-/RNAse-free tips, 100 - 1000 µl | TTM | 960 | Pipette filter tipsPP, capacity 100-1000µl, free from DNA/RNAse and ATP, length 89mm, autoclavable, sterile, 10 racks of 96 pcs | 35 months | 3 weeks | General Cargo | Product linked to 106057 |  | 2.2 | 49 | 26 | 11 | 0.014014 | 0 | 58.97 | USD | 0 | AHN Biotechnologie 2-201-96-0 | CE |  | 39269097
106061 | Hand-dispenser pipette, 5 dosage steps, for combitips - LPA | TTM | 1 | Handy Step Brand, repetitive pipette is contoured to fit your hand comfortably. All controls are designed for intuitive use with one hand. With wall bracket. | N/A | 3 weeks | General Cargo |  |  | 0.3 | 24 | 17 | 6 | 0.002448 | 0 | 211.48 | USD | 0 | Brand Handystep 612-3858 VWR |  |  | 84798997
106395 | Tips for Multipette/Combi 5 ml - bag of 10 | TTM | 100 | Dispenser tips classic, 5ml, Compatible with: Ritter® pipette® and pipette® pro, Eppendorf® Multipette® 4780, Brand® HandyStep® and HandyStep®electronic, Minilab 100/101, Easy Step, Distriman® | 24 months | 3 weeks | General Cargo | Product linked to 106115 |  | 0.8 | 28 | 22 | 16 | 0.009856 | 0 | 62.7 | USD | 0 | VWR Collection 613-1012 VWR | CE |  | 39269097
106396 | Tips for Multipette/Combi 5 ml, individually packed | TTM | 100 | Dispenser tips classic, 5ml, individually packed Compatible with: Ritter® pipette® and pipette® pro, Eppendorf® Multipette® 4780, Brand® HandyStep® and HandyStep®electronic, Minilab 100/101, Easy Step, Distriman® | 24 months | 3 weeks | General Cargo | Product linked to 106115 |  | 0.8 | 28 | 22 | 15 | 0.00924 | 0 | 64.34 | USD | 0 | VWR Collection 613-1012 VWR | CE |  | 39269097
106397 | Tips for Multipette/Combi 12.5ml- bag of 10 | TTM | 100 | Dispenser tips classic, 12.5ml Compatible with: Ritter® pipette® and pipette® pro, Eppendorf® Multipette® 4780, Brand® HandyStep® and HandyStep®electronic, Minilab 100/101, Easy Step, Distriman® | 24 months | 3 weeks | General Cargo | Product linked to 106115 |  | 0.9 | 29 | 19 | 13 | 0.007163 | 0 | 72.83 | USD | 0 | VWR Collection 613-1013 VWR | CE |  | 39269097
106398 | Tips for Multipette/Combi 12.5 ml - individually packed | TTM | 100 | Dispenser tips classic, 12.5ml individually packed Compatible with: Ritter® pipette® and pipette® pro, Eppendorf® Multipette® 4780, Brand® HandyStep® and HandyStep®electronic, Minilab 100/101, Easy Step, Distriman® | 24 months | 3 weeks | General Cargo | Product linked to 106115 |  | 1 | 28 | 22 | 15 | 0.009405 | 0 | 72.83 | USD | 0 | VWR Collection 613-1013 VWR | CE |  | 39269097
106412 | Long 1 ml tips with filter | TTM | 800 | Molecular Bio Products *ART* Barrier tips, with Micro Point Cat. No. 2079 100-1000µl packaging 8 x 100 sterile | 24 months | 3 weeks | General Cargo | Product linked to 106075 or 106057 |  | 2.2 | 28 | 23 | 27 | 0.017388 | 0 | 97.03 | USD | 0 | Thermo Scientific ART ART 1000 µl | CE |  | 39269097
106068 | Stand (96) for 0.2 ml pcr tubes / PCR Rack | TTM | set of 10 | Set of Rotilabo®-PCR-stands, Made of SBS. With 96 holes in 8 x 12 array, for vials 0.2 ml, individual vials, on strips or plates, in different colours 4x blue, 4x red, 2x white Dimensions L 127 x W 85 x H 16 mm. | N/A | 3 weeks | General Cargo | Product linked to 106385 |  | 0.2 | 13 | 9 | 10 | 0.00117 | 0 | 24.58 | USD | 0 | Ratiolab 6114102, 6114111 | CE |  | 39269097
106344 | Cryo-vial, sterile with cap, 2 ml | TTM | 1000 | Cryo-vial, sterile with cap, 2 ml Material: PP screw cap same material, dimensions: 12.5x49 mm, with stand | 35 months | 3 weeks | General Cargo |  |  | 2.5 | 36 | 21 | 23 | 0.017388 | 0 | 332.15 | USD | 0 | Brand 114832 479-1203 VWR | CE |  | 39269097
106350 | Cryo-vial, sterile with cap, 1.5 ml | TTM | 1000 | Sigma Cryo-vial, sterile with cap, 1.5ml, V4131, packaging in box of 2 x 500 pcs. | N/A | 3 weeks | General Cargo |  |  | 1 | 23 | 21 | 15 | 0.007245 | 0 | 251.96 | USD | 0 | Kleinfeld Labortechnik R334-5SPR | CE |  | 39269097
106052 | Desktop Deep freeze storage box with lid for 1.5 ml cryovials - Cardboard cryobox | TTM | 1 | Cardboard cryobox, 50mm high, white, 136x136mm, with 9x9 inlet | N/A | 3 weeks | General Cargo | Product linked to 106350 |  | 0.1 | 1.6 | 13.6 | 5 | 0.0001088 | 0 | 2.46 | USD | 0 | VWR Collection 479-1422, 479-0714 | CE |  | 48196000
106495 | Desktop Deep freeze storage box with lid for 1.5 ml cryovials - Plastic | TTM | 1 | Simport Cryo-storage-box with lid plastic, suitable for cryovial 1ml to 2ml, for storage of up to 100 pcs. Operating range -196 °C to +121 °C, autoclavable, with numeric position label, colour (blue, red or green) | N/A | 3 weeks | General Cargo |  |  | 0.2 | 15 | 15 | 56 | 0.0126 | 0 | 4.19 | USD | 0 | VWR Collection 479-0487 VWR | CE |  | 39239090
106446 | Reagent Quick Chiller | TTM | 1 | Quick Chill™ unit, polycarbonate; non-toxic insulating solution, with handle. For use in molecular biology applications such as ethanol precipitation of DNA/RNA samples. For repeatedly use for up to 45 minutes, holds 12 x 1.5 or 0.5 ml microcentrifuge tubes (16 inserts for 0.5 ml included). Stackable design. | N/A | 3 weeks | General Cargo |  |  | 1.3 | 12 | 16 | 14 | 0.002688 | 0 | 109.44 | USD | 0 | Nalgene 2600646 | CE |  | 39269097
106217 | Racks with lid (plastic preferred), suitable for cryovial | TTM | 1 | Simport Cryo-storage-box with lid plastic, suitable for cryovial 1ml to 2ml, for storage of up to 81 pcs. Operating range -196 °C to +121 °C, autoclavable, with numeric position label, colour (blue, red or green) | N/A | 3 weeks | General Cargo | Product linked to 106344 |  | 0.1 | 13 | 13 | 6 | 0.001014 | 0 | 5.5 | USD | 0 | Kleinfeld Labortechnik T314-281B | CE |  | 39239090
106383 | Cryo-vial, sterile with cap, for one hand operation | TTM | 1000 | Cryovial PP Y sterilized external Th (SAL10-6), with lid, 2ml, temperature up to -196°C, free of RNAsen, DNAsen, DNA and Endotoxine, packaging 1000pcs | 35 months | 3 weeks | General Cargo |  |  | 2.7 | 30 | 21 | 25 | 0.01575 | 0 | 255.14 | USD | 0 | Kleinfeld Labortechnik 1 x T310-2A, 2 x T312-7 | CE |  | 39269097
106410 | Racks for cryovials for One hand operation | TTM | 1 | Rack for cryovials made of white PC, the snap in attachment in the base stops the vias from turning when opening and closing the box and also enables speedy and convenient one-hand operation. With alpha numerical coding. Dimensions: L197xW102xH28mm 50 holes in a 5x 10 array. autoclavable | N/A | 3 weeks | General Cargo | Product linked to 106383 |  | 0.1 | 22 | 11 | 3 | 0.000726 | 0 | 12.88 | USD | 0 | Kleinfeld Labortechnik T315-2 | CE |  | 39269097
106357 | Cryo-tags | TTM | 1 | Labels white, for 2ml tubes, packed in rolls of 1000 pcs | N/A | 3 weeks | General Cargo |  |  | 0.05 | 9.5 | 9.5 | 4 | 0.000361 | 0 | 26.31 | USD | 0 | Labelsolutions UK Limited LSFTT-SG-82MR-P, 32 x 10 mm | SDS |  | 48219010
106385 | PCR tubes | TTM | 1000 | Single PCR Tubes with Attached Caps, 0.2 ml and 0.5 ml PCR tubes with attached caps are compatible with leading thermal cyclers with heated lids, autoclavable, packaging 1000 pcs | 35 months | 3 weeks | General Cargo |  |  | 0.2 | 25 | 14 | 11 | 0.00385 | 0 | 46.16 | USD | 0 | BRAND 7813 05 | CE |  | 39269097
106051 | Desktop deep freeze storage box with lid for 0.2ml PCR tubes | TTM | 1 | High resistance PP. Designed to accommodate 96 individual 0.2 ml tubes or 12 or 8 strips of 8- or 12- tubes, Tube wells are easily identified by imprinted numbers and letters, Removable, hinged lid, in different colours | N/A | 3 weeks | General Cargo | Product linked to 106385 |  | 0.07 | 13.5 | 10 | 3.5 | 0.0004725 | 0 | 3.1 | USD | 0 | Heathrow HEA23461A 211-3505 VWR
106421 | Standard reaction tubes 1.5 ml | TTM | 500 | Microcentrifuge Tubes 1.5ml with attached screw caps, PP, graduated, temp. -196°C, +121°C autoclavable, PP screw cap with silicone seal, with frosted marking field, can withstand centrifugation to 20000g | 35 months | 3 weeks | General Cargo |  |  | 1 | 23 | 21 | 15 | 0.007245 | 0 | 108.12 | USD | 0 | Kleinfeld Labortechnik T334-5SPR | CE |  | 39269097
106422 | Standard reaction tubes 2 ml | TTM | 500 | Microcentrifuge Tubes 2 ml with attached screw caps, PP, graduated, -196°C, +121°C autoclavable, PP screw cap with silicone seal, with frosted marking, can withstand centrifugation to 20000g | 35 months | 3 weeks | General Cargo |  |  | 1 | 23 | 22 | 15 | 0.00759 | 0 | 109.44 | USD | 0 | Kleinfeld Labortechnik T334-7SPR | CE |  | 39269097
106246 | Racks, desinfectable / autoclavable for 1.5/2 ml tubes | TTM | 1 | Rack, PP autoclavable, for tubes 0,2 - 2,0 ml 20positions | N/A | 3 weeks | General Cargo |  |  | 0.02 | 21.5 | 9 | 5 | 0.0009675 | 0 | 3.61 | USD | 0 | Heathrow HEA4284A 211-3901 VWR | CE |  | 39239090
106384 | PP-tubes for centrifuge, non sterile, 15 ml | TTM | 500 | Centrifuge tube , 15ml, SUPERCLEAR™, material PP, with lid, packaging 500pcs | 35 months | 3 weeks | General Cargo |  |  | 5 | 40 | 40 | 50 | 0.08 | 0 | 85.34 | USD | 0 | Greiner Bio-One 188271 | CE |  | 39269097
106433 | Rack for 15 ml PP-tubes -universal test tube rack 17mm | TTM | 1 | Universal test tube rack, autoclavable temperature resistant from -15 to 130 °C suitable for Ø 17mm 250x105x70 mm, 60 pos. White, art. No. Marien 5610123 | N/A | 3 weeks | General Cargo | Product linked to 106342 |  | 1.2 | 25 | 12 | 17 | 0.0021 | 0 | 12.99 | USD | 0 | Kleinfeld Labortechnik 7460001 | CE |  | 39233010
106340 | PP-tubes for centrifuge non sterile 50 ml | TTM | 500 | Centrifuge tubes 50ml, super clear™, PP aor Ps, conical, printer graduation, extra thick walls to allow centrifugation up to 15000g, with extra large, solvent resistant, white labelling area and black graduations flat PE caps can close with quick 3/4 turn, allowing one handed operation, patented earth friendly® racks can be labelled with lab markers, 1 rack contains 25tubes | 35 months | 3 weeks | General Cargo |  |  | 8.05 | 60 | 40 | 37 | 0.0888 | 0 | 74.1 | USD | 0 | Greiner Bio-One 227261 | CE |  | 39269097
106413 | Rack for 50 ml PP-tubes | TTM | 1 | Rack, for centrifuge tubes, 3x6 positions, for 50ml centrifuge tubes or test tubes up to 30mm diameter | N/A | 3 weeks | General Cargo | Product linked to 106340; 106341 |  | 1.9 | 40 | 37 | 20 | 0.0296 | 0 | 13.54 | USD | 0 | Kleinfeld Labortechnik S600-30B | CE |  | 39269000
106454 | Rack 15 ml / 50 ml - 4-Way microtube rack | TTM | 1 | Rack four different usable sides, high chemical resistant, high temperature resistant and autoclavable | N/A | 3 weeks | General Cargo |  |  | 0.2 | 18 | 6 | 9 | 0.000972 | 0 | 8.55 | USD | 0 | VWR Collection 211-0203 VWR | CE |  | 39269097
106434 | Disposable Pasteur pipettes, graduated, non sterile, 155 mm, 3 ml | TTM | 500 | TTM Pasteur pipette unsterile length 155mm, non toxic and controlled, packaging 500 pcs | N/A | 3 weeks | General Cargo |  |  | 1 | 36 | 17 | 16 | 0.009792 | 0 | 9.34 | USD | 0 | Herenz 1131303 | CE |  | 39269097
106556 | Forceps, individually wrap, sterile | TTM | 50 | Megro, Pinzette sterile, white plastic, individual packaging, 145 mm, box of 50 pcs | 35 months | 3 weeks | General Cargo |  |  | 0.25 | 15 | 5 | 8 | 0.0006 | 0 | 7.56 | USD | 0 | Ratiomed 126300 Megro | CE |  | 39239000
106415 | Filter paper - sheets | TTM | 100 | Macherey-Nagel, Filter paper MN 616md, filtration time 55, 580x580mm, weight 85g/m², 100 pc | N/A | 3 weeks | General Cargo |  |  | 3.2 | 60 | 60 | 3 | 0.0108 | 0 | 118.4 | USD | 0 | Machery Nagel MN 616 md | CE |  | 48232000
106401 | Parafilm sealing film | TTM | 1 | Parafilm, M for use between -40 to +50°C, length 58m | N/A | 3 weeks | General Cargo |  |  | 0.7 | 13 | 11 | 14 | 0.002002 | 0 | 26.14 | USD | 0 | Brand 701605 291-1213 VWR | CE |  | 39191080
106361 | Parafilm with dispenser | TTM | 1 | Parafilm, M for use between -40 to +50°C, length 58m, with dispenser box | N/A | 3 weeks | General Cargo |  |  | 0.7 | 13 | 11 | 14 | 0.002002 | 0 | 26.14 | USD | 0 | Brand 701605 291-1213 VWR | CE |  | 39191080
14. Molecular testing: Chemicals for gel electrophoresis
106321 | Agarose | TTM | 100 g | Agarose SeaKem LE(Lonza) | 47 months | 3 weeks | General Cargo |  |  | 0.15 | 6 | 6 | 13.5 | 0.000486 | 0 | 246.67 | USD | 0 | Lonza 733-1535 VWR | SDS/ CoA |  | 38220000
106322 | Boric acid | TTM | 1 kg | Boric acid, for analyses EMSURE® ACS, ISO, Reag. Ph Eur > 99.5% | 47 months | 3 weeks | General Cargo |  |  | 1.2 | 20 | 20 | 25 | 0.01 | 0 | 52.36 | USD | 0 | Merck 1001651000 | SDS/ CoA |  | 28100090
106323 | Sodium acetate trihydrate | TTM | 1 kg | Sodium acetate trihydrate, analytical grade | 12 months | 3 weeks | General Cargo |  |  | 1.05 | 9 | 9 | 21.5 | 0.0017415 | 0 | 40.4 | USD | 0 | Merck 1062671000 | SDS/ CoA |  | 29152900
106324 | Ethidium bromide powder | TTM | 1 g | Ethidium bromide, for bio chemical usage | 24 months | 3 weeks | General Cargo |  |  | 0.2 | 10 | 10 | 12 | 0.0012 | 0 | 57.32 | USD | 0 | Merck Millipore 1.11615.0001 | SDS/ CoA |  | 29339980
106325 | TRIS | TTM | 1 kg | TRIS, > 99%, buffer grade ACS, Reag. Ph Eur 99% | 47 months | 3 weeks | General Cargo |  |  | 1.1 | 2 | 12 | 18 | 0.002592 | 0 | 216.41 | USD | 0 | Merck Millipore 1.08382.1000 | SDS/ CoA |  | 29221985
106326 | EDTA | TTM | 250 g | EDTA Bio chemical | 47 months | 3 weeks | General Cargo |  |  | 0.3 | 10 | 10 | 12 | 0.0012 | 0 | 23.22 | USD | 0 | AppliChem A1103.0250 | SDS/ CoA |  | 29224985
106327 | Gel load buffer | TTM | 1 | Carl Roth 6x Gel load buffer (with Ficoll), 5x1.8ml | 35 months | 3 weeks | Keep cool 2-8°C |  |  | 0.2 | 10 | 10 | 4 | 0.0004 | 0 | 38.48 | USD | 0 | Carl Roth X905.1 | SDS/ CoA |  | 29054500
106334 | DNA ladder | TTM | 1 | Roth DNA Ladder kombi, 100-10000 bp, 20μ + 50μ + gel buffer/ CL22.1. Control for separation of DNA fragments | 35 months | 3 weeks | Keep cool 2-8°C |  |  | 0.8 | 15 | 15 | 20 | 0.0045 | 0 | 98.17 | USD | 0 | Carl Roth CL22.1 | SDS/ CoA |  | 38220000
106336 | Ethidium bromide solution, | TTM | 5 ml | Ethidium bromide solution, 500 µg/mL in H2O | 24 months | 3 weeks | General Cargo |  |  | 0.5 | 20 | 20 | 10 | 0.004 | 0 | 20.62 | USD | 0 | Sigma-Aldrich E1385-5ML | SDS/ CoA |  | 38220000
15. Media kitchen: Equipment
106099 | Universal Hot-air oven - Memmert UFE 600 | TTM | 1 | Memmert UFE 600 has been phased out; the replacement is UF 260, forced air circulation by quiet air turbine adjustable in 10% steps, continuous adjustment of pre heated fresh air admixture, fuzzy-supported PID microprocessor, solid state switching unit, 1 P | N/A | 3 weeks | General Cargo | · This item needs to be ordered with item 106100; · 2 years warranty; · The transport of this item must be upright |  | 166 | 93 | 92 | 153 | 1.3091 | 0 | 3493.31 | USD | 0 | Memmert UF260 inkl. 2 E28891 |  |  | 84198998
106100 | Perforated stainless steel shelf for Universal Hot-air oven | TTM | 1 | Stainless steel Grid for Memmert UF 260 E3 | N/A | 3 weeks | General Cargo | Product linked to 106099 |  | 2 | 80 | 60 | 10 | 0.048 | 0 | 113.48 | USD | 0 | Memmert B29725 |  |  | 84198998
106101 | Laboratory washer-disinfector | TTM | 1 | Washer disinfector for laboratory glass, LAB 500 SCL (500D has been phased out), ISO 15883-1/3, KIWA, CSA-US, CEI EN 61010-1, EMC and SPRI, chamber volume 171lt, 5600W, pump 550W, Noise 52 dB(A), permitted temp. +5°C/ +40°C, programmable detergent pump, 5 washing programs, inbuilt water softener, wrasencondensator | N/A | 3 weeks | General Cargo | · This item needs to be ordered with: 1. 106102 2. 106103 3. 106104 4. 106105 · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 93 | 60 | 63 | 85 | 0.3213 | 0 | 5328.18 | USD | 0 | Steelco LAB 500 SC DL | CE/ ISO |  | 84222000
106102 | Washer-disinfector 1/2 insert | TTM | 1 | Insert with 28 spring hooks for laboratory glassware | N/A | 3 weeks | General Cargo | Product linked to 106101 |  | 8 | 55 | 52 | 56 | 0.16016 | 0 | 142.08 | USD | 0 | Steelco C61 | CE/ ISO |  | 84229090
106103 | 1/2 insert for baskets | TTM | 1 | Upper basket with washing arm | N/A | 3 weeks | General Cargo | Product linked to 106101 |  | 10 | 55 | 52 | 56 | 0.16016 | 0 | 233.84 | USD | 0 | Steelco C721 | CE/ ISO |  | 84229090
106104 | Upper basket/ injector for nozzles | TTM | 1 | Upper basket with 36 jets mm 80/110h diameter 6mm | N/A | 3 weeks | General Cargo | Product linked to 106101 |  | 12 | 55 | 52 | 56 | 0.16016 | 0 | 1117.93 | USD | 0 | Steelco C711 | CE/ ISO |  | 84229090
106105 | 1/4 insert, for test tubes 160x16 | TTM | 1 | Net basket 1/4mm 200h/7.87"h | N/A | 3 weeks | General Cargo | Product linked to 106101 |  | 10 | 55 | 52 | 56 | 0.16016 | 0 | 193.39 | USD | 0 | Steelco C 70 | CE/ ISO |  | 84229090
106134 | Sterile filling station of culture media 1000ml | TTM | 1 | Dropping funnel with pressure compensation, PTFE plug, made of borosilicate glass 3.3 with graduation, cylindrical, With pressure compensation tube, NS core, NS sleeve and cock plug with threaded safety device. Din 12567, ISO 4800 | N/A | 3 weeks | General Cargo |  |  | 1.5 | 15 | 10 | 35 | 0.00525 | 0 | 82.29 | USD | 0 | Lenz Laborglasinstrumente 4.3770.14 + 1.4000.29 | CE |  | 39269097
106135 | Sterile filling station of culture media 250ml | TTM | 1 | Dropping funnel with pressure compensation, PTFE plug, made of borosilicate glass 3.3 with graduation, cylindrical, With pressure compensation tube, NS core, NS sleeve and cock plug with threaded safety device. Din 12567, ISO 4800 | N/A | 3 weeks | General Cargo |  |  | 0.4 | 42 | 13 | 10 | 0.00546 | 0 | 56.7 | USD | 0 | Lenz Laborglas instrumente 4.3749.14 + 1.4000.29 | CE |  | 39269097
106136 | Automatic filling station | TTM | 1 | Automatic filling station, PD 5206, flow rate 1.3-3.900ml, flow rate accuracy ±2 %, rpm 24-600, speed setting and control digital, vol. dosing 0.1-9,999, interval dosing 0.1-9,999 in breaks 0.1sec-750h. | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 9 | 46 | 26 | 32 | 0.038272 | 0 | 3042.48 | USD | 0 | Heidolph PD 5206, SP quick 1,6 mm, 525-33000-00 | CE |  | 84138100
106138 | Inspissator for coagulation of egg-based culture media | TTM | 1 | Inspissator TBT 100-IN control unit, tank, quilt& blanket, range ambient +5°C to 90°C, with std constant level device fitted, to hold up to 156 test tubes (16x150mm) or 162 universal containers or glass bijous. Unit does not include electronic timer, racks are slanted but not adjustable, water level sensor: the grant inspissator has a constant level device and is fitted with a visible alarm. | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 4.3 | 104 | 60 | 38 | 0.23712 | 0 | 4931 | USD | 0 | Grant Instruments TBT-T100IN | CE |  | 84198998
106127 | Fully automatic water distil | TTM | 1 | Fully automatic water distil: water still for single distillation (fully automatic) 4l/h, with 16 ltr. storage tank, suitable for both bench and wall mounting, very good distillate quality; conductivity approx. 2.3 µs/ cm at 25°C, the distillate storage tank accepts twice the hourly capacity of still, electronic level switch switches the still off when storage tank is full and restarts it automatically when distillate is withdrawn, electronic impurity detector switches the unit off in case of high degree impurities of water, thermostat low water cut off, economical energy consumption by distillation of the heated cooling water, this unit is specially been used in research and development e.g. for preparing bacteriological and medical samples. | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 24 | 69 | 43 | 59 | 0.17505 | 0 | 1952.43 | USD | 0 | GFL 2004 | CE |  | 84194000
106137 | Stative / Stand | TTM | 1 | Stand plate: Steel, varnished: Firm, non-slip stand. For a wide variety of uses. With threaded hole on face side. With rubber feet for a firm footing , easy to clean, robust, rod stainless steel art. no. 241-7133, rust-proof, 12 mm diameter, very good chemical resistance. Three-pronged Clamp art. no. 241-7433: Zinc die-casting, powder-coated, the prongs of the clamps are covered with a plastic coating to protect the items clamped, without bosshead for fixing to stand, with wing screws for rapid easy fixing Bosshead art.no. 241-7179, clamping range up to 16 mm, angle: 90 + 180, for quick, easy assembly | N/A | 3 weeks | General Cargo |  |  | 4 | 32 | 32 | 20 | 0.02048 | 0 | 51.76 | USD | 0 | Usbeck + Juchheim 2005 + 2104 + 4710 + 2201 | CE |  | 39269097
106140 | Laboratory floor-standing refrigerator 140L | TTM | 1 | Liebherr/Vaust Laboratory refrigerator FKUv 1610/ electronic same unit like LKUv 1610 | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 40 | 82 | 60 | 61.5 | 0.30258 | 0 | 1295.86 | USD | 0 | Vaust FKUv 1610 | CE |  | 84182151
106424 | Laboratory freezer 140 L | TTM | 1 | Vaust GGU 1500/ex: same basic model as Liebherr but with Vaust electronic. | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz; · To be transported and store upright |  | 43 | 60 | 62 | 82 | 0.30504 | 0 | 1060.7 | USD | 0 | Liebherr / Vaust GGU 1500/ex | CE |  | 84182110
106142 | Deep freezer 70L / - 50 °C to - 85 °C | TTM | 1 | Chest Freezer: user friendly microprocessor technique, conformable control panel with digital display to enter and read out all security relevant data, adjustable alarm limits for over and under temperatures (1-20K), additional control and alarm option, maintains all display and alarm functions for 60 hours in case of power failure, electronic alarm device and potential free contact for connection to an internal alarm system, RS 232 serial interface for trouble free data transfer, inside cabinet made completely of stainless steel and ex proof, energy saving maintenance-free cooling unit, both not flammable refrigerant and the all around special insulation (up to 150 mm) are CFC free (heat emission is only approx. 400W), temperature range from -50°C to -85°C, capacity 5 lt, unit is included with the following accessories: 6x Storage rack for boxes 50mm height art. no. 6901, 48x storage boxes art. no. 6970 | N/A | 3 weeks | General Cargo Large/Dangerous | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 221 | 101 | 86 | 128 | 1.1118 | 0 | 11778.32 | USD | 0 | GFL 6380, 6901, 6970 | CE | Large/Dangerous goods UN 3161 Class 2 Packing group N/A | 84183020
106130 | Benchtop pH meter | TTM | 1 | Hanna Instruments bench pH meter, ph211: with large multifunction display, which indicates the pH- value or V-value and temperature, automatic calibration with 5 pre- programmed std buffers (pH 4.01; 6.86; 7.01; 9.18; 10.01), graphic symbols show the user how to calibrate and use the unit, temperature compensation automatic and manual, pH 0.00 to 14.00, mV ± 1999 (ORP), Range, ± 399.9 (ISE) , °C 0.0 to 100.0, pH 0.01, mV 1 (ORP) Resolution 0.1 (ISE), °C 0.1, pH ± 0.01 m'V ± 1 (ORP), Accuracy ± 0.2 (ISE) ° C ± 0.5 | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 1.1 | 26 | 20 | 10 | 0.0052 | 0 | 455.93 | USD | 0 | Hanna Instruments HI 2211-02, HI7001-0P, HI7000-6P, HI7000-9P | CE |  | 90278011
106132 | Magnetic stirrer with heating plate | TTM | 1 | Heating and magnetic stirrer RCT: with basic safety control, adjustable safety circuit of hotplate temperature (50-360 °C), direct connection of medium temperature sensor PT 1000 possible. Including protective hood and Pt 1000 sensor. Magnet stirrer: max stirring volume 20l, motor power input/ output 16/9W, rpm 50 to 1500 min. Heater function: heating power 600 w, digital, temperature range room temperature to 310 °C | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 3.7 | 46 | 18 | 14 | 0.011592 | 0 | 806.49 | USD | 0 | IKA 0003810000, 0003378000, 0001545100, 0002437700, 0003547700 | CE |  | 84798200
106557 | Magnetic bars set of 12 | TTM | 1 set of 12 | Magnetic followers, cylindric PTFE coated: set of 12 magnetic bares, 2 x 15 mm, 2 x 20 mm, 2 x 25 mm, 1 x 30 mm, 2 x 40 mm, 2x 50 mm, 1 x 60 mm length | N/A | 3 weeks | General Cargo |  |  | 0.8 | 10 | 10 | 4 | 0.0004 | 0 | 39.3 | USD | 0 | Cowie 001.3009 | CE |  | 39269097
106426 | Blender 2x Speed | TTM | 1 | Waring Labormixer Reliable devices with powerful motors. For thorough mixing, dispersion and emulsification. Special cutter heads and high speed operation make the laboratory mixer suitable for many applications. Basic devices with capacity of 1 l , heat resistant glass beakers, Supplied with motor, cutter head, beaker and lid, 18000-22000rpm | N/A | 3 weeks | General Cargo | · This item needs to be ordered with item 106427; · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 5 | 25 | 25 | 47 | 0.029375 | 0 | 355.21 | USD | 0 | Wamhoff 40110 | CE |  | 84798200
106427 | Glass receptable for blender 1l | TTM | 1 | Waring glass beaker 1l | N/A | 3 weeks | General Cargo | Product linked to 106426 |  | 1 | 20 | 20 | 30 | 0.012 | 0 | 95.9 | USD | 0 | Wamhoff 40211 | CE |  | 84389000
106128 | Electronic analytical balance for DST | TTM | 1 | Analytical Balance ABT 120-5DNM, with single-cell weighing system, weighing range: 42 – 120 g, readout: 0,01 – 0,1 mg, reproducibility: 0,02 – 0,1 mg, linearity: 0,05 – 0,2 mg, incl. thermal printer with data interface RS-232 | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 7.6 | 22 | 35 | 34 | 0.02618 | 0 | 2256.67 | USD | 0 | Kern ABT 120-5DNM | CE |  | 90160010
106129 | Electronic precision balance | TTM | 1 | Kern PCB 3500-2 precision balance, measuring range max 3.5 kg, readout 10mg, reproducibility 10mg, linearity ± 30 mg, pre tare function of known containers, data transfer for example to Printer or PC, battery option, auto of function | N/A | 3 weeks | General Cargo | · Essential for the preparation of ZN reagents; · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 1.4 | 24 | 36 | 18 | 0.015552 | 0 | 296.01 | USD | 0 | Kern PCB 3500-2 | CE |  | 90160010
106592 | Deep Freezer 550L / - 40 °C to - 86 °C | TTM | 1 | Deep Freezer 550L / - 40 °C to - 86 °C | N/A | 3 - 4 weeks | General Cargo Large/Dangerous | · This item needs to be ordered with: 1. 106593 2. 106594 3. 106595 · 2 years warranty; · Voltage / Amp / Frequency: 230 V, 50/60 Hz |  | 215 | 107 | 102 | 231 | 2.5211 | 0 | 11291.28 | USD | 0 | Arctiko ULUF 550 | EN ISO 13485:2012; ISO 9001:2008 | Large/Dangerous goods UN 3161 Class 2 Packing group N/A | 84184080
106593 | Stainless steel rack for Deep Freezer | TTM | 1 | Stainless steel rack for Deep Freezer | N/A | 3 - 4 weeks | General Cargo | Product linked to 106592 |  | 4.2 | 22.1 | 13.9 | 55.8 | 0.017141 | 0 | 190.48 | USD | 0 | Arctiko UDR50-44 | EN ISO 13485:2012; ISO 9001:2008 |  | 84189990
106594 | Cryo Gloves | TTM | 1 | Cryo Gloves | N/A | 3 - 4 weeks | General Cargo | Product linked to 106592 |  | 0.198 | 48.5 | 6 | 15 | 0.004365 | 0 | 159.39 | USD | 0 | Tempshield Cryo-Gloves | CE |  | 39262000
106595 | Voltage Stabilizer 2000W | TTM | 1 | Voltage Stabilizer 2000W | N/A | 3 - 4 weeks | General Cargo | · Product linked to 106592; · 2 years warranty; · Voltage / Amp / Frequency: min. 171 - 272 V, 8 Amp |  | 6.5 | 27.5 | 38.5 | 16 | 0.01694 | 0 | 286 | USD | 0 | Sollatek SVS08-22 | CE |  | 85363010
106107 | Separate water supply for deionised water + conductivity measure | TTM | 1 | Separate water supply unit wall mounted, stainless steel, for deionised water, pressure up to 10 bar, 300l/min, quality 0.1-20µS/cm, water temp. Max 30°C, with digital conductivity measure, measuring range 0-199.9 µS/cm, including 1.5m R3/4" hoses and mounting kit, with digital display | N/A | 3 weeks | General Cargo | This item needs to be ordered with: 1. 106124 2. 106125 3. 106126 |  | 24 | 25 | 25 | 70 | 0.04375 | 0 | 1340.26 | USD | 0 | Orben 18140022.1, 1814.235, 127442CAG, 1814.066 | CE |  | 90251920
106124 | Ion exchanger cartridge | TTM | 1 | Ion Exchange Cartridge, particle system, connected directly to water tab, pressure resistant up to 10bar, DIN285 for laboratory glassware washer, V&A stainless steel can be cleaned and sanitized, unlimited regeneration of the mixed bed resin, environmentally friendly, Flow rate 300/h, capacity 2100L, water quality 0.1-20µS/cm, including RDS quick connection 3/4" | N/A | 3 weeks | General Cargo | Product linked to 106107 |  | 15 | 45 | 45 | 24 | 0.0486 | 0 | 616.68 | USD | 0 | Orben 18140022.1, 127442CAG | CE |  | 84212100
106125 | Conductivity measuring device | TTM | 1 | Conductivity meter LKM Type a 1.0 with quick connections, conductivity in µS/cm and temp in °C, conductivity measuring range: 0.055-199.9 µS/cm, temp range 0.1-99.9 °C, automatic temp compensation, 230V/ 50Hz | N/A | 3 weeks | General Cargo | Product linked to 106107 |  | 0.5 | 17 | 14 | 8 | 0.001904 | 0 | 254.89 | USD | 0 | Orben 1202593-ANA | CE |  | 84219900
106126 | Ion exchanger resin for laboratory grade water | TTM | 1 | Regenerative resin 1l | N/A | 3 weeks | General Cargo | Product linked to 106107 |  | 1.3 | 40 | 50 | 60 | 0.12 | 0 | 17.76 | USD | 0 | Orben 891000 | CE |  | 39140000
16. Media kitchen: Consumables (plastic and glassware)
106143 | Balloon HDPE 30 L | TTM | 1 | Bürkle, Canister wide mouth with screw cap, capacity 31Lt. HDPE, transparent. Extra wide filling opening also makes cleaning inside much easier, which is important particularly when storing and transporting sensitive substances. UV-stabilised. Good chemical resistance. Heat resistant up to 110 °C. With threaded connector on the side just above the canister bottom, and drain cock 216-2851 | N/A | 3 weeks | General Cargo | This item needs to be ordered with item 106145 |  | 3.5 | 50 | 40 | 30 | 0.06 | 0 | 88.72 | USD | 0 | Bürkle 0431-1030 | CE |  | 39269097
106144 | Balloon HDPE 10 L | TTM | 1 | Bürkle, Canister wide mouth with screw cap, capacity 10Lt. HDPE, transparent. Extra wide filling opening also makes cleaning inside much easier, which is important particularly when storing and transporting sensitive substances. UV-stabilised. Good chemical resistance. Heat resistant up to 110 °C. With threaded connector on the side just above the canister bottom, and drain cock 216-2850 | N/A | 3 weeks | General Cargo | This item needs to be ordered with item 106145 |  | 2.8 | 40 | 30 | 25 | 0.03 | 0 | 51.32 | USD | 0 | Bürkle 0431-1010 | CE |  | 39269097
106145 | Stop cock PP | TTM | 1 | Bürkle stop cock PP 3/4 thread 216-2854 | N/A | 3 weeks | General Cargo | Product linked to 106143 and 106144 |  | 0.3 | 15 | 10 | 10 | 0.0015 | 0 | 15.83 | USD | 0 | Bürkle 0530-1000 | CE |  | 39269097
106146 | Measuring pipettes 1 ml - Glass | TTM | 10 | VWR Blaubrand®, measuring pipettes 1 ml, graduated in brown 0,01 gdt., conformity certified, with ring mark and lettering in blue, pack of 10 pieces, ISO 835, 612-3152 | N/A | 3 weeks | General Cargo |  |  | 0.5 | 4 | 4 | 35 | 0.00056 | 0 | 29.1 | USD | 0 | Brand 278 17 | CE |  | 70179000
106147 | Measuring pipettes 2 ml - Glass | TTM | 10 | VWR Blaubrand®, measuring pipettes 2 ml, graduated in brown 0,01 gdt., conformity certified, with ring mark and lettering in blue, pack of 10 pieces, ISO 835, 612-3154 | N/A | 3 weeks | General Cargo |  |  | 0.5 | 4 | 4 | 35 | 0.00056 | 0 | 29.1 | USD | 0 | Brand 278 19 | CE |  | 70179000
106148 | Measuring pipettes 5 ml - Glass | TTM | 10 | VWR Blaubrand®, measuring pipettes 5 ml, graduated in brown 0,01 gdt., conformity certified, with ring mark and lettering in blue, pack of 10 pieces, ISO 835, 612-3157 | N/A | 3 weeks | General Cargo |  |  | 0.5 | 4 | 4 | 35 | 0.00056 | 0 | 34.04 | USD | 0 | Brand 278 21 | CE |  | 70179000
106149 | Measuring pipettes 10 ml - Glass | TTM | 10 | VWR Blaubrand®, measuring pipettes 10 ml, graduated in brown 0,01 gdt., conformity certified, with ring mark and lettering in blue, pack of 10 pieces, ISO 835, 612-3159 | N/A | 3 weeks | General Cargo |  |  | 0.3 | 38 | 6 | 5 | 0.00114 | 0 | 34.86 | USD | 0 | Brand 278 23 | CE |  | 70179000
106152 | Woulff's flask | TTM | 1 | Duran® woulff's bottle kit, H995.1: Pressure plus-screw thread bottle 1000ml, 1 top without filter plate; 2x plastic cutting ring Ø outer 8.6 mm; 2x screw cap GL 14, 1x screw cap GL 45 and 1x silicone ring | N/A | 3 weeks | General Cargo |  |  | 1 | 20 | 20 | 25 | 0.01 | 0 | 54.3 | USD | 0 | Duran Group H995.1, C217.1, L992.1, L993.1, L994.1, L995.1 Carl Roth | CE |  | 70172000
106157 | Swan neck bottles - Set | TTM | 1 | Swan neck bottles, natural colour with lid and dispenser tube, set of one each 250, 500, 1000ml | N/A | 3 weeks | General Cargo |  |  | 0.1 | 18 | 13 | 36 | 0.008424 | 0 | 3.55 | USD | 0 | Mraz 1100/250, 1100/500, 1100/1000, 2 x 1500/25, 1500/28 | CE |  | 39239000
106158 | Bowl | TTM | 1 | Universal bowl, 12l, PP transparent, multi usage for storage and laboratory utensils, 510x345x110mm | N/A | 3 weeks | General Cargo |  |  | 0.3 | 52 | 35 | 12 | 0.02184 | 0 | 12.2 | USD | 0 | 216-1466 VWR Kartell 548 | CE |  | 39239000
106159 | Erlenmeyer flasks - 2000ml | TTM | 10 | Erlenmeyer flask, borosilicate Din/ ISO 3585, wide neck, graduated, 2000ml | N/A | 3 weeks | General Cargo |  |  | 0.04 | 16 | 16 | 29 | 0.007424 | 0 | 78.6 | USD | 0 | Herenz 1070238 | CE |  | 70172000
106160 | Erlenmeyer flasks - 1000ml | TTM | 10 | Erlenmeyer flask, borosilicate Din/ ISO 3585, wide neck, graduated, 1000ml | N/A | 3 weeks | General Cargo |  |  | 2.9 | 70 | 28 | 24 | 0.04704 | 0 | 43.08 | USD | 0 | Herenz 1070237 | CE |  | 70172000
106161 | Erlenmeyer flasks - 500ml | TTM | 10 | Erlenmeyer flask, borosilicate Din/ ISO 3585, wide neck, graduated, 500ml | N/A | 3 weeks | General Cargo |  |  | 2.4 | 56 | 23 | 20 | 0.02576 | 0 | 27.29 | USD | 0 | Herenz 1070236 | CE |  | 70172000
106162 | Erlenmeyer flasks - 250ml | TTM | 10 | Erlenmeyer flask, borosilicate Din/ ISO 3585, wide neck, graduated, 250ml | N/A | 3 weeks | General Cargo |  |  | 1.8 | 47 | 19 | 17 | 0.015181 | 0 | 19.24 | USD | 0 | Herenz 1070234 | CE |  | 70172000
106163 | Erlenmeyer flasks - 100ml | TTM | 10 | Erlenmeyer flask, borosilicate Din/ ISO 3585, wide neck, graduated, 100ml | N/A | 3 weeks | General Cargo |  |  | 1 | 38 | 15 | 13 | 0.00741 | 0 | 16.28 | USD | 0 | Herenz 1070232 | CE |  | 70172000
106164 | Erlenmeyer flasks - 50ml | TTM | 10 | Erlenmeyer flask, borosilicate Din/ ISO 3585, wide neck, graduated, 50ml | N/A | 3 weeks | General Cargo |  |  | 0.5 | 27 | 11 | 10 | 0.00297 | 0 | 17.92 | USD | 0 | Herenz 1070231 | CE |  | 70172000
106165 | Measuring cylinder - 1000ml - 0.5ml graduated | TTM | 1 | Duran® measuring cyl. Glass A, blue graduated, hexagonal foot and protective collar made of PP can be slipped on easily, DIN 12680-2, comes with batch identification, autoclavable, capacity 1000ml graduation 10ml | N/A | 3 weeks | General Cargo |  |  | 0.6 | 47 | 15 | 14 | 0.00987 | 0 | 39.35 | USD | 0 | Hirschmann 2270193 | CE |  | 70172000
106166 | Measuring cylinder - 500ml - 2.5ml graduated | TTM | 1 | Duran® measuring cyl. Glass A, blue graduated, hexagonal foot and protective collar made of PP can be slipped on easily, DIN 12680-2, comes with batch identification, autoclavable, capacity 500ml graduation 1ml | N/A | 3 weeks | General Cargo |  |  | 0.7 | 13 | 13 | 36 | 0.006084 | 0 | 24.93 | USD | 0 | Hirschmann 2270190 | CE |  | 70172000
106167 | Measuring cylinder - 250ml - 1ml graduated | TTM | 1 | Duran® measuring cyl. Glass A, blue graduated, hexagonal foot and protective collar made of PP can be slipped on easily, DIN 12680-2, comes with batch identification, autoclavable, capacity 500ml graduation 5ml | N/A | 3 weeks | General Cargo |  |  | 0.5 | 42 | 12 | 10 | 0.00504 | 0 | 15.01 | USD | 0 | CX78.1 Carl Roth Hirschmann 2270186 | CE |  | 70172000
106168 | Measuring cylinder - 100ml - 0.5ml graduated | TTM | 1 | Duran® measuring cyl. Glass A, blue graduated, hexagonal foot and protective collar made of PP can be slipped on easily, DIN 12680-2, comes with batch identification, autoclavable, capacity 100ml graduation 0,5ml | N/A | 3 weeks | General Cargo |  |  | 0.06 | 8 | 8 | 26 | 0.001664 | 0 | 8.43 | USD | 0 | CX77.1 Carl Roth Hirschmann 2270180 | CE |  | 70172000
106169 | Measuring cylinder - 50ml - 0.5ml graduated | TTM | 1 | Duran® measuring cyl. Glass A, blue graduated, hexagonal foot and protective collar made of PP can be slipped on easily, DIN 12680-2, comes with batch identification, autoclavable, capacity 50ml graduation 0.5ml | N/A | 3 weeks | General Cargo |  |  | 0.1 | 25 | 8 | 7 | 0.0014 | 0 | 7.48 | USD | 0 | CX76.1 Carl Roth Hirschmann 2270175 | CE |  | 70172000
106170 | Measuring cylinder - 25ml - 0.25ml graduated | TTM | 1 | Duran® measuring cyl. Glass A, blue graduated, hexagonal foot and protective collar made of PP can be slipped on easily, DIN 12680-2, comes with batch identification, autoclavable, capacity 25ml graduation 0,25ml | N/A | 3 weeks | General Cargo |  |  | 0.05 | 6 | 6 | 17 | 0.000612 | 0 | 6.77 | USD | 0 | Hirschmann 2270170 | CE |  | 70172000
106171 | Measuring cylinder - 1000ml - 10ml graduated | TTM | 1 | Measuring cyl., transparent PMP (TPX®), tall form, autoclavable, 1000ml, graduation 10ml, No meniscus, chemically non-absorbent, non-wetting surface, excellent chemical resistance, Can be used with liquids up to +170 °C, ISO 6706 | N/A | 3 weeks | General Cargo |  |  | 0.3 | 12 | 12 | 15 | 0.00216 | 0 | 19.73 | USD | 0 | 612-4074 VWR Kartell 2566 | CE |  | 39239000
106172 | Measuring cylinder - 500ml - 5ml graduated | TTM | 1 | Measuring cyl., transparent PMP (TPX®), tall form, autoclavable, 500ml, graduation 5.0ml, No meniscus, chemically non-absorbent, non-wetting surface, excellent chemical resistance, Can be used with liquids up to +170 °C, ISO 6706 | N/A | 3 weeks | General Cargo |  |  | 0.3 | 12 | 12 | 15 | 0.00216 | 0 | 11.75 | USD | 0 | 612-4073 VWR Kartell 2565 | CE |  | 39239000
106173 | Measuring cylinder - 250ml - 2ml graduated | TTM | 1 | Measuring cyl., transparent PMP (TPX®), tall form, autoclavable, 250ml, graduation 2.0ml, No meniscus, chemically non-absorbent, non-wetting surface, excellent chemical resistance, Can be used with liquids up to +170 °C, ISO 6706 | N/A | 3 weeks | General Cargo |  |  | 0.3 | 12 | 12 | 15 | 0.00216 | 0 | 8.63 | USD | 0 | 612-4072 VWR Kartell 2564 | CE |  | 39239000
106174 | Measuring cylinder - 100ml - 1ml graduated | TTM | 1 | Measuring cyl., transparent PMP (TPX®), tall form, autoclavable, 100ml, graduation 1.0ml, No meniscus, chemically non-absorbent, non-wetting surface, excellent chemical resistance, Can be used with liquids up to +170 °C, ISO 6706 | N/A | 3 weeks | General Cargo |  |  | 0.3 | 12 | 12 | 15 | 0.00216 | 0 | 5.34 | USD | 0 | 612-4071 VWR Kartell 2563 | CE |  | 39239000
106175 | Measuring cylinder - 50ml - 1ml graduated | TTM | 1 | Measuring cyl., transparent PMP (TPX®), tall form, autoclavable, 50ml, graduation 1.0ml, No meniscus, chemically non-absorbent, non-wetting surface, excellent chemical resistance, Can be used with liquids up to +170 °C, ISO 6706 | N/A | 3 weeks | General Cargo |  |  | 0.3 | 12 | 12 | 15 | 0.00216 | 0 | 4.68 | USD | 0 | 612-4070 VWR Kartell 2562 | CE |  | 39239000
106176 | Thick-walled glass flask - 2500ml | TTM | 1 | Culture media bottle, 2500ml, clear Duran glass | N/A | 3 weeks | General Cargo |  |  | 1 | 25 | 25 | 40 | 0.025 | 0 | 57.9 | USD | 0 | Duran Group 214516604 | CE |  | 70172000
106177 | Thick-walled glass flask - 2000ml | TTM | 1 | Thick walled glass bottle, Grade 3.3 borosilicate glass, with graduations, DIN GL 45 thread, pouring ring and blue PP screw cap, transparent, very good chemical resistants, autoclavable, according to ISO 4796, with ring and cap of PP, 2000 ml | N/A | 3 weeks | General Cargo |  |  | 1 | 25 | 25 | 40 | 0.025 | 0 | 8.9 | USD | 0 | Kimble 632414321950 | CE |  | 70172000
106178 | Thick-walled glass flask - 1000ml | TTM | 1 | Thick walled glass bottle, Grade 3.3 borosilicate glass, with graduations, DIN GL 45 thread, pouring ring and blue PP screw cap, transparent, very good chemical resistants, autoclavable, according to ISO 4796, with ring and cap of PP, 1000 ml | N/A | 3 weeks | General Cargo |  |  | 0.6 | 9 | 9 | 24 | 0.001944 | 0 | 2.74 | USD | 0 | Kimble 632414321940 | CE |  | 70172000
106179 | Thick-walled glass flask - 500ml | TTM | 1 | Thick walled glass bottle, Grade 3.3 borosilicate glass, with graduations, DIN GL 45 thread, pouring ring and blue PP screw cap, transparent, very good chemical resistants, autoclavable, according to ISO 4796, with ring and cap of PP, 500 ml | N/A | 3 weeks | General Cargo |  |  | 3.9 | 46 | 19 | 21 | 0.018354 | 0 | 2.46 | USD | 0 | VWR Collection 215-1594 | CE |  | 70172000
106182 | Screw caps with weld in PFTE membrane | TTM | 1 | Screw caps with weld in PFTE membrane, 0.2 µm pore size, autoclavable to 140 °C | N/A | 3 weeks | General Cargo |  |  | 0.08 | 6 | 6 | 3 | 0.000108 | 0 | 5.39 | USD | 0 | Duran Group 292402807 | CE |  | 39235010
106183 | Laboratory beaker, low form with graduation and spout, 2000 ml | TTM | 1 | Laboratory beaker, low form with graduation and spout, 2000 ml | N/A | 3 weeks | General Cargo |  |  | 0.3 | 16 | 16 | 19 | 0.004864 | 0 | 6.34 | USD | 0 | Herenz 1070139 | CE |  | 70172000
106184 | Laboratory beaker, low form with graduation and spout, 1000 ml | TTM | 1 | Laboratory beaker, low form with graduation and spout, 1000 ml | N/A | 3 weeks | General Cargo |  |  | 0.3 | 13 | 13 | 15 | 0.002535 | 0 | 3.79 | USD | 0 | Herenz 1070138 | CE |  | 70172000
106185 | Laboratory beaker, low form with graduation and spout, 600 m | TTM | 1 | Laboratory beaker, low form with graduation and spout, 600 ml | N/A | 3 weeks | General Cargo |  |  | 0.2 | 11 | 11 | 13 | 0.001573 | 0 | 2.61 | USD | 0 | Herenz 1070136 | CE |  | 70172000
106186 | Laboratory beaker, low form with graduation and spout, 250 m | TTM | 1 | Laboratory beaker, low form with graduation and spout, 250 ml | N/A | 3 weeks | General Cargo |  |  | 0.2 | 9 | 9 | 10 | 0.00081 | 0 | 1.77 | USD | 0 | Herenz 1070134 | CE |  | 70172000
106187 | Laboratory beaker, low form with graduation and spout, 100 m | TTM | 1 | Laboratory beaker, low form with graduation and spout, 100 ml | N/A | 3 weeks | General Cargo |  |  | 0.2 | 7 | 7 | 8 | 0.000392 | 0 | 1.54 | USD | 0 | Herenz 1070132 | CE |  | 70172000
106188 | Laboratory beakers; 100 ml (ETFE) | TTM | 1 | Beakers, Low Form, with spout and easy-to-read black printed graduations, very good chemical resistance, good thermal resistance from –100…+150 °C, Autoclavable, ISO 7056, 100ml, graduated | N/A | 3 weeks | General Cargo |  |  | 0.2 | 15 | 15 | 20 | 0.0045 | 0 | 12.66 | USD | 0 | Vitlab 110404 | CE |  | 70172000
106189 | Laboratory beakers; 250 ml (ETFE) | TTM | 1 | Beakers, Low Form, with spout and easy-to-read black printed graduations, very good chemical resistance, good thermal resistance from –100…+150 °C, Autoclavable, ISO 7056, 250ml, graduated | N/A | 3 weeks | General Cargo |  |  | 0.2 | 15 | 15 | 20 | 0.0045 | 0 | 21.75 | USD | 0 | Vitlab 110604 | CE |  | 70172000
106190 | Laboratory beakers; 500 ml (ETFE) | TTM | 1 | Beakers, Low Form, with spout and easy-to-read black printed graduations, very good chemical resistance, good thermal resistance from –100…+150 °C, Autoclavable, ISO 7056, 500ml, graduated | N/A | 3 weeks | General Cargo |  |  | 0.2 | 20 | 20 | 25 | 0.01 | 0 | 35.66 | USD | 0 | Vitlab 110904 | CE |  | 70172000
106191 | Laboratory beakers; 1000 ml (ETFE) | TTM | 1 | Beakers, Low Form, with spout and easy-to-read black printed graduations, very good chemical resistance, good thermal resistance from –100…+150 °C, Autoclavable, ISO 7056, 1000ml, graduated | N/A | 3 weeks | General Cargo |  |  | 0.2 | 20 | 20 | 25 | 0.01 | 0 | 59.41 | USD | 0 | Vitlab 111004 | CE |  | 70172000
106192 | Laboratory beakers; 2000 ml (PP) | TTM | 1 | Beakers; 2000 ml; PP, ISO 7056 and DIN 12331 max temp. 135°C; low form; with spout and graduation, autoclavable | N/A | 3 weeks | General Cargo |  |  | 0.2 | 17 | 15 | 19 | 0.004845 | 0 | 5.82 | USD | 0 | Hecht 1935/2000 | CE |  | 70172000
106193 | Volumetric flasks 1000 ml | TTM | 1 | Volumetric Flask, BLAUBRAND®, Class A, blue graduation, conformity certificate, DURAN®, with PP stopper calibrated to contain (TC,In), DIN/ ISO 1042, toleranz 0.4ml, 1000ml | N/A | 3 weeks | General Cargo |  |  | 0.8 | 16 | 16 | 35 | 0.00896 | 0 | 29.46 | USD | 0 | Brand 372 53 | CE |  | 70172000
106194 | Volumetric flasks 500 ml | TTM | 1 | Volumetric Flask, BLAUBRAND®, Class A, blue graduation, conformity certificate, DURAN®, with PP stopper calibrated to contain (TC,In), DIN/ ISO 1042, toleranz 0.25ml, 500ml | N/A | 3 weeks | General Cargo |  |  | 0.5 | 11.5 | 11.5 | 27.5 | 0.0036369 | 0 | 18.64 | USD | 0 | Brand 372 52 | CE |  | 70172000
106195 | Volumetric flasks 250 ml | TTM | 1 | Volumetric Flask, BLAUBRAND®, Class A, blue graduation, conformity certificate, DURAN®, with PP stopper calibrated to contain (TC,In), DIN/ ISO 1042, toleranz 0.15ml, 250ml | N/A | 3 weeks | General Cargo |  |  | 0.6 | 9 | 9 | 24 | 0.001944 | 0 | 17.2 | USD | 0 | Brand 372 51 | CE |  | 70172000
106196 | Volumetric flasks 200 ml | TTM | 1 | Volumetric Flask, BLAUBRAND®, Class A, blue graduation, conformity certificate, DURAN®, with PP stopper calibrated to contain (TC,In), DIN/ ISO 1042, toleranz 0.15ml, 200ml | N/A | 3 weeks | General Cargo |  |  | 0.6 | 12 | 12 | 28 | 0.004032 | 0 | 14.68 | USD | 0 | Brand 372 50 | CE |  | 70172000
106197 | Volumetric flasks 100 ml | TTM | 1 | Volumetric Flask, BLAUBRAND®, Class A, blue graduation, conformity certificate, DURAN®, with PP stopper calibrated to contain (TC,In), DIN/ ISO 1042, toleranz 0.10ml, 100ml | N/A | 3 weeks | General Cargo |  |  | 0.2 | 18 | 10 | 8 | 0.00144 | 0 | 10.21 | USD | 0 | Brand 372 49 | CE |  | 70172000
106198 | Volumetric flasks 50 ml | TTM | 1 | Volumetric flasks 50 ml, borosilicate 3.3, class A, autoclavable, DIN/ ISO 1042, calibrated IN, with 1 circular mark, toleranz ± 0.06 ml, per piece | N/A | 3 weeks | General Cargo |  |  | 0.4 | 10 | 10 | 20 | 0.002 | 0 | 7.48 | USD | 0 | 612-5077 VWR Brand 37248 | CE |  | 70172000
106199 | Volumetric flasks 25 ml | TTM | 1 | Volumetric flasks 25 ml, borosilicate 3.3, class A, autoclavable, DIN/ ISO 1042, calibrated IN, with 1 circular mark, toleranz ± 0.04 ml, per piece | N/A | 3 weeks | General Cargo |  |  | 0.1 | 5 | 5 | 12.5 | 0.0003125 | 0 | 6.66 | USD | 0 | 612-5076 VWR Brand 37247 | CE |  | 70172000
106200 | Volumetric flasks 20 ml | TTM | 1 | Volumetric flasks 20 ml, borosilicate 3.3, class A, autoclavable, DIN/ ISO 1042, calibrated IN, with 1 circular mark, toleranz ± 0.04 ml, per piece | N/A | 3 weeks | General Cargo |  |  | 0.1 | 4.5 | 4.5 | 12 | 0.000243 | 0 | 6.66 | USD | 0 | 612-5075 VWR Brand 37246 | CE |  | 70172000
106201 | Laboratory glassware drying rack | TTM | 1 | Draining rack, wire, pre coated, for 24 glasses/ 20 flasks | N/A | 3 weeks | General Cargo |  |  | 2 | 50 | 20 | 100 | 0.1 | 0 | 77.81 | USD | 0 | WCF 5281-01-11 | CE |  | 73262000
106202 | Funnel, soda lime glass, plain, short stem, diameter = 75 mm | TTM | 1 | Funnel vitlab, PP plain with short stern, Ø 75mm | N/A | 3 weeks | General Cargo |  |  | 0.8 | 10 | 10 | 15 | 0.0015 | 0 | 0.67 | USD | 0 | Vitlab 41394 | CE |  | 39239000
106203 | Funnel, soda lime glass, plain, short stem, diameter = 125 mm | TTM | 1 | Funnel soda lime glass, DIN 12445, plain angle 60°, with short and angled -ground stern, Ø 125mm | N/A | 3 weeks | General Cargo |  |  | 1 | 15 | 15 | 20 | 0.0045 | 0 | 7.4 | USD | 0 | Bochemia Christal Handelsgesellschaft mbH 221-1727 VWR | CE |  | 70171000
106204 | Powder funnel, diameter 100mm | TTM | 1 | Powder funnel , diameter 100mm, PP transparent | N/A | 3 weeks | General Cargo |  |  | 0.2 | 12 | 12 | 15 | 0.00216 | 0 | 1.8 | USD | 0 | Vitlab 70994 | CE |  | 39239000
17. Media kitchen: Other consumables
106108 | Aqua purificator cabinet for two de-mineralising cartridges | TTM | 1 | Stainless steel storage cabinet for 2x cartridges | N/A | 3 weeks | General Cargo |  |  | 37 | 35 | 68 | 90 | 0.2142 | 0 | 617.5 | USD | 0 | Gede ST200 | CE |  | 73269098
106109 | Pipette box with lid for 20 - 25 cm long pipettes | TTM | 1 | Pipette box aluminium/stainless steel with lid, square, silicon insert, desinfectable and autoclavable/sterilisable up to +205 °C , length total 290 mm, usable length of 257 mm | N/A | 3 weeks | General Cargo |  |  | 0.9 | 40 | 19 | 15 | 0.0114 | 0 | 65.61 | USD | 0 | Bochem 8871 | CE |  | 73269098
106110 | Pipette box with lid for 35 cm long pipettes | TTM | 1 | Pipette box aluminium/stainless steel with lid, square, silicon insert, desinfectable and autoclavable/sterilisable up to +205 °C , length total 430 mm, usable length of 412 mm | N/A | 3 weeks | General Cargo |  |  | 0.5 | 80 | 80 | 43 | 0.2752 | 0 | 96.58 | USD | 0 | VWR Collection 612-2065 | CE |  | 73269098
106111 | Inlay for pipette box | TTM | 1 | Inlay for pipette box Vitan 75 shoe A temp -10 to > +200°C | N/A | 3 weeks | General Cargo |  |  | 0.05 | 3 | 3 | 1 | 9e-06 | 0 | 7 | USD | 0 | Deutsch & Neumann 520 3015 | CE |  | 73269098
106112 | Pipette washer | TTM | 1 | Kartell Automatic Pipette Washer, with 2x baskets and pipette handle, 170 Ø | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 35 | 17 | 17 | 99 | 0.028611 | 0 | 179 | USD | 0 | Kartell 217 VWR 149-1110, -31, -30 | CE |  | 39269097
106360 | Aluminium foil | TTM | 1 | Aluminum foil, 600mm width, 100m long, 0,030 mm in dispenser box | N/A | 3 weeks | General Cargo |  |  | 5.3 | 67 | 11 | 9 | 0.006633 | 0 | 84 | USD | 0 | Korff 60150 | CE |  | 76072099
106363 | Tube brush | TTM | 5 | Rotalab® tub brush, length 300mm, Ø25mm, packaging 5 pcs | N/A | 3 weeks | General Cargo |  |  | 0.1 | 32 | 7 | 2 | 0.000448 | 0 | 12.74 | USD | 0 | Hecht 2760/2 | CE |  | 96039099
106364 | Erlenmeyer, bottle brush | TTM | 1 | Erlenmeyer, bottle brush length 470mm | N/A | 3 weeks | General Cargo |  |  | 0.02 | 48 | 8 | 48 | 0.018432 | 0 | 5.54 | USD | 0 | Reitenspiess 23630234 | CE |  | 96039099
106379 | Salt for dish washer | TTM | 12kg | Salt: Kontrakalk, 1 box of 6x2 kg | N/A | 3 weeks | General Cargo |  |  | 12 | 27 | 18.5 | 23 | 0.011489 | 0 | 13.3 | USD | 0 | HENKEL-ECOLAB GH WOOD DIVISION 1161340 | SDS |  | 34022090
106380 | Rinse aid for dish washer | TTM | 1 L | Rinse aid, NEODISCHER N, 1lt bottles | N/A | 3 weeks | Dangerous goods |  |  | 1.5 | 11 | 6 | 24.5 | 0.001617 | 0 | 12.52 | USD | 0 | Dr. Weigert 420187 | SDS | UN 1805 Class 8 Packing group III | 34029090
106150 | Desiccator | TTM | 1 | Desiccator "Novus", art. No. 1164/20: stress, vacuum density and pressure tested, with std ground stopcock in the interchangeable cover, Porcelain for desiccators 1175/20, with various perforations and one large hole in centre | N/A | 3 weeks | General Cargo |  |  | 3 | 30 | 30 | 40 | 0.036 | 0 | 210.69 | USD | 0 | Duran Group, Haldenwanger, Lenz Laborglasinstrumente 24 770 61, 244206109 , 119C/190, 2.8224.04 | CE |  | 70172000
106151 | Membrane vacuum pump for desiccator | TTM | 1 | Laborport® N86KN.18: Mini Diaphragm Vacuum pump and compressor single head pump, dry running device used in a wide range of laboratory application, the unit is maintenance free; Transfers, compress and pump down without contamination; the heart of this very compact pump is a KNF structured diaphragm; This patented diaphragm was stressed-optimized using the Finite Elements method; Technical data: delivery (l/min)¹ 6; ultimate vacuum (mbar abs.) 100; operating (bar g) 2.4; permissible glass and ambient temperature +5 to +40°C, 65W, | N/A | 3 weeks | General Cargo | Product linked to 106153 |  | 2.3 | 18 | 16 | 93 | 0.026784 | 0 | 303.29 | USD | 0 | KNF Neuberger 046208/064242 | CE |  | 70172000
106153 | Vacuum tube, 3 m | TTM | 1 | Vacuum tube, 3 m | N/A | 3 weeks | General Cargo | Product linked to 106151 |  | 0.5 | 30 | 30 | 7 | 0.0063 | 0 | 20.12 | USD | 0 | Deutsch & Neumann 311 0816 | CE |  | 39239000
106154 | Silica gel for desiccator, 0.5 kg | TTM | 1 kg | Silica gel for desiccator, 1 kg, with humidity indicator | N/A | 3 weeks | General Cargo |  |  | 1 | 20 | 15 | 25 | 0.0075 | 0 | 55.25 | USD | 0 | Merck Millipore 1.01969.1000 | CE |  | 38249097
106155 | Bunsen burner | TTM | 1 | Teclu gas burner: diameter of flame exit: outer Ø 18mm, inner Ø mm | N/A | 3 weeks | General Cargo |  |  | 0.5 | 12 | 12 | 16 | 0.002304 | 0 | 35.27 | USD | 0 | Juchheim 1262DP | CE |  | 84198998
106156 | Tripod | TTM | 1 | Tripod stand with fire mat | N/A | 3 weeks | General Cargo |  |  | 0.5 | 17 | 17 | 25 | 0.007225 | 0 | 10.45 | USD | 0 | Usbeck 451-0093, 451-9201 VWR | CE |  | 73259990
106206 | Strainer, stainless steel, 1 grip, 2 eyes, diameter 20 cm | TTM | 1 | Strainer, stainless steel, diameter 190mm | N/A | 3 weeks | General Cargo |  |  | 0.9 | 22 | 22 | 24 | 0.011616 | 0 | 14.2 | USD | 0 | Fackelmann 510-9954 VWR | CE |  | 70179000
106207 | Set of spatula | TTM | 1 | Set of spatula: stainless steel, electrolytically polished, length approx. 150 mm, blade, 40 mm length, width ranging 3,6,9 mm, 6 pieces in most common sizes of double and micro spatulas, 3 of them with spoon 9x5, 5x3, 9x5mm | N/A | 3 weeks | General Cargo |  |  | 0.6 | 20 | 20 | 10 | 0.004 | 0 | 18.43 | USD | 0 | RSG Solingen, Buchem VWR 231-2237, 231-2239, 231-0047, 231-1355, 231-0239, 231-2244 | CE |  | 90189049
106208 | Forceps; pointed; straight; 18/8 stainless steel, polished | TTM | 1 | Forceps, stainless steel, anatom. pointed 130mm, polished | N/A | 3 weeks | General Cargo |  |  | 0.1 | 4 | 4 | 15 | 0.00024 | 0 | 1.39 | USD | 0 | Herenz 1110211 | CE |  | 82032010
106353 | Filter paper round 185 mm | TTM | 100 | Filter paper round, weight approx. 85 g/m², retention > 4µm, diameter 185 mm, packs of 100 | N/A | 3 weeks | General Cargo |  |  | 0.3 | 19 | 19 | 3 | 0.001083 | 0 | 35.35 | USD | 0 | Whatman (GE Healthcare) 515-4306 VWR GE 10311814 | CE |  | 48232000
106354 | Filter paper round 150 mm | TTM | 100 | Filter paper round, weight approx. 85 g/m², retention > 4µm, diameter 150 mm, packs of 100 | N/A | 3 weeks | General Cargo |  |  | 0.2 | 16 | 16 | 3 | 0.000768 | 0 | 24.66 | USD | 0 | Whatman (GE Healthcare) 515-4305 VWR GE10311812 | CE |  | 48232000
18. Autoclave basic unit for media kitchen
106116 | Autoclave basic unit for Media Kitchen | TTM | 1 | Autoclave type VX-75, with integrated stainless steel structure, Vertical top-loading unit, 25 programs, RS232 and RS 485 for external data transfer, 80ltr. Chamber dimension 400x 600mm, max pressure 5bar/ 150°C, 380V, 50/60 Hz | N/A | 6 weeks | General Cargo | · GDF has 2 options for Autoclave basic unit for Media Kitchen ( Systec and SHP) please select the brand desired; · This item needs to be ordered with the below items: 1. 106118 2. 106119 3. 106120 4. 106121 5. 106590 6. 106490 · 2 years warranty |  | 184 | 55 | 78 | 96 | 0.41184 | 0 | 15459.12 | USD | 0 | Systec 1075
106118 | Exhaust air filtration system for Autoclave | TTM | 1 kit | Exhaust air filtration system for above mentioned autoclaves | N/A | 6 weeks | General Cargo | Product linked to 106116 |  | 5 | 20 | 20 | 30 | 0.012 | 0 | 3248.71 | USD | 0 | Systec 8030
106119 | Autoclave stainless steel buckets | TTM | 1 kit | Autoclave basket stainless steel bucket for sterilisation, 360x355 mm | N/A | 6 weeks | General Cargo | Product linked to 106116 |  | 1.5 | 40 | 40 | 40 | 0.064 | 0 | 287.13 | USD | 0 | Systec 7210
106120 | Autoclave stainless steel baskets | TTM | 1 | Autoclave basket 360x225 mm | N/A | 6 weeks | General Cargo | Product linked to 106116 |  | 3 | 40 | 40 | 40 | 0.064 | 0 | 1043.44 | USD | 0 | Systec 7240
106121 | Printer for Autoclave | TTM | 1 | Printer for autoclave | N/A | 6 weeks | General Cargo | Product linked to 106116 |  | 2 | 30 | 30 | 40 | 0.036 | 0 | 1346.85 | USD | 0 | Systec 8011
106580 | Compressor for Autoclave 106116 | TTM | 1 | Compressor Planet-Air L-S50-15, 1-motor with suction capacity 50 I/min, container volume 15 liter, with pressure reducer, water strainer, quick coupler and 0,01 micro filter. | N/A | 6 weeks | General Cargo | · Product linked to 106116; · 2 years warranty; · Voltage / Amp / Frequency: 230 V |  | 30 | 40 | 40 | 50 | 0.08 | 0 | 1429.86 | USD | 0 | Planet-Air L-S50-15
106490 | Installation and Training - Autoclave | TTM | 1 | Installation of autoclave in the laboratory with if needed additional pre-installation work as plumbing, electricity. Training on use and troubleshooting of autoclave | N/A | On demand | Service | · Product linked to 106116; · For each autoclave please calculate 5 “units” for the installation, as 5 days is estimated for the completion of the installation of one autoclave. E.g: for 1 autoclave= 5 installations. Final invoice will be on actuals |  | N/A | N/A | N/A | N/A | N/A | N/A | 900 | USD | 0 | N/A | N/A |  | N/A
106116 | Autoclave basic unit for Media Kitchen | TTM | 1 | Autoclave Laboclav 80MSLV with blue lid for waste management, tank vol. 30lt, chamber size 82lt, max pressure 2.8 bar, max temp. 143°C, voltage 3N 400V~(±5%), 50 Hz, 16A, 4.5KW, with liquid control unit, liquid sterilization and recooling system | N/A | 6 weeks | General Cargo | · GDF has 2 options for Autoclave basic unit for Media Kitchen ( Systec and SHP) please select the brand desired; · This item needs to be ordered with the below items: 1. 106118 2. 106119 3. 106120 4. 106121 5. 106590 6. 106490 · 2 years warranty |  | 196 | 78 | 94 | 65 | 0.47658 | 0 | 21323.38 | USD | 0 | SHP
106118 | Exhaust air filtration system for Autoclave | TTM | 1 kit | Exhaust air filter system for Laboclav 80V and 80 MSLV: filter element 0,2 µm, Pt 100 sensor, filter control, stainless steel housing | N/A | 6 weeks | General Cargo | Product linked to 106116 |  | 15 | 70 | 50 | 80 | 0.28 | 0 | 2664.09 | USD | 0 | SHP
106119 | Autoclave stainless steel buckets | TTM | 1 kit | Autoclave basket stainless steel for Laboclav 80, size: 80x400x300 mm | N/A | 6 weeks | General Cargo | Product linked to 106116 |  | 1.5 | 10 | 30 | 40 | 0.012 | 0 | 427.57 | USD | 0 | SHP
106120 | Autoclave stainless steel baskets | TTM | 1 | Autoclave basket with drip tray for Laboclav 80, size: 100x400x250 mm | N/A | 6 weeks | General Cargo | Product linked to 106116 |  | 3 | 10 | 10 | 50 | 0.02 | 0 | 180.07 | USD | 0 | SHP
106121 | Printer for Autoclave | TTM | 1 | Printer for autoclave Laboclav 55-195 | N/A | 6 weeks | General Cargo | Product linked to 106116 |  | 2 | 20 | 20 | 10 | 0.004 | 0 | 1053.8 | USD | 0 | SHP
106580 | Compressor for Autoclave 106116 | TTM | 1 | Compressor Planet-Air L-S50-15, 1-motor with suction capacity 50 I/min, container volume 15 litter, with pressure reducer, water strainer, quick coupler and 0,01 micro filter. | N/A | 6 weeks | General Cargo | · Product linked to 106116; · 2 years warranty; · Voltage / Amp / Frequency: 230 V |  | 2 | 40 | 40 | 50 | 0.08 | 0 | 1053.8 | USD | 0 | SHP
106490 | Installation and Training - Autoclave | TTM | 1 | Installation of autoclave in the laboratory with if needed additional pre-installation work as plumbing, electricity. Training on use and troubleshooting of autoclave | N/A | On demand | Service | · Product linked to 106116; · For each autoclave please calculate 5 “units” for the installation, as 5 days is estimated for the completion of the installation of one autoclave. E.g: for 1 autoclave= 5 installations. Final invoice will be on actuals |  | N/A | N/A | N/A | N/A | N/A | N/A | 900 | USD | 0 | N/A | N/A |  | N/A
19. Autoclave basic unit for waste management
106117 | Autoclave basic unit for Waste Management | TTM | 1 | Autoclave type VX-75, with integrated stainless steel structure, Vertical top-loading unit, 25 programs, RS232 and RS 485 for external data transfer, 80ltr. Chamber dimension 400x 600mm, max pressure 5bar/ 150°C, 380V, 50/60 Hz, with integrated vaccum system for waste autoclaving | N/A | 6 weeks | General Cargo | · GDF has 2 options for Autoclave basic unit for Waste Management (Systec and SHP). Please select the brand desired; · This item needs to be ordered with the below items: 1. 106118 2. 106119 3. 106120 4. 106121 5. 106490 · 2 years warranty |  | 184 | 55 | 78 | 96 | 0.41184 | 0 | 20121.28 | USD | 0 | Systec 1075, 9603
106118 | Exhaust air filtration system for Autoclave | TTM | 1 kit | Exhaust air filtration system for above mentioned autoclaves | N/A | 6 weeks | General Cargo | Product linked to 106117 |  | 5 | 20 | 20 | 30 | 0.012 | 0 | 3248.71 | USD | 0 | Systec 8030
106119 | Autoclave stainless steel buckets | TTM | 1 kit | Autoclave basket stainless steel bucket for sterilisation, 360x355 mm | N/A | 6 weeks | General Cargo | Product linked to 106117 |  | 1.5 | 40 | 40 | 40 | 0.064 | 0 | 287.13 | USD | 0 | Systec 7210
106120 | Autoclave stainless steel baskets | TTM | 1 | Autoclave basket 360x225 mm | N/A | 6 weeks | General Cargo | Product linked to 106117 |  | 3 | 40 | 40 | 40 | 0.064 | 0 | 1043.44 | USD | 0 | Systec 7240
106121 | Printer for Autoclave | TTM |  | Printer for autoclave | N/A | 6 weeks | General Cargo | Product linked to 106117 |  | 2 | 30 | 30 | 40 | 0.036 | 0 | 1346.85 | USD | 0 | Systec 8011
106490 | Installation and Training - Autoclave | TTM | 1 | Installation of autoclave in the laboratory with if needed additional pre-installation work as plumbing, electricity. Training on use and troubleshooting of autoclave | N/A | On demand | Service | · Product linked to 106117; · For each autoclave please calculate 5 “units” for the installation, as 5 days is estimated for the completion of the installation of one autoclave. E.g: for 1 autoclave= 5 installations. Final invoice will be on actuals |  | N/A | N/A | N/A | N/A | N/A | N/A | 900 | USD | 0 | N/A | N/A |  | N/A
106117 | Autoclave basic unit for Waste Management | TTM | 1 | Autoclave Laboclav 80V with red lid for waste management, tank vol. 30lt, chamber size 82lt, max pressure 2.8 bar, max temp. 143°C, voltage 3N 400V~(±5%), 50 Hz, 16A, 4.5KW | N/A | 6 weeks | General Cargo | · GDF has 2 options for Autoclave basic unit for Waste Management (Systec and SHP). Please select the brand desired; · This item needs to be ordered with the below items: 1. 106118 2. 106119 3. 106120 4. 106121 5. 106490 · 2 years warranty |  | 196 | 78 | 94 | 65 | 0.47658 | 0 | 17168.58 | USD | 0 | SHP
106118 | Exhaust air filtration system for Autoclave | TTM | 1 kit | Exhaust air filter system for Laboclav 80V and 80 MSLV: filter element 0,2 µm, Pt 100 sensor, filter control, stainless steel housing | N/A | 6 weeks | General Cargo | Product linked to 106117 |  | 15 | 70 | 50 | 80 | 0.28 | 0 | 2664.09 | USD | 0 | SHP
106119 | Autoclave stainless steel buckets | TTM | 1 kit | Autoclave basket stainless steel for Laboclav 80, size: 80x400x300 mm | N/A | 6 weeks | General Cargo | Product linked to 106117 |  | 1.5 | 10 | 30 | 40 | 0.012 | 0 | 427.57 | USD | 0 | SHP
106120 | Autoclave stainless steel baskets | TTM | 1 | Autoclave basket with drip tray for Laboclav 80, size: 100x400x250 mm | N/A | 6 weeks | General Cargo | Product linked to 106117 |  | 3 | 40 | 10 | 50 | 0.02 | 0 | 180.07 | USD | 0 | SHP
106121 | Printer for Autoclave | TTM |  | Printer for autoclave Laboclav 55-195 | N/A | 6 weeks | General Cargo | Product linked to 106117 |  | 2 | 20 | 20 | 10 | 0.004 | 0 | 1053.8 | USD | 0 | SHP
106490 | Installation and Training - Autoclave | TTM | 1 | Installation of autoclave in the laboratory with if needed additional pre-installation work as plumbing, electricity. Training on use and troubleshooting of autoclave | N/A | On demand | Service | · Product linked to 106117 · For each autoclave please calculate 5 “units” for the installation, as 5 days is estimated for the completion of the installation of one autoclave. E.g: for 1 autoclave= 5 installations. Final invoice will be on actuals |  | N/A | N/A | N/A | N/A | N/A | N/A | 900 | USD | 0 | N/A | N/A |  | N/A
20. Waste management: Chemicals and consumables
106328 | Formalin | TTM | 1 l | Formaldehyde solution min. 37% GR for analysis stabilized with about 10% methanol ACS,Reag. Ph Eur 1 * 1 l | 12 months | 3 weeks | Dangerous goods |  |  | 1.6 | 25 | 25 | 30 | 0.01875 | 0 | 33.82 | USD | 0 | Merck 1040031000 | SDS/ CoA | UN 2209 Class 8 Packing group III | 29121100
106624 | Stable chlorine disinfectant | TTM | 100 | For simultaneous disinfecting in the laboratory and medical field. Based on NaDCC (sodium dichloroisocyanurate). Weight of tablet 6 g. Recommended active chlorine content for acinetobacter bactericidal, (Clostridium difficile (also spores), E. Coli, Enterococcus, Klebsiella, Pseudomonas and Staphylococcus (MRSA) (tested acc. to EN 1276), contains 1000 ppm - this corresponds to a dosage of one tablet in 1 liter water. | N/A | 3 weeks | Dangerous goods |  |  | 0.01 | 13 | 6 | 3 | 0.000234 | 0 | 30.25 | USD | 0 | Medentech S11184 | SDS | UN 3077 Class 9 Packing group III | 1
106331 | Phenol | TTM | 1 Kg | Phenol, >99.5% p.a. detached crystals, Carbolic acid, Hyrdoxybenzene, C6H6O, M94,11g/mol, bp 181,8°C, d 1,06, flp 79°C mp 40,8°C, ADR 6,1II/WGK 2, EG-No 2036327 plastic bottle 1kg | 47 months | 3 weeks | Dangerous goods |  |  | 2.4 | 15 | 15 | 31 | 0.006975 | 0 | 69.66 | USD | 0 | Carl Roth 0040.2 | SDS/ CoA | UN 1671 Class 6.1 Packing group II | 29071100
106333 | Isopropanol | TTM | 5 L | Ethanol/ Isopropanol, 70 %, 5 L bottles | 47 months | 3 weeks | Dangerous goods |  |  | 4.8 | 22 | 28 | 32 | 0.019712 | 0 | 15.01 | USD | 0 | Kost Alkohole H079150 | SDS/ CoA | UN 1219 Class 3 Packing group II | 29051200
106250 | Stainless steel bucket including lid | TTM | 1 | Contacto stainless steel bucket with lid, 10lt | N/A | 3 weeks | General Cargo |  |  | 10 | 25 | 25 | 30 | 0.01875 | 0 | 55.71 | USD | 0 | Contacto 415/010, 418/010 | CE |  | 73269098
106251 | Discard bin 2 Litters | TTM | 1 | Bench top bin for small items like pipette tips and micro centrifuge tubes, hinged swinging lid, printed with biohazard symbol, holds 250x250mm autoclavable bags | N/A | 3 weeks | General Cargo |  |  | 0.6 | 21 | 21 | 23 | 0.010143 | 0 | 42.42 | USD | 0 | Bel-Art H131940011 | CE |  | 39269097
106252 | Funnel, stainless steel 18/8 | TTM | 1 | Contacto funnel stainless steel, diameter of ramp 15 mm, top diameter 120 mm, total height approx. 165 mm | N/A | 3 weeks | General Cargo |  |  | 0.5 | 12 | 12 | 15 | 0.00216 | 0 | 13.32 | USD | 0 | Contacto 5420/120 | CE |  | 90189084
106253 | Safety box for liquid | TTM | 30 | Round disposal box 2.0 lt made of PP2 with screw cap and label Large opening, suitable for ESR pipettes. Include 30 pcs | N/A | 3 weeks | General Cargo |  |  | 4.3 | 61 | 37 | 52 | 0.11736 | 0 | 59.33 | USD | 0 | Sarstedt 771515 | CE |  | 39269097
106255 | Waste containers for solids | TTM | 6 | Carton with integral biohazard bag HDPE, Snug fitting lid has flap through which waste is inserted, When carton is filled, internal safety cap is pulled into place and the entire unit is ready for incineration, size 196x196x254mm | N/A | 3 weeks | General Cargo |  |  | 2.1 | 51 | 43 | 11 | 0.024123 | 0 | 54.31 | USD | 0 | Kleinfeld Labortechnik F13205-0002 | CE |  | 48191000
106256 | Waste container for liquids | TTM | 1 | Sekuroka safety waste disposable system, made of break resistant (fluorinated) HDPE, system consists of waste container and removable safety funnel with lid, reduces emissions of volatile substances in the laboratory, integrated ventilating valve with PTFE syringe filter to prevent overflowing, removable sieve insert, 4 litre | N/A | 3 weeks | General Cargo |  |  | 0.6 | 15 | 15 | 30 | 0.00675 | 0 | 55.46 | USD | 0 | Nalgene 6379-0004 | CE |  | 39269097
106348 | Transparent polypropylene bags | TTM | 1000 | Sekuroka® disposable bag Made of 40 µm thick PP. Heat stable to +134 °C. Transparent, with „Biohazard“ warning, size 420x600mm | N/A | 3 weeks | General Cargo |  |  | 9 | 45 | 36 | 11 | 0.01782 | 0 | 175.96 | USD | 0 | VWR Collection 129-0588 VWR | CE |  | 39232990
106358 | Sterile indicator tape, hot air oven | TTM | 1 | Drit-heat indicator tape, specially designed for dry heat sterilisation, colour change from green to brown, roll of 50m, 19 mm width | N/A | 3 weeks | General Cargo |  |  | 0.1 | 12.5 | 12.5 | 2 | 0.0003125 | 0 | 13.32 | USD | 0 | Vereinigte Papierwarenfabrik 3FKLB410112 | SDS |  | 39239000
106359 | Sterile indicator tape, autoclave | TTM | 1 | Serilisation adhesive tape, The colour of the sensitive ink changes after sterilisation. The adhesive tape indicates that the expected temperature was reached during the sterilisation process, roll of 50m, 19mm width | N/A | 3 weeks | General Cargo | This is a roll of 50 meters. |  | 0.1 | 13 | 13 | 3 | 0.000507 | 0 | 2.13 | USD | 0 | Vereinigte Papierwarenfabrik 3FKLB410102 | SDS |  | 39239000
106409 | Waste containers for sharps | TTM | 10 | Sekuroka®-safety bin, standard, capacity 2.1lt, in box of 10 pcs | N/A | 3 weeks | General Cargo |  |  | 0.1 | 15.5 | 15.5 | 19.5 | 0.0046849 | 0 | 16.11 | USD | 0 | Sarstedt 77.3894.021 | CE |  | 39269097
21. Biosafety cabinets (BSC)
106076 | Biosafety Cabinet - 3 filters, 5 feet | TTM | 1 | ScanLaf, Mars Pro 1500, class II Biological safety Cabinet, EN 123980, with 3 Hepa filter H14, 99,999% efficient against 0.3µm particles, with motorised front window, digital control level, noise level 56dB, Integrated UV Light, USB port and gas tap | N/A | 6 weeks | General Cargo | · GDF recommends to provide the labs with 3-filter-BSC’s (Mars 1500 Pro); · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50 Hz or 110V/ 60 Hz; · Product brochure available here: https://www.labogene.com/Class-2-Cabinets--Mars-Pro |  | 275 | 80 | 180 | 212 | 3.0528 | 0 | 18951.5469 | USD | 0 | Mars 1500 Pro
106077 | Base Support for BSC - 5 feet | TTM | 1 | Base Support frame for BSC Mars 1500 | N/A | 6 weeks | General Cargo | Not needed for a 3-filter-BSC (included in article 106076) |  | 70 | 150 | 92 | 90 | 1.242 | 0 | 971.4062 | USD | 0 | Labogene
106078 | Thimble for Class IIA, hood for BSC | TTM | 1 | Ducting kit with anti-blow back valve | N/A | 6 weeks | General Cargo | · Please inform if the vertical or horizontal version is needed; · GDF recommends the vertical version - minimal room height: 250 cm) |  | 50 | 140 | 110 | 60 | 0.924 | 0 | 1989.1872 | USD | 0 | Labogene
106471 | Fromaldehyde/neutralization fumigation top plate | TTM | 1 | Formaldehyde/ neutralization fumigation top plate | N/A | 6 weeks | General Cargo | Not needed for a 3-Filter-BSC. But in some countries it is mandatory. |  | 22 | 160 | 68 | 30 | 0.3264 | 0 | 1143.9142 | USD | 0 | Labogene
106472 | HVP fumigation plate | TTM | 1 | HVP fumigation plate | N/A | 6 weeks | General Cargo | Not needed for a 3-Filter-BSC. But in some countries it is mandatory. |  | 18 | 160 | 68 | 30 | 0.3264 | 0 | 931.9381 | USD | 0 | Labogene
106469 | Double Exhaust filter | TTM | 1 | Double HEPA Exhaust Filter and frame for MARS 900/1200 - dimensions 610 x 457 x 115 mm Double HEPA Exhaust Filter and frame for MARS 1500/1800 - dimensions 915 x 457 x 115 mm | N/A | 6 weeks | General Cargo | Not needed for a 3-Filter-BSC. |  | 15 | 160 | 68 | 25 | 0.272 | 0 | 1989.1872 | USD | 0 | Labogene
106285 | Exhaust HEPA Air Filter | TTM | 1 | Exhaust HEPA filter for Mars 1200 Exhaust HEPA filter for Mars 1500/1800 | N/A | 6 weeks | General Cargo | GDF recommends that these filters are always available for the replacement of non-functioning filters (1 x per 1 BSC) |  | 8 | 160 | 68 | 15 | 0.1632 | 0 | 633.4614 | USD | 0 | Labogene
106047 | Main HEPA Filter for air down flow | TTM | 1 | Main HEPA filter for Mars/Fortuna 1200 Main HEPA filter for Mars/Fortuna 1500 Main HEPA filter for Mars/Fortuna 1800 | N/A | 6 weeks | General Cargo | GDF recommends that these filters are always available for the replacement of non-functioning filters (1 x per 1 BSC) |  | 15 | 160 | 68 | 20 | 0.2176 | 0 | 1266.9228 | USD | 0 | Labogene
106123 | Pre HEPA filter for down flow and in take air | TTM | 1 | V-shaped H-14 prefilters for Mars Pro 1200 - To order Qty : 12 V-shaped H-14 prefilters for Mars Pro 1500 - To order Qty : 15 V-shaped H-14 prefilters for Mars Pro 1800 - To order Qty : 18 | N/A | 6 weeks | General Cargo | GDF recommends that these filters are always available for the replacement of non-functioning filters (15 x per 1 BSC) |  | 5 | 39 | 39 | 28 | 0.042588 | 0 | 170.5346 | USD | 0 | Labogene
106489 | Installation and Training - BSC | TTM | 1 | Installation of BSC in the laboratory. If necessary duct the BSC outside with a Thimble construction Certification of BSC after installation Training on use and troubleshooting of BSC and if necessary on certification | N/A | On demand | Service | For each BSC please calculate 5 “units” for the installation, as 5 days is estimated for the completion of the installation of one BSC. E.g: for 1 BSC= 5 installations. Final invoice will be on actuals |  | N/A | N/A | N/A | N/A | N/A | N/A | 900 | USD | 0 | N/A | N/A |  | N/A
106076 | Biosafety Cabinet - 2 filters, 5 feet | TTM | 1 | ScanLaf, Mars 1500, class II Biological safety Cabinet, EN 123980, with 3 Hepa filter H14, 99,999% efficient against 0.3µm particles, with motorised front window, digital control level, noise level 56dB, Integrated UV Light, USB port and gas tap | N/A | 6 weeks | General Cargo | · 2 years warranty; · Product brochure available here: https://www.labogene.com/Class-2-Cabinets--Mars |  | 275 | 80 | 180 | 212 | 3.0528 | 0 | 16417.04 | USD | 0 | Mars 1500
106077 | Base Support for BSC - 5 feet | TTM | 1 | Base Support frame for BSC Mars 1500 | N/A | 6 weeks | General Cargo | Needed if you do not have laboratory benches or the like where you can put the BSC on. |  | 70 | 150 | 92 | 90 | 1.242 | 0 | 971.4062 | USD | 0 | Labogene
106078 | Thimble for Class IIA, hood for BSC | TTM | 1 | Ducting kit with anti blow back valve | N/A | 6 weeks | General Cargo | · Please inform if the vertical or horizontal version is needed; · GDF recommends the vertical version - minimal room height: 250 cm) |  | 50 | 140 | 110 | 60 | 0.924 | 0 | 1989.1872 | USD | 0 | Labogene
106471 | Fromaldehyde/neutralization fumigation top plate | TTM | 1 | Formaldehyde/ neutralization fumigation top plate | N/A | 6 weeks | General Cargo | Recommended to order with this BSC |  | 22 | 160 | 68 | 30 | 0.3264 | 0 | 1143.9142 | USD | 0 | Labogene
106472 | HVP fumigation plate | TTM | 1 | HVP fumigation plate | N/A | 6 weeks | General Cargo | Optional item, in addition to item 106471. |  | 18 | 160 | 68 | 30 | 0.3264 | 0 | 931.9381 | USD | 0 | Labogene
106469 | Double Exhaust filter | TTM | 1 | Double HEPA Exhaust Filter and frame for MARS 900/1200 - dimensions 610 x 457 x 115 mm Double HEPA Exhaust Filter and frame for MARS 1500/1800 - dimensions 915 x 457 x 115 mm | N/A | 6 weeks | General Cargo | If you need a 3-filter-BSC but do not have the space for a 3-filter-BSC, you can use this to “make” a 3-filter-BSC of a 2-filter-BSC. |  | 15 | 160 | 68 | 25 | 0.272 | 0 | 1989.1872 | USD | 0 | Labogene
106285 | Exhaust HEPA Air Filter | TTM | 1 | Exhaust HEPA filter for Mars 1200 Exhaust HEPA filter for Mars 1500/1800 | N/A | 6 weeks | General Cargo | GDF recommends that these filters are always available for the replacement of non-functioning filters (1 x per 1 BSC) |  | 8 | 160 | 68 | 15 | 0.1632 | 0 | 633.4614 | USD | 0 | Labogene
106047 | Main HEPA Filter for air down flow | TTM | 1 | Main HEPA filter for Mars/Fortuna 1200 Main HEPA filter for Mars/Fortuna 1500 Main HEPA filter for Mars/Fortuna 1800 | N/A | 6 weeks | General Cargo | GDF recommends that these filters are always available for the replacement of non-functioning filters (1 x per 1 BSC) |  | 15 | 160 | 68 | 20 | 0.2176 | 0 | 1266.9228 | USD | 0 | Labogene
106123 | Pre HEPA filter for down flow and in take air | TTM | 1 | V-shaped H-14 prefilters for Mars Pro 1200 - To order Qty : 12 V-shaped H-14 prefilters for Mars Pro 1500 - To order Qty : 15 V-shaped H-14 prefilters for Mars Pro 1800 - To order Qty : 18 | N/A | 6 weeks | General Cargo | Not needed for a 2-filter BSC |  | 5 | 39 | 39 | 28 | 0.042588 | 0 | 170.5346 | USD | 0 | Labogene
106489 | Installation and Training - BSC | TTM | 1 | Installation of BSC in the laboratory. If necessary duct the BSC outside with a Thimble construction Certification of BSC after installation Training on use and troubleshooting of BSC and if necessary on certification | N/A | On demand | Service | For each BSC please calculate 5 “units” for the installation, as 5 days is estimated for the completion of the installation of one BSC. E.g: for 1 BSC= 5 installations. Final invoice will be on actuals |  | N/A | N/A | N/A | N/A | N/A | N/A | 900 | USD | 0 | N/A | N/A |  | N/A
22. Biosafety: Consumables
106079 | Thermic anemometer with telescope probe | TTM | 1 | Anemometer TSI 9565X and telescopic probe 1962 - to measure velocity of air or gas flow and temperature, with memory function and one test probe | N/A | 3 weeks | General Cargo | · Product linked to 106076; · 2 years warranty; · Voltage / Amp / Frequency: battery 1.5 V |  | 4.5 | 50 | 30 | 20 | 0.03 | 0 | 1993.15 | USD | 0 | Driesen& Kern TSI 9565X & TSI962 & Tasche | CE |  | 90258040
106080 | UV lamps, 40 W, 90 cm, incl. holder | TTM | 1 | UVP, UV lamp 40 W , 365Nm, strip light unit, with electronical ballast, incl. 40W, UVC power 11W ozon free, | N/A | 4 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 1.8 | 47 | 15 | 12 | 0.00846 | 0 | 272.98 | USD | 0
106082 | Swivel stool/laboratory chair | TTM | 1 | Swivel stool (laboratory chair), vertical adjustable 420 - 620 mm, seat and back made of PU foam, disinfect able, 5 wheels, without armrests | N/A | 3 weeks | General Cargo |  |  | 15.6 | 68 | 65 | 31 | 0.13702 | 0 | 195.03 | USD | 0 | VWR Collection 132-0406 | CE |  | 94013000
106083 | Mobile container | TTM | 1 | Mobile container, lockable, on 4 castors, disinfect able, worktop 30 mm thick, Melamin coated, high chemical resistance, temperature resistance to 90 °C, with edges of the worktop lipped, four drawers, smooth action and metal rails with anti bocking system for total extraction, h x w x d: 730 x 436 x 590 mm | N/A | 3 weeks | General Cargo |  |  | 30 | 60 | 60 | 50 | 0.18 | 0 | 583.79 | USD | 0 | Caspar & Labora US-1x20-f-s-CC30-Sonder | CE |  | 94029000
106608 | Trolley | TTM | 1 | Trolley, 18/0 stainless steel, tubular frame: 25 mm ø, thick shelves: 0,8 mm, with 76 x 44 cm surfaces, outer dimension: 83 x 51 cm, four 125 mm ø swivel wheels (2 with brakes), 40 kg maximum load per shelf. | N/A | 3 - 4 weeks | General Cargo |  |  | 12.6 | 100 | 60 | 45 | 0.27 | 0 | 144.79 | USD | 0 | Contacto 7004/003
106089 | Plastic-foil welder 30 cm | TTM | 1 | Plastic welder, welding rod 300mm, 500W, 230v 50Hz | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: 230V/ 50/60Hz, 110V 60Hz |  | 5 | 50 | 20 | 15 | 0.015 | 0 | 205.72 | USD | 0 | Burghardt ME300HI | CE |  | 39269097
106362 | Plastic-foil | TTM | 1 | Plastic foile, PE, roll of 100m, 300x0.2mm | N/A | 3 weeks | General Cargo |  |  | 9.5 | 23.5 | 23.5 | 32 | 0.017672 | 0 | 68.9 | USD | 0 | Losch, Rainer 129-2134 VWR | CE |  | 39269097
106090 | Dispenser for disinfectant/soap | TTM | 1 | TTM soap and disinfectant dispenser , wall mounted, for 1 lt bottles, | N/A | 3 weeks | General Cargo |  |  | 1 | 33 | 24 | 12 | 0.009504 | 0 | 55.09 | USD | 0 | TEMDEX 119611 | CE |  | 39269097
106091 | Paper towel box | TTM | 1 | TTM Wall mounted towel holder for minimum 200 towels | N/A | 3 weeks | General Cargo |  |  | 2.2 | 37 | 28 | 14 | 0.014504 | 0 | 21.29 | USD | 0 | 4178 Scholl | CE |  | 48030090
106092 | Goggles | TTM | 1 | UVEX goggle, EN166and 170 | N/A | 3 weeks | General Cargo |  |  | 0.01 | 16.5 | 7 | 8 | 0.000924 | 0 | 11.34 | USD | 0 | UVEX 9166.065 | CE |  | 39269097
106094 | First Aid Kit for scientific laboratories | TTM | 1 | First Aid Kit for scientific laboratories, content according DIN 13169, with additional eye rinsing bottle, wall mounted, lockable cupboard, door with sealing gasket K932.1 | N/A | 3 weeks | General Cargo |  |  | 0.2 | 27 | 12 | 7 | 0.002268 | 0 | 126.61 | USD | 0 | Söhngen, Bartels + Rieger GmbH Paris, Barikos KS (954004) | CE |  | 30059099
106095 | Stainless steel discard funnel | TTM | 1 | Contacto funnel stainless steel, diameter of ramp 22 mm, top diameter 150 mm, total height approx. 170 mm | N/A | 3 weeks | General Cargo |  |  | 0.8 | 17.5 | 17.5 | 18.5 | 0.0056656 | 0 | 16.87 | USD | 0 | Contacto 5420/150 | CE |  | 90189084
106088 | Electronic maxima-minima-thermometer | TTM | 1 | Electronic maxima-minima thermometer, measuring indoor and outdoor temperatures, memory function for min/max values, Out door temperature alarm with adjustable upper and lower limit values, Outer probe with 1.5m lead, measuring range, indoor -10 to + 50°C, outdoor -50 to +70°C, resolution 0.1°C, accuracy ±1°C (0 to + 50°C), ±2°C (rest measuring range), 110 x 70x 20mm, inclusive 1 spare battery | N/A | 3 weeks | General Cargo | · 2 years warranty; · Voltage / Amp / Frequency: battery 1.5 V |  | 0.2 | 12 | 9 | 3 | 0.000324 | 0 | 12.25 | USD | 0 | Carl Roth AN34.1 | CE |  | 90251920
106442 | Deep-freeze thermometer | TTM | 1 | Carl roth deep freeze thermometer | N/A | 3 weeks | General Cargo |  |  | 0.5 | 10 | 10 | 5 | 0.0005 | 0 | 5.18 | USD | 0 | Carl Roth H717.1 | CE |  | 90251180
106494 | Calibrated thermometer | TTM | 1 | Measuring range -50.0 to +300.0 ° C Accuracy ±0.2 °C (-50 to +200 °C) ±0.3 °C (rest meas. range) Resolution 0.1 °C Working temperature -20 to +50 °C Dimensions (W x D x H) 64 x 40 x 182 mm Weight 171 g Delivery incl. thermometer testo 112 incl. Pt 100 Immersion-/Penetration probe (L 115 mm, Ø 5 mm) with attached cable (cable length approx. 1.2 m), Certificate of Calibration and E-Block battery 9 V. | N/A | 3 weeks | General Cargo |  |  | 0.5 | 7 | 4 | 19 | 0.000532 | 0 | 328.07 | USD | 0 | Dostmann 5000-0700, 6000-1006 | CE |  | 90251920
106084 | Safety-Pipetting Ball | TTM | 1 | Pileusball, STD | N/A | 3 weeks | General Cargo |  |  | 0.8 | 20 | 20 | 20 | 0.008 | 0 | 2.94 | USD | 0 | Deutsch & Neumann 470 0002 | CE |  | 40149000
106086 | Set of mechanical pipetting aids | TTM | 1 | Pipump, pipette filler set : 1132005, size 0 (yellow 0-0.02ml), 1132006 size 2 (blue 0-2ml), 1132007, size 10 (green, 0-10ml) | N/A | 3 weeks | General Cargo |  |  | 0.5 | 18 | 6 | 8 | 0.000864 | 0 | 34.4 | USD | 0 | Herenz 1132005, 1132006, 1132007 | CE |  | 39269097
106087 | Transport box for cultures with BioPack insert | TTM | 1 | Alex Breuer GmbH BIO bottle transport box for cultures with BioPack insert, IATA approved, 2 lt, height inside 210mm, including label and absorbate material | N/A | 3 weeks | General Cargo |  |  | 0.7 | 13 | 13 | 25 | 0.004225 | 0 | 22.23 | USD | 0 | Alex Breuer 90030000.01 | CE |  | 48191000
106294 | Disinfectant for cleaning instrument | TTM | 2 L | Bomix® plus 2lt, Aldehyde-free instrument decontaminant | 35 months | 3 weeks | Dangerous goods |  |  | 2.3 | 19 | 14.5 | 28 | 0.007714 | 0 | 47.44 | USD | 0 | Paul Hartmann 9746092 | SDS | UN 1903 Class 8 Packing group III | 38089490
106304 | Hand pump for Disinfectant for cleaning instrument | TTM | 1 | Hand pump for cleaning instrument 10ml dosage | N/A | 3 weeks | General Cargo |  |  | 0.6 | 12.5 | 7.5 | 34 | 0.0031875 | 0 | 7.02 | USD | 0 | Paul Hartmann 8126602 | N/A |  | 84135040
106370 | Disinfectant for floors | TTM | 1 | Medizid AF, disinfectant for floor, 2lt jerry can | N/A | 3 weeks | Dangerous goods |  |  | 5.2 | 19 | 15 | 25.5 | 0.0072675 | 0 | 13.79 | USD | 0 | Meditrade Medizid AF | SDS | UN 1903 Class 8 Packing group III | 38085000
106371 | Disinfectant for BSCs surface | TTM | 5 L | Bacillol AF, alcoholic fast disinfectant, 5lt jerry can | N/A | 3 weeks | Dangerous goods |  |  | 5.1 | 21 | 29 | 32 | 0.019488 | 0 | 38.11 | USD | 0 | Paul Hartmann9733894 | SDS | UN 1987 Class 3 Packing group III | 38085000
106372 | Spray head and 1L bottles | TTM | 1 set | Spray head with 2 x 1lt bottle | N/A | 3 weeks | General Cargo |  |  | 0.2 | 20 | 12 | 25 | 0.006 | 0 | 5.32 | USD | 0 | Mraz 2 x 1100/1000, 1 x 1200/28S | N/A |  | 84135040
106373 | Liquid soap | TTM | 1 L | Premium liquid soap, 1 lt bottle | N/A | 3 weeks | General Cargo |  |  | 1.1 | 9 | 9 | 22 | 0.001782 | 0 | 3.14 | USD | 0 | Paul Hartmann 9813291 | SDS |  | 34012090
106374 | Disinfectant for hands | TTM | 1 L | Disinfectant sterilium classic, 1lt bottle | 35 months | 3 weeks | Dangerous goods |  |  | 1.1 | 9 | 9 | 21 | 0.00174 | 0 | 9.09 | USD | 0 | Hartmann 975123 | SDS | UN 1987 Class 3 Packing group III | 34012090
106375 | Single-use paper towels | TTM | 5000 | Paper towels, single use 5000 per box, folded for wall dispenser | N/A | 3 weeks | General Cargo |  |  | 12 | 61 | 40 | 26 | 0.06344 | 0 | 15.21 | USD | 0 | Fripa 401 1103 | CE |  | 73249000
106378 | Detergent/washing Powder | TTM | 10 kg | Detergent kit DECONEX 22 HPF/ average for 3 month | 24 months | 3 weeks | Dangerous goods |  |  | 10 | 15.5 | 12 | 21 | 0.003906 | 0 | 160.96 | USD | 0 | Borer Chemie / Deconex, Rixius (Kanister) 516-900 VWR | SDS | UN 3266 Class 8 Packing group II | 34029090
106345 | Latex gloves size S | TTM | 1000 | Noba medical examination gloves, latex, powder free, non sterile, packaging 10x 100 pcs | 35 months | 3 weeks | General Cargo |  |  | 7 | 42 | 27 | 26 | 0.029484 | 0 | 50.32 | USD | 0 | NOBA 905451 | CE |  | 40151900
106346 | Latex gloves size M | TTM | 1000 | Noba medical examination gloves, latex, powder free, non sterile, packaging 10x 100 pcs | 35 months | 3 weeks | General Cargo |  |  | 8 | 42 | 27 | 26 | 0.029484 | 0 | 50.32 | USD | 0 | NOBA 905452 | CE |  | 40151900
106347 | Latex gloves size L | TTM | 1000 | Noba medical examination gloves, latex, powder free, non sterile, packaging 10x 100 pcs | 35 months | 3 weeks | General Cargo |  |  | 9 | 42 | 27 | 26 | 0.029484 | 0 | 50.32 | USD | 0 | NOBA 905453 | CE |  | 40151900
106551 | Laboratory coat size L - Disposable - Sterile | TTM | 50 | Laboratory coat made of Polypropylene- single use, white, 50 g/m², size L, packaging in box of 50 | N/A | 3 weeks | General Cargo |  |  | 5 | 37 | 28 | 34 | 0.035224 | 0 | 46.94 | USD | 0 | Wiros 25001305 | CE |  | 62113210
106552 | Laboratory coat size M - Disposable - Sterile | TTM | 50 | Laboratory coat made of Polypropylene- single use, white, 50 g/m², size M, packaging in box of 50 | N/A | 3 weeks | General Cargo |  |  | 5 | 37 | 28 | 35 | 0.03626 | 0 | 46.94 | USD | 0 | Wiros 25001304 | CE |  | 62113210
106553 | Laboratory coat size S - Disposable - Sterile | TTM | 50 | Laboratory coat made of Polypropylene- single use, white, 50 g/m², size S, packaging in box of 50 | N/A | 3 weeks | General Cargo |  |  | 5 | 33 | 28 | 40 | 0.03696 | 0 | 46.94 | USD | 0 | Wiros 25001303 | CE |  | 62113210
106503 | Laboratory coat cotton - size S | TTM | 1 | Laboratory coat with cuffs, 100 % Cotton, twill 3/1, white, 205 g/m², Rivet closing at front, 3 pockets | N/A | 3 weeks | General Cargo |  |  | 0.75 | 35 | 37 | 2.5 | 0.0032375 | 0 | 19.07 | USD | 0 | Pikkemaat 12-102-S | CE |  | 62114210
106504 | Laboratory coat cotton - size M | TTM | 1 | Laboratory coat with cuffs, 100 % Cotton, twill 3/1, white, 205 g/m², Rivet closing at front, 3 pockets | N/A | 3 weeks | General Cargo |  |  | 0.75 | 35 | 37 | 2.5 | 0.0032375 | 0 | 19.07 | USD | 0 | Pikkemaat 12-102-M | CE |  | 62114210
106505 | Laboratory coat cotton - size L | TTM | 1 | Laboratory coat with cuffs, 100 % Cotton, twill 3/1, white, 205 g/m², Rivet closing at front, 3 pockets | N/A | 3 weeks | General Cargo |  |  | 0.75 | 35 | 37 | 2.5 | 0.0032375 | 0 | 19.07 | USD | 0 | Pikkemaat 12-102-L | CE |  | 62114210
106369 | Masks FFP2 (N95) - 3M 9320 | TTM | 10 | 3M, foldable mask, FFP2 protection level, STD: EN 149:2001FFP2, CE Marked, Assigned Protection Factor: 10xWEL (Nominal Protection Factor 12x WEL) | N/A | 3 weeks | General Cargo |  |  | 0.17 | 21 | 12 | 8.5 | 0.00214 | 0 | 20.06 | USD | 0 | 9320+ |  |  | 63079009
106376 | Cotton wool | TTM | 1 | Nobavit, Cotton wool, 100% cotton, 1kg 40cm width | N/A | 3 weeks | General Cargo |  |  | 1.2 | 42 | 42 | 18 | 0.031752 | 0 | 5.49 | USD | 0 | NOBA 110061 | CE |  | 30059010
106377 | Tissue pulp | TTM | 1 | Bench Surface protection, High-quality, smooth, absorbent paper that quickly absorbs liquid spills with a thick, laminated polyethylene layer that prevents seeping onto the work surface, 570x450mm, 50pcs per box | N/A | 3 weeks | General Cargo |  |  | 2.6 | 67 | 47 | 3 | 0.009447 | 0 | 29.6 | USD | 0 | VWR Collection 115-9220 VWR | CE |  | 48115900
106402 | Plastic bottle | TTM | 1 | Plastic bottle PE-LD with lid, natural colour/ 500ml | N/A | 3 weeks | General Cargo |  |  | 0.05 | 8 | 8 | 18 | 0.001152 | 0 | 0.72 | USD | 0 | Mraz 1100/500 + 1200/25 | CE |  | 39269097
106497 | Nitril gloves, size 6-7 - Disposable | TTM | 100 | Disposable nitril gloves, light blue powder free, with structured finger tips length 300mm, strength 0.10 mm, size 6-7 , in spender box, 100 pcs per box | N/A | 3 weeks | General Cargo |  |  | 0.9 | 28 | 12 | 8 | 0.002688 | 0 | 15.92 | USD | 0 | VWR Collection 112-2371 VWR | CE |  | 40151900
106498 | Nitril gloves, size 7-8 - Disposable | TTM | 100 | Disposable nitril gloves, light blue powder free, with structured finger tips length 300mm, strength 0.10 mm, size 7-8 , in spender box, 100 pcs per box | N/A | 3 weeks | General Cargo |  |  | 0.9 | 28 | 12 | 8 | 0.002688 | 0 | 15.92 | USD | 0 | VWR Collection 112-2372 VWR | CE |  | 40151900
106499 | Nitril gloves, size 8-9 - Disposable | TTM | 100 | Disposable nitril gloves, light blue powder free, with structured finger tips length 300mm, strength 0.10 mm, size 8-9 , in spender box, 100 pcs per box | N/A | 3 weeks | General Cargo |  |  | 0.9 | 28 | 12 | 8 | 0.002688 | 0 | 15.92 | USD | 0 | VWR Collection 112-2373 VWR | CE |  | 40151900
106403 | Nitril gloves, size 6-7 - Non-disposable | TTM | 12 | Gloves SOLVEX NITRIL green, 0.38mm G7, category III, packaging 12 pcs per box | 35 months | 3 weeks | General Cargo |  |  | 0.7 | 38 | 18 | 9 | 0.006156 | 0 | 24.27 | USD | 0 | Ansell 37-675-070 | CE |  | 40151900
106404 | Nitril gloves, size 7-8 - Non-disposable | TTM | 12 | Gloves SOLVEX NITRIL green, 0.38mm G8, category III, packaging 12 pcs per box | 35 months | 3 weeks | General Cargo |  |  | 0.7 | 38 | 18 | 9 | 0.006156 | 0 | 24.27 | USD | 0 | Ansell 37-675-080 | CE |  | 40151900
106405 | Nitril gloves, size 8-9 - Non-disposable | TTM | 12 | Gloves SOLVEX NITRIL green, 0.38mm G9, category III, packaging 12 pcs per box | 35 months | 3 weeks | General Cargo |  |  | 0.7 | 38 | 18 | 9 | 0.006156 | 0 | 24.27 | USD | 0 | Ansell 37-675-080 | CE |  | 40151900
106435 | Surgical gowns non sterile - Size S | TTM | 20 | Lab coat NOBADRESS single usage, material fibrous web, sterile, single packed, blue, packaging of 20 pcs per box | 35 months | 3 weeks | General Cargo |  |  | 5 | 40 | 30 | 39 | 0.0468 | 0 | 57.84 | USD | 0 | NOBA 621120 | CE |  | 48185000
106436 | Surgical gowns non sterile - Size M | TTM | 20 | Lab coat NOBADRESS single usage, material fibrous web, sterile, single packed, blue, packaging of 20 pcs per box | 35 months | 3 weeks | General Cargo |  |  | 5 | 40 | 30 | 39 | 0.0468 | 0 | 59.46 | USD | 0 | NOBA 621130 | CE |  | 48185000
106437 | Surgical gowns non sterile - Size L | TTM | 20 | Lab coat NOBADRESS single usage, material fibrous web, sterile, single packed, blue, packaging of 20 pcs per box | 35 months | 3 weeks | General Cargo |  |  | 5 | 40 | 30 | 39 | 0.0468 | 0 | 60.75 | USD | 0 | NOBA 621150 | CE |  | 48185000
106500 | Surgical gown - Cotton - S | TTM | 1 | Surgical gown with cuffs, 100% Cotton, plain weave, white, 175 g/m², Closing with binding bands out of main fabric | N/A | 3 weeks | General Cargo |  |  | 0.75 | 35 | 37 | 2.5 | 0.0032375 | 0 | 17.84 | USD | 0 | Pikkemaat 11-102-S | CE |  | 62114210
106501 | Surgical gown - Cotton - M | TTM | 1 | Surgical gown with cuffs, 100% Cotton, plain weave, white, 175 g/m², Closing with binding bands out of main fabric | N/A | 3 weeks | General Cargo |  |  | 0.75 | 35 | 37 | 2.5 | 0.0032375 | 0 | 17.84 | USD | 0 | Pikkemaat 11-102-M | CE |  | 62114210
106502 | Surgical gown - Cotton - L | TTM | 1 | Surgical gown with cuffs, 100% Cotton, plain weave, white, 175 g/m², Closing with binding bands out of main fabric | N/A | 3 weeks | General Cargo |  |  | 0.75 | 35 | 37 | 2.5 | 0.0032375 | 0 | 17.84 | USD | 0 | Pikkemaat 11-102-L | CE |  | 62114210
106440 | Shoe covers | TTM | 1000 | PE shoe cover 21x40, blue, 21cm high, with rubber band, extra strong, rhombus foil with ABS structure, box of 1000 pcs packed in bags of 100pcs | N/A | 3 weeks | General Cargo |  |  | 6.3 | 39 | 28 | 39 | 0.042588 | 0 | 22.85 | USD | 0 | Wiros 50190122 | CE |  | 39262000
106441 | Hair covers | TTM | 2000 | Hair cover, PP barrett type, white Ø 48 cm, packaging 1000 pcs per bag in box of 2000 pcs | N/A | 3 weeks | General Cargo |  |  | 4.6 | 34 | 28 | 50 | 0.0476 | 0 | 22.2 | USD | 0 | Wiros 10000102 | CE |  | 39262000
106349 | Tri-pod stand for waste bags of ca. 2 litres | TTM | 1 | Tri-pod stand for waste bags of 2 litres capacity | N/A | 3 weeks | General Cargo | This item needs to be ordered with item 106399 |  | 0.3 | 20 | 20 | 25 | 0.01 | 0 | 10.87 | USD | 0 | Ratiolab 7001400 | CE |  | 73269028
106399 | Plastic bags made from PP | TTM | 100 | Disposable bag, autoclavable up to 121°C, 200x300mm, 40µm, box of 100pcs | N/A | 3 weeks | General Cargo | Product linked to 106349 |  | 0.5 | 23 | 12 | 9 | 0.002484 | 0 | 7.89 | USD | 0 | Ratiolab 7001000 | CE |  | 39232990
106444 | Plastic bags - 30L | TTM | 200 | Bags, Biohazard 27lt, with temperature indicator, text is in white before autoclaving, than black. The text 'autoclaved' indicates successful autoclaving, with highly visible label 'Biohazard'; bags meet ASTM D 709-98 Standard; packaging 200 pcs per box | N/A | 3 weeks | General Cargo | Product linked to 106445 |  | 5.7 | 54 | 35 | 11 | 0.02079 | 0 | 79.1 | USD | 0 | VWR Collection 129-0025 VWR | CE |  | 39232100
106445 | Stand for 30L plastic bags | TTM | 1 | Clavies® Biohazard bag holder for 30lt bags | N/A | 3 weeks | General Cargo | Product linked to 106444 / 106348 |  | 2 | 82 | 37 | 6 | 0.018204 | 0 | 47.77 | USD | 0 | Kleinfeld Labortechnik F13192-0002 | CE |  | 39232990
106507 | Smock stick starter kit | TTM | 1 | Dräger, Smock sticks starter kit with rubber ball pump, rubber caps, 10 smoke tubes, tube opener | N/A | 3 weeks | General Cargo | This item needs to be ordered with item 106508 |  | 0.24 | 6.5 | 14.5 | 0.5 | 4.7125e-05 | 0 | 90.74 | USD | 0 | Dräger CH 00 216 | CE |  | 28070000
106508 | Refill: Smock stick pack - 10 pcs | TTM | 1 | Dräger, Refill pack for smock sticks 10 pcs | N/A | 3 weeks | Dangerous goods | Product linked to 106507 |  | 0.24 | 0.24 | 0.24 | 0.24 | 1.3824e-08 | 0 | 50.61 | USD | 0 | Dräger CH 25 301 | CE | UN 3260 Class 8 Packing group: II | 28070000

## Order Summary
| TB Diagnostics Catalogue Product Categories | TB Diagnostics Order Summary
1. GDF microscopy kits
2. LED / Light microscopy: Instruments, reagents and consumables |  | Supplier | Lead time EXW after payment | Type of cargo (volume in m3) | [5 empty] | Cost of services (maintenance, training, etc) | Cost EXW per supplier
3. Molecular testing: GeneXpert system |  |  |  | General Cargo | Dangerous goods | Keep cool 2-8°C | General Cargo Large/Dangerous | General Cargo Large item | General Cargo special provision A67
4. Molecular testing: Loop-mediated isothermal amplification (TB-LAMP) |  | BD |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5. Processing and transport: Equipment, reagents and consumables |  | Carl Zeiss |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
6. Solid culture: Instruments, reagents and consumables |  | Cepheid |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
7. Liquid culture: Instruments, reagents and consumables |  | HAIN |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
8. Solid and liquid culture DST: First-line drug pure substances |  | HUMAN |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
9. Solid and liquid culture DST: Second-line drug pure substances |  | MEG |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
10. Solid and liquid culture: Rapid tests for MTB identification |  | SD |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11. Solid and liquid culture: Chemicals for MTB identification |  | SVIZERA |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
12. Molecular testing: LPA equipment and reagents |  | TAUNS |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
13. Molecular testing: Consumables |  | TTM |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
14. Molecular testing: Chemicals for gel electrophoresis | [9 empty] | Total EXW cost | 0
15. Media kitchen: Equipment | [9 empty] | Total volume (m3) | 0
16. Media kitchen: Consumables (plastic and glassware)
17. Media kitchen: Other consumables
18. Autoclave basic unit for media kitchen
19. Autoclave basic unit for waste management
20. Waste management: Chemicals and consumables
21. Biosafety cabinets (BSC)
22. Biosafety: Consumables