## Introduction
Score-TB Package
Building quality-assured TB testing and management capacity using SLIPTA methodology
Version 2.1 - September 2020
Introduction
Despite the fact that laboratory results influence 70% of medical diagnoses, laboratory services in low- and middle-income countries (LMICs) have long been a neglected component of health care systems. TB laboratories, which are an essential component in all stages of the TB care cascade, are no exception. A key intervention to strengthen laboratory services is the implementation of a quality management system (QMS). A QMS is defined by the International Organization for Standardization (ISO) as the “management system to direct and control an organization with regard to quality”. Hence it is the system (“the set of interrelated or interacting elements”) aimed at implementing and operationalizing quality management in an organization. Standardization of testing through implementation of a QMS has been shown to improve the quality of testing by reducing testing errors.
Several tools and initiatives to assist laboratories implement quality improvement activities have been developed. One of the most successful approaches to QMS improvement is the Strengthening Laboratory Management toward Accreditation (SLMTA) approach, first described by Yao et al. The SLMTA approach is often used in conjunction with the Stepwise Laboratory Improvement Process Towards Accreditation (SLIPTA) checklist. The SLIPTA checklist was developed by WHO Regional Office for Africa (WHO-AFRO) and partners in 2010 in recognition of the gap between the current state of laboratory quality and the requirements of the ISO 15189:2007 standard. In 2015, the SLIPTA checklist was adapted to incorporate the requirements of the ISO 15189:2012 standard and became known as ‘SLIPTA v2:2015’.
The Foundation for Innovative New Diagnostics (FIND) has reported on the development of a TB laboratory-specific approach called TB Strengthening Laboratory Management toward Accreditation (TB SLMTA). The program is based on the existing successful SLMTA approach and utilized a revised checklist (TB Harmonized Checklist) based on SLIPTA, but incorporating some elements from the GLI tool with a focus on the technical side of TB laboratory testing.
In 2019, an additional technical revision was made to the TB SLMTA Harmonized Checklist to include TB testing methods not included in previous revisions. The current major revision concerns the incorporation of the TB SLMTA Harmonized Checklist into the ‘Score-TB Package’, which also includes an electronic version of the checklist referred to as the ‘e-tool’. This e-tool substantially increases user-friendliness and reduces the risk for errors by automating the calculation of assessment scores and presenting these in a reporting worksheet to visualize strengths and weaknesses of a laboratory’s QMS (the SLIPTA score) and TB testing methods.
Target audience
The Score-TB Package is intended to inform Ministries of Health officials, health facility- and laboratory managers, donors, implementing partners, quality assurance personnel, program managers and supervisory staff at national, regional and facility level on requirements for delivering quality-assured laboratory testing for TB and ensuring effective use of laboratory resources as well as data for patient management and surveillance in LMIC.
Acknowledgements
The development of the Score-TB Package was led by FIND with contributions from Andre Trollip and Heidi Albert (FIND) and Tjeerd Datema and Linda Oskam (DATOS). The TB Lab Quality Scorecards draw from a number of existing tools, including CDC Laboratory Assessment of Antimicrobial Resistance Testing Capacity checklist, India Council for Medical Research (ICMR) TB Checklist, the WHO-AFRO SLIPTA checklist and FIND’s TB Laboratory Quality Management Systems Towards Accreditation Harmonized Checklist.

## Instructions
Score TB Package v2.1 - September 2020
Instructions
The eTool is divided into the following color-coded TABS:
Introduction & Instructions sections (grey)
This and the previous tab.
Set Audit Scope
On this tab the assessor needs to select which tests are performed by the laboratory. Links are provided to the scorecards that need to be completed during the assessment based on the answers. The General TB Module needs to be completed for all laboratories, regardless of the test menu.
TB sections (green)
All TB scorecards (green) include a general information section for collection of quantitative data and a technical information section. In the technical information section questions can be answered with a drop-down menu. Based on the answer selected, a score is automatically calculated for each question. Assessors can also provide additional information and comments in the comments-section of each question.
General TB Module
This module corresponds to the General Procedures scorecard. These questions need to be answered for all laboratories as they are related to general procedures that need to be in place in all laboratories regardless of test menu.
Smear
This module corresponds to the Smear Microscopy tuberculosis technical scorecard and is compulsory for all laboratories conducting smear microscopy. Please refer to the Score TB Package User Guide for more details on completing this section.
Culture
This module corresponds to the TB culture for detection and identification of mycobacteria technical scorecard and is compulsory for all laboratories performing TB culture. Please refer to the Score TB Package User Guide for more details on completing this section.
DST
This module corresponds to the Phenotypic Drug Susceptibility Testing (DST) technical scorecard and is compulsory for all laboratories performing DST. Please refer to the Score TB Package User Guide for more details on completing this section.
Xpert
This module corresponds to the Xpert MTB/RIF (includes Xpert MTB/RIF Ultra) tuberculosis technical scorecard and is compulsory for all laboratories conducting Xpert MTB/RIF and/or Xpert MTB/RIF Ultra testing. Please refer to the Score TB Package User Guide for more details on completing this section.
TB LAMP
This module corresponds to the Loop-Mediated Isothermal Amplification (TB-LAMP) tuberculosis technical scorecard and is compulsory for all laboratories conducting TB-LAMP testing. Please refer to the Score TB Package User Guide for more details on completing this section.
LF-LAM
This module corresponds to the Lateral Flow Urine Lipoarabinomannan Assay (LF-LAM) tuberculosis technical scorecard and is compulsory for all laboratories conducting LF-LAM testing. Please refer to the Score TB Package User Guide for more details on completing this section.
LPA
This module corresponds to the Line Probe Assay (includes MTBDRplus and MTBDRsl and CM) tuberculosis technical scorecard and is compulsory for all laboratories performing LPA. Please refer to the Score TB Package User Guide for more details on completing this section.
Truenat
This module corresponds to the Truenat (includes Truenat MTB, Truenat MTB Plus and Truenat MTB-Rif Dx) technical scorecard and is compulsory for all laboratories performing the Truenat assays. Please refer to the Score TB Package User Guide for more details on completing this section.
Previous Audit Information (yellow)
If available, previous audit information (for both the Score TB Package and the SLIPTA checklist) can be inserted on this TAB. Inserting data on this TAB enables comparison on the TB Summary Report and SLIPTA Summary Report TABs of the current assessment scores with scores of previous assessments and analyze progress.
Summary sections (blue)
These sections provide a calculated summary score for the TB and SLIPTA assessments. Scores are automatically updated when the TB (green) & SLIPTA (purple) sheets are completed. Because the TB scorecards are integrated in the SLIPTA checlist, the SLIPTA score automatically incorporates the score on the TB scorecards if they are completed. If not, only the score on the SLIPTA questions is provided. Scores of both the TB scorecards and SLIPTA are only displayed when the scorecards/checklists have been completed.
SLIPTA sections (purple)
To perform the SLIPTA assessment all the questions in all tabs should be completed:
- SLIPTA laboratory profile: This TAB corresponds to the "Part I: Laboratory Profile" of the SLIPTA checklist.
- SLIPTA_S1: This TAB corresponds to the "Part II: Laboratory Audits Section 1: Documents & Records" of the SLIPTA checklist.
- SLIPTA_S2: This TAB corresponds to the "Part II: Laboratory Audits Section 2: Management Reviews" of the SLIPTA checklist.
- SLIPTA_S3: This TAB corresponds to the "Part II: Laboratory Audits Section 3: Organization & Personnel" of the SLIPTA checklist.
- SLIPTA_S4: This TAB corresponds to the "Part II: Laboratory Audits Section 4: Client Management & Customer Service" of the SLIPTA checklist.
- SLIPTA_S5: This TAB corresponds to the "Part II: Laboratory Audits Section 5: Equipment" of the SLIPTA checklist.
- SLIPTA_S6: This TAB corresponds to the "Part II: Laboratory Audits Section 6: Evaluation and Audits" of the SLIPTA checklist.
- SLIPTA_S7: This TAB corresponds to the "Part II: Laboratory Audits Section 7: Purchasing & Inventory" of the SLIPTA checklist.
- SLIPTA_S8: This TAB corresponds to the "Part II: Laboratory Audits Section 8: Process Control" of the SLIPTA checklist.
- SLIPTA_S9: This TAB corresponds to the "Part II: Laboratory Audits Section 9: Information Management" of the SLIPTA checklist.
- SLIPTA_S10: This TAB corresponds to the "Part II: Laboratory Audits Section 10: Identification of Non Conformities, Corrective and Preventive Actions" of the SLIPTA checklist.
- SLIPTA_S11: This TAB corresponds to the "Part II: Laboratory Audits Section 11: Occurrence/Incident Management & Process Improvement" of the SLIPTA checklist
- SLIPTA_S12: This TAB corresponds to the "Part II: Laboratory Audits Section 12: Facilities and Biosafety" of the SLIPTA checklist.
Completing the eTool:
The suggested order for completing the eTool is shown below:

## Set Audit Scope
Score-TB Package v2.1 - September 2020
Set Audit Scope
Does the laboratory perform smear microscopy?
Does the laboratory perform TB culture for detection and identification of mycobacteria?
Does the laboratory perform drug susceptibility testing of M. tuberculosis complex?
Does the laboratory perform Xpert MTB/RIF or Xpert MTB/RIF Ultra testing?
Does the laboratory perform TB-LAMP testing?
Does the laboratory perform LF-LAM testing?
Does the laboratory perform LPA?
Does the laboratory perform Truenat testing?
Next step(s):
- Complete the General TB Module (click here)
Yes
No

## General TB Module
Score-TB Package v2.1 - September 2020
General TB Module
A. General Information
Name of Assessor(s)
Title & organization of Assessor
Name of laboratory being assessed
Type of laboratory | [7 empty] | Choose from drop-down menu -->
Level of laboratory | [7 empty] | Choose from drop-down menu -->
Does the microbiology laboratory meet minimum space and infrastructure requirements? | [7 empty] | Choose from drop-down menu --> |  |  |  | <------- | In addition to the GLI Laboratory Safety Handbook, refer to national/local regulations, if any.
How many hospitals and/or other health care facilities does the laboratory serve? | [12 empty] | Click here to go to the GLI Laboratory Safety Handbook
Location of laboratory being assessed (City/Town, County / District / Sub-district and Country)
Details of contact person at laboratory
Name
Position
Qualification
Email
Phone
Is there clinical microbiologist and / or pathologist with experience in microbiology on staff? | [7 empty] | Choose from drop-down menu -->
- If yes, how many years' experience do they have?
- If no, what is the highest qualified member of laboratory staff?
Is the laboratory accredited? | [7 empty] | Choose from drop-down menu -->
- If yes, name of accrediting body?
- If yes, what tests is the laboratory accredited for?
B. Technical Information
G-A | Is the following equipment available, and if so, is it functional, monitored, serviced and maintained?
|  |  |  | Available | Functional | Monitored | Serviced |  | Maintained | <------- | Functional: is the equipment in working order? Monitored: is the functionality of equipment regularly checked (e.g. temperature / calibrated)? Serviced: is the equipment regularly serviced by a qualified service technician? Maintained: is the equipment regularly maintained according to the manufacturer’s recommendations (e.g. cleaning)?
| Bright field microscope | [5 empty] | NA
| Fluorescence microscope | [5 empty] | NA
| Thermometers | [5 empty] | NA
| Incubator (aerobic) | [5 empty] | NA
| Incubator (CO2) | [5 empty] | NA
| Refrigerator (2-8°C) | [5 empty] | NA
| Centrifuge | [5 empty] | NA
| Vortex | [5 empty] | NA
| Shaker | [5 empty] | NA
| Hot air oven | [5 empty] | NA
| Inspissator | [5 empty] | NA
| Water bath | [5 empty] | NA
| pH meter | [5 empty] | NA
| Magnetic stirer | [5 empty] | NA
| Freezer (-20 - -80°C) | [5 empty] | NA
| Balance / scale | [5 empty] | NA
Autoclave
Incinerator
Micropipette (0.02 ml, 0.1 ml, 1 ml)
Biosafety Cabinet Class I
Biosafety Cabinet Class II
| Auxiliary/emergency power system or standby generator |  |  |  |  | NA | NA |  | NA
| Access to sink with running water |  |  |  |  | NA | NA |  | NA
Other equipment:
Specify:
Specify:
Specify:
Specify:
Specify:
Specify:
NA = Not Applicable
G-B | How does the laboratory report results?
Smear microscopy
TB culture
DST
Xpert MTB/RIF
TB-LAMP
LF-LAM
LPA
Truenat
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G1.1 |  | Are the following processes documented? | [6 empty] | 3 | Answer all sub questions |  | 0 |  |  | 1.5
|  |  | 1. Rejection criteria specific for TB samples and tests | [12 empty] | 100
|  |  | 2. A policy for reporting critical TB test results | [12 empty] | 100
|  |  | 3. Instructions for referral of samples | [12 empty] | 100
|  |  | 4. Instructions for extra-pulmonary sample criteria for collection and testing | [12 empty] | 100
|  |  | 5. Instructions for handling samples received after hours | [12 empty] | 100
|  |  | 6. Realistic turnaround time for TB tests From sample collection to reporting | [12 empty] | 100
|  |  | 7. Definition of result categories (rare/unexpected/discordant results) | [12 empty] | 100
|  |  | 8. Confirmatory tests for unusual, discordant or unexpected TB results | [12 empty] | 100 | 800
Sub total | [8 empty] | 3 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G2.1 |  | Does the laboratory report findings / trends and other related important information to NTRL / NTP regarding: | [6 empty] | 2 | Answer all sub questions |  | 0 |  |  | 2.2
|  |  | - Test quality indicators | [12 empty] | 100
|  |  | - Sample rejection rates | [12 empty] | 100
|  |  | - Other clinically relevant changes in the laboratory such as changes to testing information, PT results, and equipment downtime? | [12 empty] | 100 | 300
G2.2 |  | Does the laboratory report cumulative test quality indicators to NTRL/NTP at least annually? | [6 empty] | 2 |  |  | 0 |  |  | 2.2
Sub total | [8 empty] | 4 |  |  | 0
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 1. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G3.1 |  | Is there evidence that laboratory staff have been trained and correctly implement the following: Review training records, competency assessment forms and duty rosters. Pay attention to date of training and scope of training compared with techniques being performed. See section 8. | [6 empty] | 3 | Answer all sub questions |  | 0 |  |  | 3.6
|  |  | 1. Sample collection and transport | [12 empty] | 100
|  |  | 2. Sample accession and registration | [12 empty] | 100
|  |  | 3. Processing of pulmonary samples and conducting TB testing | [12 empty] | 100
|  |  | 4. Processing of extra-pulmonary samples and conducting TB testing | [12 empty] | 100
|  |  | 5. Interpretation of TB test results | [12 empty] | 100
|  |  | 6. Recording & reporting of TB test results | [12 empty] | 100
|  |  | 7. QC, EQA & PT for TB tests | [12 empty] | 100
|  |  | 8. Laboratory safety related to TB testing | [12 empty] | 100 | 800
G3.2 |  | Is there evidence that laboratory staff are following the procedures described in the laboratory documentation: Directly observe procedures being performed compared to the SOP. | [6 empty] | 3 | Answer all sub questions |  | 0 |  |  | 3.7
|  |  | 1. Sample collection and transport | [12 empty] | 100
|  |  | 2. Sample accession and registration | [12 empty] | 100
|  |  | 3. Processing of pulmonary samples and conducting TB testing | [12 empty] | 100
|  |  | 4. Processing of extra-pulmonary samples and conducting TB testing | [12 empty] | 100
|  |  | 5. Interpretation of TB test results | [12 empty] | 100
|  |  | 6. Reporting of TB test results | [12 empty] | 100
|  |  | 7. QC, EQA & PT for TB tests | [12 empty] | 100
|  |  | 8. Laboratory safety related to TB testing | [12 empty] | 100 | 800
Sub total | [8 empty] | 6 |  |  | 0
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G4.1 |  | Is there evidence that: | [6 empty] | 3 | Answer all sub questions |  | 0 |  |  | 4.1
|  |  | - The laboratory has provided clients information/instructions on sample collection, storage and transportation to the laboratory? | [12 empty] | 100
|  |  | - The laboratory enforces the criteria for sample volume, container type and storage conditions? | [12 empty] | 100 | 200
G4.2 |  | Does the laboratory provide feedback to clinicians regarding: | [6 empty] | 2 | Answer all sub questions |  | 0 |  |  | 4.3
|  |  | 1. Sample quality & rejection rates | [12 empty] | 100
|  |  | 2. Appropriateness of test requisition | [12 empty] | 100
|  |  | 3. Identity & frequency of isolated or identified MTB and NTMs | [12 empty] | 100
|  |  | 4. DST including MDR/XDR-TB | [12 empty] | 100 | 400
Sub total | [8 empty] | 5 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5.
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G5.1 |  | Does the laboratory use verified/validated methods for MTB detection and DST? Includes all conventional, automated and molecular (commercial & non-commercial) methods | [6 empty] | 5 |  |  | 0 |  |  | 5.3
G5.2 |  | Is all equipment installed and placed correctly in a suitable environment? | [6 empty] | 2 |  |  | 0 |  |  | 5.1
G5.3 |  | Does the laboratory maintain all equipment for TB testing (see G-A)? | [6 empty] | 3 |  |  | 0 |  |  | 5.11
Sub total | [8 empty] | 10 |  |  | 0
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6.
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G6.1 |  | Evaluation and audits: It is recommended that internal audits be conducted at least annually. External audits are conducted less frequently-assessors should use the recommendation of local accrediting bodies to determine the frequency of external audits. | [6 empty] | 5 | Answer all sub questions |  | 0 |  |  | 6.1 & 1.5
|  |  | Does the laboratory regularly conduct internal audits including the TB laboratory? | [12 empty] | 100
|  |  | Are external audits regularly conducted including the TB section? | [12 empty] | 100
|  |  | Are audit recommendations and action plans followed up within the timeframe defined by the laboratory? | [12 empty] | 100 | 300
Sub total | [8 empty] | 5 |  |  | 0
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G7.1 |  | Has the laboratory formulated specifications for supplies and consumables and are they followed during the procurement process? | [6 empty] | 2 |  |  | 0 |  |  | 7.2
G7.2 |  | Are storage areas for reagents and supplies set up, maintained and monitored according to manufacturer’s requirements? Ensure all supplies/reagents have not expired | [6 empty] | 2 |  |  | 0 |  |  | 7.8
Sub total | [8 empty] | 4 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G8.1 |  | Does the laboratory request form require the date and time of sample collection to be recorded? | [6 empty] | 3 |  |  | 0 |  |  | 8.2 & 8.3
G8.2 |  | Does the laboratory request form have space for the diagnosis/reason for testing? | [6 empty] | 2 |  |  | 0 |  |  | 8.2 & 8.3
G8.3 |  | Does the laboratory determine the cause of failed QC (root cause analysis) for all TB tests, perform corrective actions and measure the effectiveness thereof? | [6 empty] | 2 |  |  | 0 |  |  | 8.10
G8.4 |  | Does the laboratory refrain from reporting in case of failed QC? | [6 empty] | 2 |  |  | 0 |  |  | 8.10
G8.5 |  | External quality assessment: | [6 empty] | 5 | Answer all sub questions |  | 0 |  |  | 8.14
|  |  | Is the laboratory enrolled in an interlaboratory comparison, (blinded) rechecking or PT program for all TB tests evaluated? | [12 empty] | 100
|  |  | Did the laboratory pass the last 3 rounds of interlaboratory comparison or PT program testing? | [12 empty] | 100
|  |  | Does the laboratory receive onsite supervision visits as part of the National EQA program for TB testing? | [12 empty] | 100 | 300
Sub total | [8 empty] | 14 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G9.1 |  | Does the laboratory: | [6 empty] | 3 | Answer all sub questions |  | 0 |  |  | 9.3
|  |  | - report TB test results on an NTP approved TB report form? | [12 empty] | 100
|  |  | - record results correctly from the instruments and laboratory register to the report form? | [12 empty] | 100
|  |  | - report positive results or DST resistance results to NTP/NTRL? | [12 empty] | 100 | 300
Sub total | [8 empty] | 3 |  |  | 0
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10.
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G10.1 |  | Are all nonconforming activities identified and documented adequately? | [6 empty] | 5 |  |  | 0 |  |  | 10.1
G10.2 |  | Is root cause analysis performed and corrective action implemented for all non-conforming work? | [6 empty] | 3 |  |  | 0 |  |  | 10.2
Sub total | [8 empty] | 8 |  |  | 0
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G11.1 |  | Is the following performance indicator collected? - Number of samples for TB testing rejected (disaggregated by type of test and origin) | [6 empty] | 2 |  |  | 0 |  |  | 11.4 / 11.5
G11.2 |  | Are aggregate reports shared periodically with clinicians/NTP/NTRL (as applicable)? Assessors should review the guidance documents of the relevant authorities such as MoH, NTP and/or surveillance committees to determine the frequency that the laboratory should share their reports. If no recommendations exist, this should be at least quarterly. | [6 empty] | 2 |  |  | 0 |  |  | 11.4 / 11.5
G11.3 |  | Do reports for clinicians/NTP/NTRL (as applicable) include at a minimum the number of samples, isolated or identified organisms and DST patterns? | [6 empty] | 2 |  |  | 0 |  |  | 11.4 / 11.5
Sub total | [8 empty] | 6 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition, assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer |  | Score |  | Comments | Related to SLIPTA question
G12.1 |  | Are the following PPE items used when processing sputum and extra pulmonary samples? | [6 empty] | 2 | Answer all sub questions |  | 0 |  |  | 12.16
|  |  | Gloves | [12 empty] | 100
|  |  | Laboratory coat or gown | [12 empty] | 100
|  |  | N95 respirator | [12 empty] | 100 | 300
G12.2 |  | Waste management: | [6 empty] | 2 | Answer all sub questions |  | 0 |  |  | 12.4
|  |  | Does the laboratory handle waste appropriately including disposal of media and infectious material generated during testing? | [12 empty] | 100
|  |  | Are suitable disinfectants available for use when processing samples, are they freshly prepared, and is there evidence of their use? Clinical Microbiology Reviews, Jan. 1999, P. 147–179 | [12 empty] | 100 | 200
G12.3 |  | Is a biological safety cabinet (BSC) or hood available, functional, maintained and serviced for handling specimens or organisms considered to be highly contagious by air borne routes? | [6 empty] | 2 |  |  | 0 |  |  | 12.8
Sub total | [8 empty] | 6 |  |  | 0
Totals
[9 empty] | 74 |  |  | 0
Links to other AMR modules
- Smear module
- Culture module
- DST module
- Xpert module
- TB-LAMP module
- LF-LAM module
- LPA module
- Truenat module
Yes
No
N/A
Yes
Partial
No
N/A
Public
Private
Academic
NGO
Other
National
Reference
Provincial / County
District / Sub-district
Zonal / Peripheral
Field
Other
Manual, paper-based
Electronic, commercial LIS
Electronic, in-house LIS
(Smart)phone
E-mail

## Smear
Score-TB Package v2.1 - September 2020
Smear Microscopy
S-A | How many tests were performed last year?
| Smear microscopy (diagnostic) |  |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Number of samples received: | [7 empty] | 0
| Number of samples rejected: | [7 empty] | 0
| Scanty positive: | [7 empty] | 0
| Positive: | [7 empty] | 0
| Negative: | [7 empty] | 0
| Smear microscopy (follow-up) |  |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Number of samples received: | [7 empty] | 0
| Number of samples rejected: | [7 empty] | 0
| Scanty positive: | [7 empty] | 0
| Positive: | [7 empty] | 0
| Negative: | [7 empty] | 0
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition to the General Procedures (Section 1), assessors should review the following:
Question no. |  | Question | [7 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
S1.1 |  | Does the laboratory have documentation covering the following processes? | [7 empty] | 3 | Answer all sub questions | 0 |  | 1.5
|  |  | 1. Production of stains for Ziehl Nielsen (ZN) and/or fluorescent (Auramine O) staining | [11 empty] | 100
|  |  | 2. Sample collection and transport | [11 empty] | 100
|  |  | 3. Processing of samples and conducting smear microscopy | [11 empty] | 100
|  |  | 4. Recording & reporting procedures Smear Microscopy results conforming to WHO standards | [11 empty] | 100
|  |  | 5. Interlaboratory comparison or proficiency testing (PT) for Smear Microscopy | [11 empty] | 100
|  |  | 6. Laboratory safety related to Smear Microscopy | [11 empty] | 100 | 600
S1.2 |  | Are the documents complete, in-date and witnessed by all staff performing smear microscopy? | [7 empty] | 2 |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
Sub total | [9 empty] | 5 |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the General Procedures (Section 2).
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the General Procedures (Section 3).
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition to the General Procedures (Section 4), assessors should review the following:
Question no. |  | Question | [7 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
S4.1 |  | Is there evidence that the laboratory has provided clients information/instructions on interpretation of smear microscopy results? | [7 empty] | 2 |  | 0 |  | 4.1
Sub total | [9 empty] | 2 |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5. In addition, assessors should review the General Procedures (Section 5).
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. In addition, assessors should review the General Procedures (Section 6).
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition to the General Procedures (Section 7), assessors should review the following:
Question no. |  | Question | [7 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
S7.1 |  | Are all reagents stored at the correct temperature, under the right conditions (in the dark where appropriate) and in date (according to manufacturer's requirements)? | [7 empty] | 2 | Answer all sub questions | 0 |  | 7.10
|  |  | - ZN stains | [11 empty] | 100
|  |  | - Fluorescent stains | [11 empty] | 100 | 200
Sub total | [9 empty] | 2 |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition to the General Procedures (Section 8), assessors should review the following:
Question no. |  | Question | [7 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
QUALITY CONTROL
S8.1 |  | Please answer: | [7 empty] | 2 | Answer all sub questions | 0 |  | 8.8
|  |  | - Do QC records for ZN stains demonstrate their ability to stain acid-fast organisms? | [11 empty] | 100
|  |  | - Do QC records for fluorescent stains demonstrate their ability to stain acid-fast organisms? | [11 empty] | 100
|  |  | - Do QC records show that QC is performed for each batch of slides | [11 empty] | 100 | 300
SMEAR MICROSCOPY PROCEDURE
S8.2 |  | Please answer: | [7 empty] | 5 | Answer all sub questions | 0 |  | 8.10
|  |  | Are smears prepared on clean, un-used glass slides? | [11 empty] | 100
|  |  | Before making the smear, is the slide cleaned with alcohol and clearly labelled with the laboratory number? | [11 empty] | 100
|  |  | Is a new swab-stick (or loop) used to collect from the specimen sediment (pellet) a representative portion of the sample for smearing? | [11 empty] | 100
|  |  | Is the smear approx. 2cm x 1cm and in the center of the slide? | [11 empty] | 100
|  |  | After drying, is fixation done by gentle heating? | [11 empty] | 100
|  |  | Are materials discarded in accordance with biosafety recommendations? | [11 empty] | 100
|  |  | Is ZN staining performed according to the laboratory’s SOP: | [7 empty] | Answer all sub questions |  |  |  | 100
|  |  |  | Heated & filtered Ziehl Neelsen Carbolfuchsin 3-5 minutes | [10 empty] | 100
|  |  |  | Decolorizing solution (e.g. 3% acid alcohol) for a maximum of 3 minutes | [10 empty] | 100
|  |  |  | Counterstain (e.g. Methylene blue) 1 minute | [10 empty] | 100 | 300
|  |  | Is fluorescent staining performed according to the laboratory’s SOP: | [7 empty] | Answer all sub questions |  |  |  | 100
|  |  |  | Auramine-O 15-20 minutes | [10 empty] | 100
|  |  |  | Decolorizing solution (e.g. 3% acid alcohol) for a maximum of 2 minutes | [10 empty] | 100
|  |  |  | Counterstain (e.g. potassium permanganate) 2 minutes | [10 empty] | 100 | 300
|  |  | Are QC results reviewed before test results are reported? | [11 empty] | 100 | 900
Sub total | [9 empty] | 7 |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition, assessors should review the General Procedures (Section 9).
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10. In addition, assessors should review the General Procedures (Section 10).
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition to the General Procedures (Section 11), assessors should review the following:
Question no. |  | Question | [7 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
S11.1 |  | Are the following performance indicators collected? | [7 empty] | 5 | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | - Number of smears performed disaggregated by diagnostic and follow-up | [11 empty] | 100
|  |  | - Number and proportion of positive, scanty and negative smears disaggregated by diagnostic and follow-up | [11 empty] | 100
|  |  | - Smear microscopy TAT (from sample collection to reporting) | [11 empty] | 100 | 300
Sub total | [9 empty] | 5 |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition, assessors should review the General Procedures (Section 12).
Totals
[10 empty] | 21 |  | 0
Links to other AMR modules
- General module
- Culture module
- DST module
- Xpert module
- TB-LAMP module
- LF-LAM module
- LPA module
- Truenat module
Yes
No
N/A
Yes
Partial
No
N/A

## Culture
Score-TB Package v2.1 - September 2020
TB culture for detection and identification of mycobacteria
C-A | How many tests were performed last year?
|  |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Number of samples rejected: | [6 empty] | 0
| Solid culture |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Positive MTB: | [6 empty] | 0
| Positive NTM: | [6 empty] | 0
| Negative: | [6 empty] | 0
| Contaminated: | [6 empty] | 0
| Mix culture (NTM&TB) | [6 empty] | 0
| Subtotal: |  |  | 0 | 0 | 0 | 0 | 0
| Liquid culture |  |  | Q1 | Q2 | Q3 | Q4
| Positive MTB: | [6 empty] | 0
| Positive NTM: | [6 empty] | 0
| Negative: | [6 empty] | 0
| Contaminated: | [6 empty] | 0
| Subtotal: |  |  | 0 | 0 | 0 | 0 | 0
| Total: |  |  | 0 | 0 | 0 | 0 | 0
MTB = mycobacteria tuberculosis
NTM = Non-tuberculosis mycobacteria
Q = quarter
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition to the General Procedures (Section 1), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
C1.1 |  | Does the laboratory have documentation covering the following processes? | [6 empty] | 3 | Answer all sub questions | 0 |  | 1.5
|  |  | 1. Preparation of media for solid culture and or QC on solid & liquid media | [10 empty] | 100
|  |  | 2. Sample collection and transport | [10 empty] | 100
|  |  | 3. Processing of pulmonary samples and conducting TB culture including decontamination and inoculation of media | [10 empty] | 100
|  |  | 4. Processing of extra-pulmonary samples and conducting TB culture including decontamination and inoculation of media | [10 empty] | 100
|  |  | 5. Identification of MTB from positive cultures | [10 empty] | 100
|  |  | 6. Recording & reporting TB culture and MTB identification test results | [10 empty] | 100
|  |  | 7. Quality control procedures for TB culture | [10 empty] | 100
|  |  | 8. Interlaboratory comparison or proficiency testing (PT) for all TB tests | [10 empty] | 100
|  |  | 9. Laboratory safety related to TB testing | [10 empty] | 100 | 900
C1.2 |  | Are the documents complete, in-date and witnessed by all staff performing TB culture and identification of MTB? | [6 empty] | 2 |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
Sub total | [8 empty] | 5 |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the General Procedures (Section 2).
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the General Procedures (Section 3).
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition to the General Procedures (Section 4), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
C4.1 |  | Is there evidence that the laboratory has provided clients information/instructions on interpretation of culture test results? | [6 empty] | 2 |  | 0 |  | 4.1
Sub total | [8 empty] | 2 |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5. In addition, assessors should review the General Procedures (Section 5).
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. In addition, assessors should review the General Procedures (Section 6).
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition to the General Procedures (Section 7), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
C7.1 |  | Are all media and consumables for TB culture and MTB identification testing stored at the correct temperature and in date (according to manufacturer's requirements)? | [6 empty] | 2 | Answer all sub questions | 0 |  | 7.10
|  |  | - Liquid media | [10 empty] | 100
|  |  | - Solid media | [10 empty] | 100
|  |  | - MTB identification tests | [10 empty] | 100 | 300
Sub total | [8 empty] | 2 |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition to the General Procedures (Section 8), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
QUALITY CONTROL | [8 empty] | 3 | Answer all sub questions | 0 |  | 8.8
C8.1 |  | Does the laboratory perform QC testing on all media before use? | [11 empty] | 100
|  | Do QC records for liquid culture media demonstrate their ability to support growth of MTB? | [11 empty] | 100
|  | Do QC records for solid culture media demonstrate their ability to support growth of MTB? | [11 empty] | 100
|  | Do QC records for solid culture media demonstrate their sterility? | [11 empty] | 100
|  | Do QC records for decontamination reagents demonstrate that they are sterile? | [11 empty] | 100
|  | Do QC records for MTB identification tests indicate their ability to identify MTB from NTM? | [11 empty] | 100 | 600
C8.2 |  | Does the laboratory: | [6 empty] | 3 | Answer all sub questions | 0 |  | 8.10
|  |  | 1. Perform sterility and performance tests for every batch of culture media using certified reference strains as controls? | [10 empty] | 100
|  |  | 2. Are reference strains (MTB H37Rv) sourced from an authorized supplier? | [10 empty] | 100
|  |  | 3. Are the reference strains stored, cultured and sub-cultured in accordance with the appropriate guidelines? | [10 empty] | 100 | 300
C8.3 |  | Does laboratory record all samples in batches along with controls on a processing worksheet? | [6 empty] | 2 |  | 0 |  | 8.10
TB CULTURE PROCEDURE - Decontamination | [8 empty] | 5 | Answer all sub questions | 0 |  | 8.10
C8.4 |  | Is TB culture performed in batches corresponding to the number of centrifuge buckets possible? | [11 empty] | 100
|  | Is the correct concentration of decontamination solution used? | [11 empty] | 100
|  | Is the volume of the specimen checked and equal volume of digestion- decontamination reagent added and thoroughly mixed? | [11 empty] | 100
|  | Are the decontamination- digestion mixtures incubated at room temperature (20°C to 25°C) for 15 minutes? | [11 empty] | 100
|  | Is buffer added to fill the tube? | [11 empty] | 100
|  | Are measures in place to ensure that the buffer is not contaminated with sample material? | [11 empty] | 100
|  | Are the samples centrifuged at a Relative Centrifugal Force (RCF) of 3000g for 15–20 minutes? | [11 empty] | 100
|  | Are centrifuge buckets opened in the biological safety cabinet (BSC) after allowing aerosols to settle? | [11 empty] | 100
|  | Is the supernatant decanted into a flask with tuberculocidal disinfectant? | [11 empty] | 100
|  | Is the correct buffer solution used? | [11 empty] | 100
|  | Are the samples re-suspended in the recommended volume of buffer? | [11 empty] | 100
|  | During processing, is only one specimen tube open at a time? | [11 empty] | 100
|  | Is a fresh pipette used at every step to avoid transfer of bacilli from one specimen to the other? | [11 empty] | 100
|  | Are the transfers between tubes done without tubes touching each other to avoid cross-contamination? | [11 empty] | 100
|  | Is aerosol production minimized by avoiding splashes and squirting from pipettes? | [11 empty] | 100
|  | Are materials discarded in accordance with local biosafety recommendations? | [11 empty] | 100 | 1600
LIQUID CULTURE PROCEDURE | [8 empty] | 5 | Answer all sub questions | 0 |  | 8.10
C8.5 |  | Is MGIT tube loading/unloading performed according to the SOP? | [11 empty] | 100
|  | Is smear made from remaining pellet and results used in result interpretation? | [11 empty] | 100
|  | Is the remaining sample stored appropriately for potential inoculation if contamination is detected? | [11 empty] | 100
|  | Are MGIT tubes incubated for 42 days before being reported as negative? | [11 empty] | 100
|  | Is a ZN smear prepared from MGIT positive tubes prior to speciation? | [11 empty] | 100
|  | Is the MGIT tube re-incubated if the culture is instrument positive, ZN microscopy negative? | [11 empty] | 100
|  | Are materials discarded in accordance with local biosafety recommendations? | [11 empty] | 100 | 700
SOLID CULTURE PROCEDURE | [8 empty] | 5 | Answer all sub questions | 0 |  | 8.10
C8.6 |  | Is a disposable pipette used to inoculate each slant with 3–4 drops ensuring that the entire surface of the slant is inoculated? | [11 empty] | 100
|  | Is a smear made from the remaining pellet and results used in result interpretation? | [11 empty] | 100
|  | Is the remaining sample stored appropriately for potential re-testing if contamination is detected? | [11 empty] | 100
|  | Is solid media initially incubated for up to one week in a slanted position such that the surface of the solid media is horizontal and facing upwards? | [11 empty] | 100
|  | Are LJ caps tightened after one day drying time? | [11 empty] | 100
|  | Are slants checked at least weekly for early signs of MTB growth/contamination? | [11 empty] | 100
|  | Is solid media incubated for 56 days before being reported as negative? | [11 empty] | 100
|  | Are materials discarded in accordance with local biosafety recommendations? | [11 empty] | 100 | 800
MTB IDENTIFICATION PROCEDURE | [8 empty] | 5 | Answer all sub questions | 0 |  | 8.10
C8.7 |  | Is MTB identification performed according to the SOP or manufacturer’s instructions? | [11 empty] | 100
|  | Are materials discarded in accordance with local biosafety recommendations? | [11 empty] | 100 | 200
Sub total | [8 empty] | 28 |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition assessors should review the General Procedures (Section 9).
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10. In addition assessors should review the General Procedures (Section 10).
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition to the General Procedures (Section 11), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
C11.1 |  | Are the following performance indicators collected? | [6 empty] | 5 | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | 1. Number of TB cultures performed (disaggregated by type)? | [10 empty] | 100
|  |  | 2. Number and proportion of positive, negative and contaminated / mixed TB cultures (disaggregated by type)? | [10 empty] | 100
|  |  | 3. Number and proportion of MTB and NTM isolated (disaggregated by type)? | [10 empty] | 100
|  |  | 4. Number and proportion of MTB and NTM isolated (disaggregated by patient group)? | [10 empty] | 100
|  |  | 5. Liquid and/or solid culture TAT | [10 empty] | 100 | 500
Sub total | [8 empty] | 5 |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition assessors should review the General Procedures (Section 12).
Totals
[9 empty] | 42 |  | 0
Links to other AMR modules
- General module
- Smear module
- DST module
- Xpert module
- TB-LAMP module
- LF-LAM module
- LPA module
- Truenat module
Yes
No
N/A
Yes
Partial
No
N/A

## DST
Score-TB Package v2.1 - September 2020
Phenotypic drug susceptibility testing (DST)
D-A | How many tests were performed last year?
[5 empty] | Phenotypic DST (e.g. MGIT) |  |  |  | Totals
| Isoniazid |  |  |  | Q1 | Q2 | Q3 | Q4
| Resistant | [7 empty] | 0
| Susceptible | [7 empty] | 0
| Total: |  |  |  | 0 | 0 | 0 | 0 | 0
| Rifampicin |  |  |  | Q1 | Q2 | Q3 | Q4 | Q1
| Resistant | [7 empty] | 0
| Susceptible | [7 empty] | 0
| Total: |  |  |  | 0 | 0 | 0 | 0 | 0
Q = quarter
| OTHER DRUG SENSITIVITIES |  |  |  | METHOD: |  |  |  | METHOD:
| DRUG: |  |  |  | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [11 empty] | 0
| Susceptible | [11 empty] | 0
| Subtotal: |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| DRUG: |  |  |  | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [11 empty] | 0
| Susceptible | [11 empty] | 0
| Subtotal: |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| DRUG: |  |  |  | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [11 empty] | 0
| Susceptible | [11 empty] | 0
| Subtotal: |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| DRUG: |  |  |  | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [11 empty] | 0
| Susceptible | [11 empty] | 0
| Subtotal: |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| DRUG: |  |  |  | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [11 empty] | 0
| Susceptible | [11 empty] | 0
| Subtotal: |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| DRUG: |  |  |  | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [11 empty] | 0
| Susceptible | [11 empty] | 0
| Subtotal: |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| DRUG: |  |  |  | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [11 empty] | 0
| Susceptible | [11 empty] | 0
| Subtotal: |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Q = quarter
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition to the General Procedures (Section 1), assessors should review the following:
Question no. |  | Question | [8 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
D1.1 |  | Does the laboratory have documentation covering the following processes? | [8 empty] | 3 |  | Answer all sub questions | 0 |  | 1.5
|  |  | 1. Preparation of drug concentrations and dilutions | [13 empty] | 100
|  |  | 2. Processing of samples for DST | [13 empty] | 100
|  |  | 3. Recording and reporting DST results | [13 empty] | 100
|  |  | 4. Quality control procedures for DST | [13 empty] | 100
|  |  | 5. EQA for DST | [13 empty] | 100
|  |  | 6. Laboratory safety required for DST | [13 empty] | 100 | 600
D1.2 |  | Are the documents complete, in-date and witnessed by all staff performing DST? | [8 empty] | 2 |  |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
Sub total | [10 empty] | 5 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the General Procedures (Section 2).
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the General Procedures (Section 3).
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition to the General Procedures (Section 4), assessors should review the following:
Question no. |  | Question | [8 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
D4.1 |  | Is there evidence that the laboratory has provided clients information/instructions on interpretation of DST test results? | [8 empty] | 2 |  |  | 0 |  | 4.1
Sub total | [10 empty] | 2 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5. In addition, assessors should review the General Procedures (Section 5).
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. In addition, assessors should review the General Procedures (Section 6).
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition to the General Procedures (Section 7), assessors should review the following:
Question no. |  | Question | [8 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
D7.1 |  | Are all media and consumables for DST testing stored at the correct temperature (antibiotics should be stored in a non-defrosting freezer) and in date? (according to manufacturer's requirements - antibiotic disk cartridges and strips should be stored in a tightly sealed container with active desiccants that are replaced or recharged at least monthly). | [8 empty] | 2 |  | Answer all sub questions | 0 |  | 7.10
|  |  | 1. Antibiotics | [13 empty] | 100
|  |  | 2. Media | [13 empty] | 100 | 200
Sub total | [10 empty] | 2 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition to the General Procedures (Section 8), assessors should review the following:
Question no. |  | Question | [8 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
QUALITY CONTROL
D8.1 |  | Answer the questions below: | [8 empty] | 3 |  | Answer all sub questions | 0 |  | 8.8
|  | Does the laboratory perform QC testing on all media before use? This includes in-house made or purchased from commercial sources. | [14 empty] | 100
|  | Do QC records for each antibiotic prepared demonstrate their ability to inhibit MTB (e.g. MTB h37rV)? | [14 empty] | 100 | 200
D8.2 |  | Does the laboratory: | [8 empty] | 3 |  | Answer all sub questions | 0 |  | 8.10
|  |  | 1. Perform sterility and performance tests for every batch of culture media using certified reference strains as controls? | [13 empty] | 100
|  |  | 2. Are reference strains sourced from an authorized supplier (MTB H37Rv)? | [13 empty] | 100
|  |  | 3. Are the reference strains stored, cultured and sub-cultured in accordance with the specification from the supplier? | [13 empty] | 100 | 300
D8.3 |  | Does laboratory record all samples in batch along with controls on a processing worksheet? | [8 empty] | 3 |  |  | 0 |  | 8.10
PHENOTYPIC DST | [10 empty] | 5 |  | Answer all sub questions | 0 |  | 8.10
D8.4 |  | Are all antibiotic reconstitution/addition steps performed in a biosafety cabinet? | [14 empty] | 100
|  | Is each tube or plate labelled with the relevant drug, concentration, laboratory number and date? | [14 empty] | 100
|  | Are all antibiotics reconstituted and/or diluted according to the SOP and / or manufacturer’s instructions? | [14 empty] | 100
|  | Is a separate pipette tip used for reconstitution and/or dilution of each antibiotic? | [14 empty] | 100
|  | Is 0.8 ml MGIT SIRE Supplement to each SIRE tube and the SIRE growth control tube? | [14 empty] | 100
|  | Are the appropriate reconstituted antibiotic solutions added into each of the corresponding labeled tubes or plate? | [14 empty] | 100
|  | Is no antibiotic added to the control tube or plate? | [14 empty] | 100
|  | Is a 1:100 dilution of the MTB inoculum prepared for the control tube or plate (except for PZA, for which the dilution should be 1:10)? | [14 empty] | 100
|  | Is the MTB inoculum for the drug tests prepared according to the SOP and / or manufactures instructions? | [14 empty] | 100
|  | Does the SOP cover the following details? | [10 empty] | Answer all sub questions |  |  |  | 100
|  |  | For the direct method, adjusting inoculum size based on the number of bacilli observed in a smear? | [13 empty] | 100
|  |  | For the indirect method: | [9 empty] | Answer all sub questions |  |  |  | 100 | 200
|  |  |  | Ensuring that the inoculum is representative for the original culture (scrapping growth from as many colonies as possible) | [12 empty] | 100
|  |  |  | Standardization based on the density of cells compared with the McFarland 1.0 standard | [12 empty] | 100
|  |  |  | Preparation of the inoculum for the growth control (1% of the inoculum for the drug-containing media)" | [12 empty] | 100 | 300
|  | Is the prepared MTB inoculum added to the plates, tubes and control plates or tubes? | [14 empty] | 100
|  | Are the tubes or plates incubated according to the SOP and / or manufacturer’s instructions? | [14 empty] | 100
|  | Are the DST results read / interpreted according to the SOP and/or manufacturer’s instructions? | [14 empty] | 100
|  | Are materials discarded in accordance with local biosafety recommendations? | [14 empty] | 100 | 1400
Sub total | [10 empty] | 14 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition assessors should review the General Procedures (Section 9).
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10. In addition assessors should review the General Procedures (Section 10).
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition to the General Procedures (Section 11), assessors should review the following:
Question no. |  | Question | [8 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
D11.1 |  | Are the following performance indicators collected? | [8 empty] | 5 |  | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | 1. Number of DST tests performed | [13 empty] | 100
|  |  | 2. Number of invalid DST results | [13 empty] | 100
|  |  | 3. Number and proportion of antibiotic resistance tests performed by DST tests (disaggregated by antibiotic) | [13 empty] | 100
|  |  | 4. Number and type of discordant results | [13 empty] | 100
|  |  | 5. Average DST test TAT (from sample collection to reporting) | [13 empty] | 100 | 500
Sub total | [10 empty] | 5 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition assessors should review the General Procedures (Section 12).
Totals
[11 empty] | 28 |  |  | 0
Links to other AMR modules
- General module
- Smear module
- Culture module
- Xpert module
- TB-LAMP module | [17 empty] | Yes
- LF-LAM module | [17 empty] | No
- LPA module | [17 empty] | N/A
- Truenat module
Yes
Partial
No
N/A

## Xpert
Score-TB Package v2.1 - September 2020
Xpert MTB/RIF (includes Xpert MTB/RIF Ultra)
X-A | How many tests were performed last year?
| Xpert MTB/RIF | [5 empty] | Q1 | Q2 | Q3 | Q4
| Number of samples received | [9 empty] | 0
| Number of samples rejected | [9 empty] | 0
| MTB detected (high, medium, low & very low) (T) | [9 empty] | 0
| MTB not detected (N) | [9 empty] | 0
| MTB detected Trace (TT) | [9 empty] | 0
| MTB detected RIF resistance detected (RR) | [9 empty] | 0
| MTB detected RIF resistance indeterminate (TI) | [9 empty] | 0
| Error (I) | [9 empty] | 0
| Invalid (I) | [9 empty] | 0
| No results (I) | [9 empty] | 0
Of the MTB detected samples, how many were:
|  | - Rif not detected | [8 empty] | 0
|  | - Rif detected | [8 empty] | 0
|  | - Rif indeterminate (not trace) | [8 empty] | 0
Q = quarter
X-B | Is the following equipment available, and if so, is it functional, monitored, serviced and maintained?
|  |  |  | Available |  | Functional |  | Monitored |  | Serviced |  | Maintained |  | <------- | Functional: is the equipment in working order? Monitored: is the functionality of equipment regularly checked (e.g. temperature / calibrated)? Serviced: is the equipment regularly serviced by a qualified service technician? Maintained: is the equipment regularly maintained according to the manufacturer’s recommendations (e.g. cleaning)?
GeneXpert
| If yes, how many modules are available? |  |  |  | of |  | modules
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition to the General Procedures (Section 1), assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
X1.1 |  | Does the laboratory have documentation covering the following processes? | [10 empty] | 3 |  | Answer all sub questions | 0 |  | 1.5
|  |  | 1. Processing of samples and conducting Xpert MTB/RIF testing | [15 empty] | 100
|  |  | 2. Processing of extra-pulmonary samples and conducting Xpert MTB/RIF testing | [15 empty] | 100
|  |  | 3. Recording & reporting procedures Xpert MTB/RIF results conforming to WHO standards | [15 empty] | 100
|  |  | 4. Interlaboratory comparison or proficiency testing (PT) for Xpert MTB/RIF | [15 empty] | 100
|  |  | 5. Laboratory safety related to Xpert MTB/RIF | [15 empty] | 100 | 500
X1.2 |  | Are the documents complete, in-date and witnessed by all staff performing Xpert MTB/RIF testing? | [10 empty] | 2 |  |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
Sub total | [12 empty] | 5 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the General Procedures (Section 2).
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the General Procedures (Section 3).
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition to the General Procedures (Section 4), assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
X4.1 |  | Is there evidence that the laboratory has provided clients information/instructions on interpretation of Xpert MTB/RIF results? | [10 empty] | 2 |  |  | 0 |  | 4.1
Sub total | [12 empty] | 2 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5. In addition, assessors should review the General Procedures (Section 5).
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. In addition, assessors should review the General Procedures (Section 6).
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition to the General Procedures (Section 7), assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
X7.1 |  | Are all media and consumables stored at the correct temperature and in date (according to manufacturer's requirements)? - Xpert MTB/RIF reagents & cartridges | [10 empty] | 2 |  |  | 0 |  | 7.10
Sub total | [12 empty] | 2 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition to the General Procedures (Section 8), assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
QUALITY CONTROL
X8.1 |  | Do QC records for Xpert MTB/RIF indicate that incoming batches of Xpert MTB/RIF cartridges have been tested? | [10 empty] | 2 |  |  | 0 |  | 8.8
XPERT MBT/RIF PROCEDURE
X8.2 |  | Answer the questions below: | [10 empty] | 5 |  | Answer all sub questions | 0 |  | 8.10
|  | - Is the correct volume of SR (Substrate Reagent) added to the sample (1:2 ratio)? | [16 empty] | 100
|  | - Is the sample mixed twice before the end of incubation time? | [16 empty] | 100
|  | - Is the sample incubated for 15 – 30 minutes? | [16 empty] | 100
|  | - Is the correct (2ml) input volume transferred to the cartridge? | [16 empty] | 100
|  | - Are the patient data entered correctly on the GeneXpert instrument ? | [16 empty] | 100 | 500
X8.3 |  | Is there proof that actions are undertaken in case of errors, invalid results and no results? | [10 empty] | 2 |  |  | 0 |  | 10.3
Sub total | [12 empty] | 9 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition assessors should review the General Procedures (Section 9).
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10. In addition assessors should review the General Procedures (Section 10).
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition to the General Procedures (Section 11), assessors should review the following:
Question no. |  | Question | [10 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
X11.1 |  | Are the following performance indicators collected? | [10 empty] | 5 |  | Answer all sub questions | 0 |  | 11.4 & 11.5
Xpert MTB/RIF procedure:
|  |  | - MTB detected (high, medium, low & very low) (T) | [15 empty] | 100
|  |  | - MTB not detected (N) | [15 empty] | 100
|  |  | - MTB detected Trace (TT) | [15 empty] | 100
|  |  | - MTB detected RIF resistance detected (RR) | [15 empty] | 100
|  |  | - MTB detected RIF resistance indeterminate (TI) | [15 empty] | 100
|  |  | - Error (I) | [15 empty] | 100
|  |  | - Invalid (I) | [15 empty] | 100
|  |  | - No results (I) | [15 empty] | 100
|  |  | Number of extra-pulmonary samples tested disaggregated by type | [15 empty] | 100
|  |  | Xpert MTB/RIF TAT (from sample collection to reporting) | [15 empty] | 100 | 1000
Sub total | [12 empty] | 5 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition assessors should review the General Procedures (Section 12).
Totals
[13 empty] | 23 |  |  | 0
Links to other AMR modules
- General module
- Smear module
- Culture module
- DST module
- TB-LAMP module | [19 empty] | Yes
- LF-LAM module | [19 empty] | No
- LPA module | [19 empty] | N/A
- Truenat module
Yes
Partial
No
N/A

## LF LAM
Score-TB Package v2.1 - September 2020
Lateral Flow Urine Lipoarabinomannan Assay (LF-LAM)
L-A | How many tests were performed last year?
| LF-LAM |  |  | Q1 | Q2 | Q3 | Q4
| Number of samples received: | [6 empty] | 0
| Number of samples rejected: | [6 empty] | 0
| Positive: | [6 empty] | 0
| Negative: | [6 empty] | 0
| Indeterminate: | [6 empty] | 0
Q = Quarter
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition to the General Procedures (Section 1), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
L1.1 |  | Does the laboratory have documentation covering the following processes? | [6 empty] | 3 | Answer all sub questions | 0 |  | 1.5
|  |  | 1. Sample collection and transport | [10 empty] | 100
|  |  | 2. Processing of samples and conducting LF-LAM testing on urine samples | [10 empty] | 100
|  |  | 3. Quality control procedures for LF-LAM | [10 empty] | 100
|  |  | 4. Recording & reporting results of LF-LAM conforming to WHO standards | [10 empty] | 100
|  |  | 5. Interlaboratory comparison or proficiency testing (PT) for LF-LAM | [10 empty] | 100
|  |  | 6. Laboratory safety related to LF-LAM | [10 empty] | 100 | 600
L1.2 |  | Are the documents complete, in-date and witnessed by all staff performing LF-LAM testing? | [6 empty] | 2 |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
Sub total | [8 empty] | 5 |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the General Procedures (Section 2).
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the General Procedures (Section 3).
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition to the General Procedures (Section 4), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
L4.1 |  | Answer the questions below: | [6 empty] | 2 | Answer all sub questions | 0 |  | 4.1
|  | - Is there evidence that the laboratory has provided clients information/ instructions on collection of samples and interpretation of LF-LAM results? | [11 empty] | 100
|  | - Do urine collection instructions to patients include wiping and mid-stream catch? | [11 empty] | 100 | 200
Sub total | [8 empty] | 2 |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5. In addition, assessors should review the General Procedures (Section 5).
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. In addition, assessors should review the General Procedures (Section 6).
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition to the General Procedures (Section 7), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
L7.1 |  | Are all media and consumables stored at the correct temperature and in date (according to manufacturer's requirements)? - LF-LAM tests | [6 empty] | 2 |  | 0 |  | 7.10
Sub total | [8 empty] | 2 |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition to the General Procedures (Section 8), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
QUALITY CONTROL
L8.1 |  | Is lot to lot testing performed? | [6 empty] | 2 |  | 0 |  | 8.8
LF-LAM TESTING PROCEDURE
L8.2 |  | Answer the questions below: | [6 empty] | 5 | Answer all sub questions | 0 |  | 8.10
|  | - Is a timer available and used? | [11 empty] | 100
|  | - Are unused kits stored resealed with desiccant? | [11 empty] | 100
|  | - Is the correct volume of urine pipetted directly onto the inoculation pad? | [11 empty] | 100
|  | - Is the test is read within the appropriate timeframe? | [11 empty] | 100
|  | - Are test results interpreted correctly? | [11 empty] | 100
|  | - Are indefinite/invalid tests repeated? | [11 empty] | 100
|  | - Are materials discarded in accordance with local biosafety recommendations? | [11 empty] | 100 | 700
Sub total | [8 empty] | 7 |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition assessors should review the General Procedures (Section 9).
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10. In addition assessors should review the General Procedures (Section 10).
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition to the General Procedures (Section 11), assessors should review the following:
Question no. |  | Question | [6 empty] | Available score | Answer | Score | Comments | Related to SLIPTA question
L11.1 |  | Are the following performance indicators collected? | [6 empty] | 5 | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | - Number of LF-LAM tests performed | [10 empty] | 100
|  |  | - Number and proportion of positive, negative, indefinite and invalid results | [10 empty] | 100
|  |  | - LF-LAM TAT (from sample collection to reporting) | [10 empty] | 100 | 300
Sub total | [8 empty] | 5 |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition assessors should review the General Procedures (Section 12).
Totals
[9 empty] | 21 |  | 0
Links to other AMR modules
- General module
- Smear module
- Culture module
- DST module
- Xpert module
- TB-LAMP module
- LPA module
- Truenat module
Yes
No
N/A
Yes
Partial
No
N/A

## TB LAMP
Score-TB Package v2.1 - September 2020
Loop-Mediated Isothermal Amplification (TB-LAMP)
T-A | How many tests were performed last year?
| TB-LAMP |  |  |  |  | Q1 | Q2 | Q3 | Q4
| Number of samples received: | [8 empty] | 0
| Number of samples rejected: | [8 empty] | 0
| Positive: | [8 empty] | 0
| Negative: | [8 empty] | 0
| Indeterminate: | [8 empty] | 0
Q = Quarter
T-B | Is the following equipment available, and if so, is it functional, monitored, serviced and maintained?
|  |  |  | Available | Functional |  | Monitored |  | Serviced |  | Maintained |  | <------- | Functional: is the equipment in working order? Monitored: is the functionality of equipment regularly checked (e.g. temperature / calibrated)? Serviced: is the equipment regularly serviced by a qualified service technician? Maintained: is the equipment regularly maintained according to the manufacturer’s recommendations (e.g. cleaning)?
LAMP amplification
LAMP visualization
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition to the General Procedures (Section 1), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
T1.1 |  | Does the laboratory have documentation covering the following processes? | [9 empty] | 3 |  | Answer all sub questions | 0 |  | 1.5
|  |  | 1. Sample collection and transport | [14 empty] | 100
|  |  | 2. Processing of samples and conducting TB-LAMP | [14 empty] | 100
|  |  | 3. Quality control procedures for TB-LAMP | [14 empty] | 100
|  |  | 4. Recording & reporting test results TB-LAMP conforming to WHO standards | [14 empty] | 100
|  |  | 5. Interlaboratory comparison or proficiency testing (PT) for TB-LAMP | [14 empty] | 100
|  |  | 6. Laboratory safety related to TB-LAMP | [14 empty] | 100 | 600
T1.2 |  | Are the documents complete, in-date and witnessed by all staff performing TB-LAMP testing? | [9 empty] | 2 |  |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
Sub total | [11 empty] | 5 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the General Procedures (Section 2).
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the General Procedures (Section 3).
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition to the General Procedures (Section 4), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
T4.1 |  | Is there evidence that the laboratory has provided clients information/instructions on interpretation of TB-LAMP results? | [9 empty] | 2 |  |  | 0 |  | 4.1
Sub total | [11 empty] | 2 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5. In addition, assessors should review the General Procedures (Section 5).
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. In addition, assessors should review the General Procedures (Section 6).
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition to the General Procedures (Section 7), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
T7.1 |  | Are all media and consumables stored at the correct temperature and in date (according to manufacturer's requirements)? - TB-LAMP reagents | [9 empty] | 2 |  |  | 0 |  | 7.10
Sub total | [11 empty] | 2 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition to the General Procedures (Section 8), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
QUALITY CONTROL
T8.1 |  | Answer the questions below: | [9 empty] | 2 |  | Answer all sub questions | 0 |  | 8.10
|  | - Is lot-to-lot testing performed | [15 empty] | 100
|  | - Do QC records for TB-LAMP demonstrate reagent ability to detect positive MTB? | [15 empty] | 100 | 200
TB-LAMP TESTING PROCEDURE This procedure describes Loopamp PURE DNA Extraction and MTBBC detection (Eiken Chemical Co Ltd.)
DNA extraction
T8.2 |  | Answer the questions below: | [9 empty] | 5 |  | Answer all sub questions | 0 |  | 8.10
|  | - Is 60 µL of a suitable sputum containing a purulent portion pipetted to the Heating Tube? | [15 empty] | 100
|  | - Is 60 µL of the negative control pipetted to the negative control Heating Tube? | [15 empty] | 100
|  | - Are the Heating Tubes mixed three to ﬁve times? | [15 empty] | 100
|  | - Are the Heating Tubes loaded onto a pre-heated heating block and inactivated for 5 minutes? | [15 empty] | 100
|  | - Are the Heating Tubes cooled at room temperature for 2 minutes? | [15 empty] | 100
|  | - Are the Heating Tubes connected to the Absorbent Tube? | [15 empty] | 100
|  | - Is the Heating Tube - Absorbent Tube assembly mixed 20 times (until the powder becomes slurry)? | [15 empty] | 100
|  | - Is the Injection Cap attached to the Absorbent Tube? | [15 empty] | 100
|  | - Is 30 µL of the DNA solution dispensed to the Reaction Tube? | [15 empty] | 100
|  | - Are DNA extraction materials discarded in accordance with local biosafety recommendations? | [15 empty] | 100 | 1000
MTBC DETECTION
T8.3 |  | Answer the questions below: | [9 empty] | 5 |  | Answer all sub questions | 0 |  | 8.10
|  | - Is 30 µL of positive control reagent transferred to the positive control Reaction Tube using the dropper provided? | [15 empty] | 100
|  | - Are all Reaction Tubes spun down? | [15 empty] | 100
|  | - Are the reagents in the Reaction Tubes reconstituted by inverting the Tubes 5 times? | [15 empty] | 100
|  | - Are all Reaction Tubes spun down (a second time)? | [15 empty] | 100
|  | - Is the amplification procedure performed at 67°C for 45 minutes? | [15 empty] | 100
|  | - Are the results interpreted according to the manufacturer’s instructions? | [15 empty] | 100
|  | - Are indefinite / invalid tests repeated? | [15 empty] | 100
|  | - Are QC results reviewed before test results are reported? | [15 empty] | 100
|  | - Are MTBC detection materials discarded in accordance with local biosafety recommendations? | [15 empty] | 100 | 900
Sub total | [11 empty] | 12 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition assessors should review the General Procedures (Section 9).
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10. In addition assessors should review the General Procedures (Section 10).
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition to the General Procedures (Section 11), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
T11.1 |  | Are the following performance indicators collected? | [9 empty] | 5 |  | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | - Number of TB-LAMP tests performed | [14 empty] | 100
|  |  | - Number and proportion of positive, negative and invalid TB-LAMP results | [14 empty] | 100
|  |  | - TB-LAMP TAT (from sample collection to reporting) | [14 empty] | 100 | 300
Sub total | [11 empty] | 5 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition assessors should review the General Procedures (Section 12).
Totals
[12 empty] | 26 |  |  | 0
Links to other AMR modules
- General module
- Smear module
- Culture module
- DST module
- Xpert module
- LF-LAM module
- LPA module
- Truenat module
Yes
No
N/A
Yes
Partial
No
N/A

## LPA
Score-TB Package v2.1 - September 2020
Line Probe Assay (LPA) (includes MTBDRplus and MTBDRsl & CM [for speciation])
P-A | How many tests were performed last year?
LPA (MTBDRplus & MTBDRsl)
| ISONIAZID |  |  |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [8 empty] | 0
| Susceptible | [8 empty] | 0
| Subtotal: |  |  |  |  | 0 | 0 | 0 | 0 | 0
| RIFAMPICIN |  |  |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [8 empty] | 0
| Susceptible | [8 empty] | 0
| Subtotal: |  |  |  |  | 0 | 0 | 0 | 0 | 0
| ETHAMBUTOL |  |  |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [8 empty] | 0
| Susceptible | [8 empty] | 0
| Subtotal: |  |  |  |  | 0 | 0 | 0 | 0 | 0
| INJECTABLES |  |  |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [8 empty] | 0
| Susceptible | [8 empty] | 0
| Subtotal: |  |  |  |  | 0 | 0 | 0 | 0 | 0
| QUINOLONES |  |  |  |  | Q1 | Q2 | Q3 | Q4 | Totals
| Resistant | [8 empty] | 0
| Susceptible | [8 empty] | 0
| Subtotal: |  |  |  |  | 0 | 0 | 0 | 0 | 0
Q = quarter
P-B | How many tests were performed last year?
LPA (CM)
[6 empty] | Q1 | Q2 | Q3 | Q4 | Totals
| TOTAL: | [8 empty] | 0
Q = Quarter
P-C | Is the following equipment available, and if so, is it functional, monitored, serviced and maintained?
|  |  |  | Available | Functional |  | Monitored |  | Serviced |  | Maintained |  | <------- | Functional: is the equipment in working order? Monitored: is the functionality of equipment regularly checked (e.g. temperature / calibrated)? Serviced: is the equipment regularly serviced by a qualified service technician? Maintained: is the equipment regularly maintained according to the manufacturer’s recommendations (e.g. cleaning)?
GT-Blot
Thermocycler
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition to the General Procedures (Section 1), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P1.1 |  | Does the laboratory have documentation covering the following processes? | [9 empty] | 3 |  | Answer all sub questions | 0 |  | 1.5
|  |  | 1. Processing of LPA samples | [14 empty] | 100
|  |  | 2. Recording & reporting LPA results | [14 empty] | 100
|  |  | 3. Quality control procedures for LPA | [14 empty] | 100
|  |  | 4. Interlaboratory comparison or proficiency testing (PT) for LPA | [14 empty] | 100
|  |  | 5. Laboratory safety required for LPA from positive culture | [14 empty] | 100 | 500
P1.2 |  | Are the documents complete, in-date and witnessed by all staff performing LPA? | [9 empty] | 2 |  |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
Sub total | [11 empty] | 5 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the General Procedures (Section 2).
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the General Procedures (Section 3).
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition to the General Procedures (Section 4), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P4.1 |  | Is there evidence that the laboratory has provided clients information/instructions on interpretation of LPA test results? | [9 empty] | 2 |  |  | 0 |  | 4.1
Sub total | [11 empty] | 2 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5. In addition, assessors should review the General Procedures (Section 5).
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. In addition, assessors should review the General Procedures (Section 6).
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition to the General Procedures (Section 7), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P7.1 |  | Are all media and consumables for LPA testing stored at the correct conditions, temperature and in date (according to manufacturer's requirements - LPA light sensitive reagents should be stored in the dark)? - LPA test reagents | [9 empty] | 2 |  |  | 0 |  | 7.10
Sub total | [11 empty] | 2 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition to the General Procedures (Section 8), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
QUALITY CONTROL
P8.1 |  | Do QC records for LPA demonstrate the ability to detect MTB mutations? | [9 empty] | 2 |  |  | 0 |  | 8.9
P8.2 |  | Does the laboratory determine the cause of failed QC (root cause analysis) for LPA, perform corrective actions and measure the effectiveness thereof? | [9 empty] | 2 |  |  | 0 |  | 8.9
P8.3 |  | Does laboratory record all samples in a batch along with controls on a processing worksheet? | [9 empty] | 2 |  |  | 0 |  | 8.9
LPA PROCEDURE
Decontamination procedure from direct sputum
Refer to TECHNICAL SCORECARD: TB culture for detection and identification of mycobacteria.
Extraction method
P8.4 |  | Answer the questions below: | [9 empty] | 5 |  | Answer all sub questions | 0 |  | 8.10
|  | - Are samples used correctly collected, stored and received in timely fashion for testing according to SOPs? | [15 empty] | 100
|  | - Are cultures checked for contamination before use? | [15 empty] | 100
|  | - Is a positive and negative extraction control included with each batch? | [15 empty] | 100
|  | - Is the appropriate volume of liquid culture and/or solid isolates and/or decontaminated sample material used for extraction? Refererence: GenoLyse® kit for extraction of bacterial DNA – Instructions for Use, Hain Lifescience, 10/2011. | [15 empty] | 100
|  | - Are filter tips used to avoid cross contamination? | [15 empty] | 100
|  | - Is heater block/water bath set to correct temperature for extraction? | [15 empty] | 100
|  | - Are correct incubation times observed used for each step of the extraction? | [15 empty] | 100
|  | - Are correct centrifuge speeds used for each step of the extraction? | [15 empty] | 100
|  | - Is DNA reconstituted in correct volume of buffer? | [15 empty] | 100
|  | - Is extracted DNA stored in case retesting is required? | [15 empty] | 100 | 1000
Amplification and detection DST
P8.5 |  | Answer the questions below: | [9 empty] | 5 |  | Answer all sub questions | 0 |  | 8.10
|  | - Is the preparation of the Master Mix performed in the pre-amplification room? | [15 empty] | 100
|  | - Is the volume of amplification mixture made consistent with the SOP and manufacturer’s instructions? | [15 empty] | 100
|  | - Is a positive and negative control amplified in each batch? | [15 empty] | 100
|  | - Is amplification performed according to the SOP and / or manufactures instructions depending on sample type used? | [15 empty] | 100
|  | - Is thermocycler ramp rate correctly set? | [15 empty] | 100
|  | - Is hybridization performed in the Post-Amplification Room? | [15 empty] | 100
|  | - Is hybridization performed according to the SOP and/or manufactures instructions? | [15 empty] | 100
|  | - Are the DST results always read/interpreted according to the latest guidelines and/or manufacturer’s instructions? | [15 empty] | 100
|  | - Are materials discarded in accordance with local biosafety recommendations? | [15 empty] | 100 | 900
Sub total | [11 empty] | 16 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition assessors should review the General Procedures (Section 9).
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10. In addition assessors should review the General Procedures (Section 10).
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition to the General Procedures (Section 11), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
P11.1 |  | Are the following performance indicators collected? | [9 empty] | 5 |  | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | 1. Number of LPA tests performed | [14 empty] | 100
|  |  | 2. Number of invalid results | [14 empty] | 100
|  |  | 3. Number NTM identified by disaggregated by type if available | [14 empty] | 100
|  |  | 4. Number and proportion of drug resistance performed by LPA (disaggregated by mutation / drug) | [14 empty] | 100
|  |  | 5. Number and type of discordant results | [14 empty] | 100
|  |  | 6. Average LPA test TAT | [14 empty] | 100 | 600
Sub total | [11 empty] | 5 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition assessors should review the General Procedures (Section 12).
Totals
[12 empty] | 30 |  |  | 0
Links to other AMR modules
- General module
- Smear module
- Culture module
- DST module
- Xpert module
- TB-LAMP module
- LF-LAM module
- Truenat module
Yes
No
N/A
Yes
Partial
No
N/A

## Truenat
Score-TB Package v2.1 - September 2020
Truenat (includes Truenat MTB, Truenat MTB Plus and Truenat MTB-Rif Dx)
N-A | How many tests were performed last year?
| Truenat MTB & Truenat MTB Plus |  |  |  |  | Q1 | Q2 | Q3 | Q4 | Total
| Number of samples received | [8 empty] | 0
| Number of samples rejected | [8 empty] | 0
| Positive | [8 empty] | 0
| Negative | [8 empty] | 0
| Invalid | [8 empty] | 0
| Error | [8 empty] | 0
| Truenat MTB -Rif Dx |  |  |  |  | Q1 | Q2 | Q3 | Q4 | Total
| Number of samples processed | [8 empty] | 0
| Rif resistance detected | [8 empty] | 0
| Rif resistance indeterminate | [8 empty] | 0
| Rif resistance not detected | [8 empty] | 0
| Error | [8 empty] | 0
Q = quarter
N-B | Is the following equipment available, and if so, is it functional, monitored, serviced and maintained?
|  |  |  | Available | Functional |  | Monitored |  | Serviced |  | Maintained |  | <------- | Functional: is the equipment in working order? Monitored: is the functionality of equipment regularly checked (e.g. temperature / calibrated)? Serviced: is the equipment regularly serviced by a qualified service technician? Maintained: is the equipment regularly maintained according to the manufacturer’s recommendations (e.g. cleaning)?
Trueprep Auto
Truelab Uno Dx
Trulab Duo
Truelab Quattro
Section 1: Documents and records
All generic requirements apply, see SLIPTA Section 1. In addition to the General Procedures (Section 1), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
N1.1 |  | Does the laboratory have documentation covering the following processes? | [9 empty] | 3 |  | Answer all sub questions | 0 |  | 1.5
|  |  | 1. Sample collection and transport | [14 empty] | 100
|  |  | 2. Processing of samples and conducting Truenat | [14 empty] | 100
|  |  | 3. Quality control procedures for Truenat | [14 empty] | 100
|  |  | 4. Recording & reporting results of Truenat conforming to WHO standards | [14 empty] | 100
|  |  | 5. Interlaboratory comparison or proficiency testing (PT) for Truenat | [14 empty] | 100
|  |  | 6. Laboratory safety related to Truenat | [14 empty] | 100 | 600
N1.2 |  | Are the documents complete, in-date and witnessed by all staff performing Truenat testing? | [9 empty] | 2 |  |  | 0 |  | 1.5
See ISO15189:2012 Clause 5.5.3 for minimum requirements for a technical Standard Operating Procedure (SOP)
Sub total | [11 empty] | 5 |  |  | 0
Section 2: Management Reviews
All generic requirements apply, see SLIPTA Section 2. In addition, assessors should review the General Procedures (Section 2).
Section 3: Organization & Personnel
All generic requirements apply, see SLIPTA Section 3. In addition, assessors should review the General Procedures (Section 3).
Section 4: Client Management & Customer Service
All generic requirements apply, see SLIPTA Section 4. In addition to the General Procedures (Section 4), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
N4.1 |  | Is there evidence that the laboratory has provided clients information/instructions on interpretation of Truenat results? | [9 empty] | 2 |  |  | 0 |  | 4.1
Sub total | [11 empty] | 2 |  |  | 0
Section 5: Equipment
All generic requirements apply, see SLIPTA Section 5. In addition, assessors should review the General Procedures (Section 5).
Section 6: Evaluation and Audits
All generic requirements apply, see SLIPTA Section 6. In addition, assessors should review the General Procedures (Section 6).
Section 7: Purchasing & Inventory
All generic requirements apply, see SLIPTA Section 7. In addition to the General Procedures (Section 7), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
N7.1 |  | Are all media and consumables stored at the correct temperature and in date (according to manufacturer's requirements)? - Truenat reagents | [9 empty] | 2 |  |  | 0 |  | 7.10
Sub total | [11 empty] | 2 |  |  | 0
Section 8: Process Control
All generic requirements apply, see SLIPTA Section 8. In addition to the General Procedures (Section 8), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
QUALITY CONTROL
N8.1 |  | Do QC records for Truenat indicate that incoming batches of Truenat assays have been tested? | [9 empty] | 2 |  |  | 0 |  | 8.9
TRUENAT TESTING PROCEDURE
N8.2 |  | Answer the questions below: | [9 empty] | 5 |  | Answer all sub questions | 0 |  | 8.10
|  | - Are 2 drops of liquefaction buffer added to the sputum? | [15 empty] | 100
|  | - Is the sputum & liquefaction buffer mixed and then incubated for 5 minutes? | [15 empty] | 100
|  | - Is 0.5 ml of the liquified sputum transferred to the lysis buffer bottle? | [15 empty] | 100
|  | - Are 2 drops of liquefaction buffer added to the lysis buffer bottle? | [15 empty] | 100
|  | - Is the lysis bottle mixed and incubated for 3 minutes? | [15 empty] | 100
|  | - Are the entire contents of lysis buffer bottle transferred to the sample chamber of the cartridge using the provided transfer pipette? | [15 empty] | 100
|  | - Is the cartridge correctly inserted into the Trueprep Auto instrument? | [15 empty] | 100
|  | - Is extraction started by pressing the start button on the Trueprep Auto instrument? | [15 empty] | 100
|  | - After extraction, is the elute chamber seal on the cartridge pierced with the filter barrier tip of the 50µl Trueprep™ Precision Micropipette? | [15 empty] | 100
|  | - Is 50 µl of the elute transferred into the ECT tube before mixing well? | [15 empty] | 100
|  | - Is 6 μl of elute from ECT tube dispensed onto the centre of the white reaction well (MTB or MTB-Rif)? | [15 empty] | 100 | 1100
Sub total | [11 empty] | 7 |  |  | 0
Section 9: Information Management
All generic requirements apply, see SLIPTA Section 9. In addition assessors should review the General Procedures (Section 9).
Section 10: Identification of Nonconformities, Corrective and Preventive Actions
All generic requirements apply, see SLIPTA Section 10. In addition assessors should review the General Procedures (Section 10).
Section 11: Occurrence/Incident Management & Process Improvement
All generic requirements apply, see SLIPTA Section 11. In addition to the General Procedures (Section 11), assessors should review the following:
Question no. |  | Question | [9 empty] | Available score |  | Answer | Score | Comments | Related to SLIPTA question
N11.1 |  | Are the following performance indicators collected? | [9 empty] | 5 |  | Answer all sub questions | 0 |  | 11.4 & 11.5
|  |  | 1. Number of Truenat tests performed | [14 empty] | 100
|  |  | 2. Number and proportion of positive, negative, invalid and error Truenat results | [14 empty] | 100
|  |  | 3. Number and proportion of Rif resistance detected, Rif not detected, Rif indeterminate and error Truenat MTB-Rif Dx results | [14 empty] | 100
|  |  | 4. Truenat TAT | [14 empty] | 100 | 400
Sub total | [11 empty] | 5 |  |  | 0
Section 12: Facilities and Biosafety
All generic requirements apply, see SLIPTA Section 12. In addition assessors should review the General Procedures (Section 12).
Totals
[12 empty] | 21 |  |  | 0
Links to other AMR modules
- General module
- Smear module
- Culture module
- DST module
- Xpert module
- TB-LAMP module
- LF-LAM module
- LPA module
Yes
No
N/A
Yes
Partial
No
N/A

## Previous Audit Information
Score-TB Package v2.1 - September 2020
Previous Audit Information
Prior TB audit status | Baseline Audit |  | Audit 1 |  | Audit 2 |  | Audit 3
Date of previous audit
Please provide values for each section | Score | Percentage | Score | Percentage | Score | Percentage | Score | Percentage
General TB module
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
Smear module
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
Culture module
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
DST module
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
Xpert MTB-RIF
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
TB-LAMP
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
LF-LAM
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
LPA
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
Truenat
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Total
Prior SLIPTA audit status | Baseline Audit |  | Audit 1 |  | Audit 2 |  | Audit 3
Date of previous audit
Prior audit star rating
Please provide values for each section | Score | Percentage | Score | Percentage | Score | Percentage | Score | Percentage
1. Documents & Records
2. Management Reviews
3. Organization & Personnel
4. Client Management & Customer Service
5. Equipment
6. Internal Audit
7. Purchasing & Inventory
8. Process Control, IQA and EQA
9. Information Management
10. Corrective Action
11. Occurrence Management & Process Improvement
12. Facilities & Safety
Overall

## TB SUMMARY REPORT
Score-TB Package v2.1 - September 2020
TB Summary Report
LABORATORY DETAILS
Laboratory name:
Laboratory address:
Laboratory telephone number:
Laboratory email:
Laboratory level:
Laboratory type:
Contact person of laboratory:
AUDIT DETAILS
Audit carried out by:
Audit date:
TB AUDIT SCORE SUMMARY
|  | Cxurrent audit |  | Previous audit results, if any (select)
|  | Percentage score | Number of stars | Percentage score |  | Number of stars
General TB Module
Smear module
Culture module
DST module
Xpert MTB-RIF
TB-LAMP
LF-LAM
LPA
Truenat
Average
GENERAL TB MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 3 | 0 | 0%
2. Management Reviews |  |  |  | 4 | 0 | 0%
3. Organization and Personnel |  |  |  | 6 | 0 | 0%
4. Client Management and Customer Service |  |  |  | 5 | 0 | 0%
5. Equipment |  |  |  | 10 | 0 | 0%
6. Evaluation and Audits |  |  |  | 5 | 0 | 0%
7. Purchasing and Inventory |  |  |  | 4 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 14 | 0 | 0%
9. Information Management |  |  |  | 3 | 0 | 0%
10. Corrective Action |  |  |  | 8 | 0 | 0%
11. Occurrence Management and Process Improvement |  |  |  | 6 | 0 | 0%
12. Facilities and Safety |  |  |  | 6 | 0 | 0%
Module total |  |  |  | 74 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
SMEAR MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 5 | 0 | 0%
2. Management Reviews |  |  |  | N/A | N/A | N/A
3. Organization and Personnel |  |  |  | N/A | N/A | N/A
4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
5. Equipment |  |  |  | N/A | N/A | N/A
6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 7 | 0 | 0%
9. Information Management |  |  |  | N/A | N/A | N/A
10. Corrective Action |  |  |  | N/A | N/A | N/A
11. Occurrence Management and Process Improvement |  |  |  | 5 | 0 | 0%
12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Module total |  |  |  | 21 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
CULTURE MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 5 | 0 | 0%
2. Management Reviews |  |  |  | N/A | N/A | N/A
3. Organization and Personnel |  |  |  | N/A | N/A | N/A
4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
5. Equipment |  |  |  | N/A | N/A | N/A
6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 28 | 0 | 0%
9. Information Management |  |  |  | N/A | N/A | N/A
10. Corrective Action |  |  |  | N/A | N/A | N/A
11. Occurrence Management and Process Improvement |  |  |  | 5 | 0 | 0%
12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Module total |  |  |  | 42 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
DST MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 5 | 0 | 0%
2. Management Reviews |  |  |  | N/A | N/A | N/A
3. Organization and Personnel |  |  |  | N/A | N/A | N/A
4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
5. Equipment |  |  |  | N/A | N/A | N/A
6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 14 | 0 | 0%
9. Information Management |  |  |  | N/A | N/A | N/A
10. Corrective Action |  |  |  | N/A | N/A | N/A
11. Occurrence Management and Process Improvement |  |  |  | 5 | 0 | 0%
12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Module total |  |  |  | 28 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
XPERT MTB-RIF MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 5 | 0 | 0%
2. Management Reviews |  |  |  | N/A | N/A | N/A
3. Organization and Personnel |  |  |  | N/A | N/A | N/A
4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
5. Equipment |  |  |  | N/A | N/A | N/A
6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 9 | 0 | 0%
9. Information Management |  |  |  | N/A | N/A | N/A
10. Corrective Action |  |  |  | N/A | N/A | N/A
11. Occurrence Management and Process Improvement |  |  |  | 5 | 0 | 0%
12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Module total |  |  |  | 23 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
TB-LAMP MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 5 | 0 | 0%
2. Management Reviews |  |  |  | N/A | N/A | N/A
3. Organization and Personnel |  |  |  | N/A | N/A | N/A
4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
5. Equipment |  |  |  | N/A | N/A | N/A
6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 12 | 0 | 0%
9. Information Management |  |  |  | N/A | N/A | N/A
10. Corrective Action |  |  |  | N/A | N/A | N/A
11. Occurrence Management and Process Improvement |  |  |  | 5 | 0 | 0%
12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Module total |  |  |  | 26 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
LF-LAM MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 5 | 0 | 0%
2. Management Reviews |  |  |  | N/A | N/A | N/A
3. Organization and Personnel |  |  |  | N/A | N/A | N/A
4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
5. Equipment |  |  |  | N/A | N/A | N/A
6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 7 | 0 | 0%
9. Information Management |  |  |  | N/A | N/A | N/A
10. Corrective Action |  |  |  | N/A | N/A | N/A
11. Occurrence Management and Process Improvement |  |  |  | 5 | 0 | 0%
12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Module total |  |  |  | 21 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
LPA MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 5 | 0 | 0%
2. Management Reviews |  |  |  | N/A | N/A | N/A
3. Organization and Personnel |  |  |  | N/A | N/A | N/A
4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
5. Equipment |  |  |  | N/A | N/A | N/A
6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 16 | 0 | 0%
9. Information Management |  |  |  | N/A | N/A | N/A
10. Corrective Action |  |  |  | N/A | N/A | N/A
11. Occurrence Management and Process Improvement |  |  |  | 5 | 0 | 0%
12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Module total |  |  |  | 30 | 0 | 0%
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
TRUENAT MODULE
Section |  |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
[5 empty] | Date: |  | Date:
[5 empty] | Current audit score |  | Previous audit score
1. Documents and Records |  |  |  | 5 | 0 | 0%
2. Management Reviews |  |  |  | N/A | N/A | N/A
3. Organization and Personnel |  |  |  | N/A | N/A | N/A
4. Client Management and Customer Service |  |  |  | 2 | 0 | 0%
5. Equipment |  |  |  | N/A | N/A | N/A
6. Evaluation and Audits |  |  |  | N/A | N/A | N/A
7. Purchasing and Inventory |  |  |  | 2 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  |  | 7 | 0 | 0%
9. Information Management |  |  |  | N/A | N/A | N/A
10. Corrective Action |  |  |  | N/A | N/A | N/A
11. Occurrence Management and Process Improvement |  |  |  | 5 | 0 | 0%
12. Facilities and Safety |  |  |  | N/A | N/A | N/A
Module total |  |  |  | 21 | 0 | 0%
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12
N/A | 1. Documents & Records |  | S1
Baseline audit | 2. Management Reviews |  | S2
Audit 1 | 3. Organization & Personnel |  | S3
Audit 2 | 4. Client Management & Customer Service |  | S4
Audit 3 | 5. Equipment |  | S5
| 6. Evaluation & Audits |  | S6
| 7. Purchasing & Inventory |  | S7
| 8. Process Control, IQA & EQA |  | S8
| 9. Information Management |  | S9
| 10. Corrective Action |  | S10
| 11. Occurrence Management & Process Improvement |  | S11
| 12. Facilities & Safety |  | S12

## SLIPTA SUMMARY REPORT
Score-TB Package v2.1 - September 2020
SLIPTA Summary Report
LABORATORY DETAILS
Laboratory name:
Laboratory number:
Laboratory address:
Laboratory telephone number:
Laboratory fax number:
Laboratory email:
Laboratory level:
Laboratory type:
Head of laboratory:
Telephone number:
AUDIT DETAILS
Audit carried out by:
Audit date:
OVERALL AUDIT SCORE
| Percentage score | Star rating
Current audit: | 0% | 0 star
Select previous audit |  |  | <-- Select the dropdown if previous audit data was entered
DETAILED AUDIT SCORE
Section |  |  | Total points possible | Current Audit |  | Previous audit | Change in score since previous audit
|  |  |  | Audit date: |  | Audit date:
|  |  |  | Current audit score |  | Previous audit score
1. Documents and Records |  |  | 28 | 0 | 0%
2. Management Reviews |  |  | 14 | 0 | 0%
3. Organization and Personnel |  |  | 22 | 0 | 0%
4. Client Management and Customer Service |  |  | 10 | 0 | 0%
5. Equipment |  |  | 35 | 0 | 0%
6. Evaluation and Audits |  |  | 15 | 0 | 0%
7. Purchasing and Inventory |  |  | 24 | 0 | 0%
8. Process Control and Internal and External Quality Assessment |  |  | 32 | 0 | 0%
9. Information Management |  |  | 21 | 0 | 0%
10. Corrective Action |  |  | 19 | 0 | 0%
11. Occurrence Management and Process Improvement |  |  | 12 | 0 | 0%
12. Facilities and Safety |  |  | 43 | 0 | 0%
Total |  |  | 275 | 0 | 0%
N/A | 1. Documents & Records |  |  | S1
Baseline audit | 2. Management Reviews |  |  | S2
Audit 1 | 3. Organization & Personnel |  |  | S3
Audit 2 | 4. Client Management & Customer Service |  |  | S4
Audit 3 | 5. Equipment |  |  | S5
| 6. Evaluation & Audits |  |  | S6
| 7. Purchasing & Inventory |  |  | S7
| 8. Process Control, IQA & EQA |  |  | S8
| 9. Information Management |  |  | S9
| 10. Corrective Action |  |  | S10
| 11. Occurrence Management & Process Improvement |  |  | S11
| 12. Facilities & Safety |  |  | S12
| 13. AMR Technical proficiency |  |  | S13

## SLIPTA Laboratory Profile
SLIPTA Laboratory Profile
Site Information
Laboratory name:
Laboratory number:
Laboratory address:
Laboratory telephone number:
Laboratory fax number:
Laboratory email:
Laboratory level:
Laboratory type:
If other, please specify:
Head of Laboratory:
Telephone (Head of Laboratory):
Staff details
Number of degree holding staff: |  | Is this adequate?
Number of diploma holding staff: |  | Is this adequate?
Number of certificate holding staff: |  | Is this adequate?
Number of data clerks: |  | Is this adequate?
Number of phlebotomists: |  | Is this adequate?
Number of cleaners |  | Is this adequate?
Is the cleaner(s) dedicate to the laboratory only?
Has the cleaner(s) been trained in safe waste handling?
Number of drivers |  | Is this adequate?
Is the driver(s) dedicate to the laboratory only?
Has the driver(s) been trained in biosafety?
Number of other staff (specify positions below) |  | Is this adequate?
Supplemental details
Does the laboratory have sufficient space, equipment, supplies, personnel, infrastructure, etc. to execute the correct and timely performance of each test and maintain the quality management system?
Sufficient Space
Equipment
Supplies
Personnel
Infrastucture
Other (enter comments below)
National
Reference
Provincial
District
Zonal
Field
Public
Hospital
Private
Research
Non-hospital Outpatient Clinic
Other (please specify)
Yes
No
Insufficient Data
N/A
Baseline Audit
Audit 1
Audit 2
Audit 3
N/A
N/A
0 stars
1 star
2 stars
3 stars
4 stars
5 stars

## SLIPTA_S1
Section 1 - Documents and Records
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
1.1 | Legal Entity Does the laboratory have documentation stating its legal identity? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.2 The laboratory or the organization of which the laboratory is a part shall be an entity that can be held legally responsible for its activities. Note: Documentation could be in the form of a National Act, Company registration certificate, License number or Practice number.
1.2 | Laboratory Quality Manual Is there a current laboratory quality manual composed of the quality management system’s policies and procedures; has the manual content been communicated to staff, and is the manual understood and implemented by all staff? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.2.3 and 4.2.2.2 and 4.3 Note: A quality manual must be available that summarizes the laboratory’s quality management system, which includes policies that address all areas of the laboratory service, and identifies the goals and objectives of the quality management system. The quality manual must include policies and make reference to processes and procedures for all areas of the laboratory service and must address all the clauses of ISO15189:2012.
Does the quality manual includes the following elements:
|  | a) Quality policy statement that includes scope of service, standard of service, objectives of the quality management system, and management commitment to compliance | [5 empty] | 100
|  | b) Documented policies for the quality management system that meet the requirements of ISO15189:2012 (Refer to Question 1.5 of this checklist for list of policies required) | [5 empty] | 100
|  | c) Description of the quality management system and the structure of its documentation | [5 empty] | 100
|  | d) Reference to supporting procedures (SOPs), including managerial and technical procedures | [5 empty] | 100
|  | e) Description of the roles and responsibilities of the laboratory director, or laboratory manager, quality manager, and other key personnel (laboratory to define its key personnel) responsible for ensuring compliance | [5 empty] | 100
|  | f) Records of review and approval of the quality manual by authorized personnel | [5 empty] | 100
|  | g) Records to show that the quality manual was communicated to and understood by the lab personnel | [5 empty] | 100 | 700
1.3 | Document and Information Control System Does the laboratory have a system in place to control all documents and information (from internal and external sources)? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3 Note: There must be a procedure on document control. A document control system must be in place to ensure that records and all documents (internal and external) are current, read and understood by personnel, approved by authorized persons, reviewed periodically and revised as required. Documents must be uniquely identified to include title, page numbers, and authority of issue, document number, versions, effective date, and author. Example of external documents includes regulations, standards, guidelines, equipment user manuals, package inserts, text books.
1.4 | Document and Records Is there a list that details all documents used in the quality management system indicating their editions and distribution? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3 Note: Documents to be included on the list include Manuals, Procedures, and Processes. Work instructions. Forms, external documents.. The list could be in the form of a document master index, document log or document register. “Edition” can be regarded as synonymous with “revision or version” number for the documents.
1.5 | Laboratory Policies and Standard Operating Procedures Are policies and/or standard operating procedures (SOPs) for laboratory functions, technical and managerial procedures current, available and approved by authorized personnel? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.3 and 5.5 Note: The laboratory must define who is authorized to approve documents for its intended use. The approver should not be the author but can be the reviewer.
Has the laboratory defined Policies and/or SOPs that address the following:
|  | Ethical conduct How the laboratory will: 1) minimize activities that would diminish confidence in the laboratory's competence, impartiality, and judgment; 2) perform work within relevant legal requirements; 3) ensure confidentiality; 4) handle human samples, tissues or their remains as per regulations; 5) identify and avoid potential conflicts of interest and commercial, financial, political or other pressures that may affect the quality and integrity of operations? |  |  |  |  | ISO15189:2012 Clause 4.1.1.3 Note: Laboratories shall uphold the principle that the welfare and interest of the patient are paramount and patients should be treated fairly and without discrimination | 100
|  | Document Control How the laboratory will: 1) control all internal and external documents; 2) create documents; 3) identify documents; 4) review documents; 5) approve documents; 6) capture current versions and their distribution by means of a list; 7) handle amendments; 8) identify changes; 9) handle obsolete documents; 10) retain documents; 11) prevent the unintended use of any obsolete document; 12) ensure safe disposal of documents? |  |  |  |  | ISO15189:2012 Clause 4.3 and 4.13 Note: Documents that should be considered for document control are those that may vary based on changes in versions or time. Examples include policy statements, instructions for use, flow charts, procedures, specifications, forms, calibration tables, biological reference intervals and their origins, charts, posters, notices, memoranda, software documentation, drawings, plans, agreements, and documents of external origin such as regulations, standards and text books from which examination procedures are taken. | 100
|  | Control of Records How the laboratory will: 1) identify; 2) collect; 3) index; 4) access; 5) store; 6) maintain; 7) amends; 8) dispose of safely; 9) define the retention period for the identified records? |  |  |  |  | ISO15189 :2012 Clause 4.13 Note: Records can be in any form or type of medium providing they are readily accessible and protected from unauthorized alterations. Legal liability concerns regarding certain types of procedures (e.g. histology examinations, genetic examinations, pediatric examinations) may require the retention of certain records for much longer periods than for other records. For some records, especially those stored electronically, the safest storage may be on secure media and an offsite location. Type of records will include but not be limited to quality records, technical records, personnel records, test request and results records, | 100
|  | Communication (internal and external) How the laboratory will: 1) ensure effective communication with staff and users of the laboratory; 2) handle staff suggestions for improvement; 3) communicate with stakeholders on the effectiveness of the quality management system across all processes; 4) capture records of all communications; 5) retain and maintain all records of communication, requests, inquiries, verbal discussions and requests for additional examinations, meeting agendas, and meeting minutes)? |  |  |  |  | ISO15189:2012 Clause 4.1.2.6 and 4.14 Note: Laboratory management must ensure that appropriate communication processes are established between the laboratory and its stakeholders and that communication takes place regarding the effectiveness of the laboratory’s pre-examination, examination and post-examination processes and quality management system. | 100
|  | Service Agreements How the laboratory will: 1) establish service agreements; 2) review service agreements; 3) handle walk in patients (if applicable); 4) inform customers and users of any changes that affect the results of the requisition stated on the service agreement; 5) communicate to the requester of any work that has been referred; 6) retain records of communication? |  |  |  |  | ISO15189:2012 Clause 4.4.1 and 5.4 Notes: By accepting a requisition form from an authorized requester, the laboratory is considered to have entered into a service agreement. Customers and users may include clinicians, health care organizations, third party payment organizations or agencies, pharmaceutical companies, and patients. | 100
|  | Examination by Referral Laboratories and Consultants How the laboratory will: 1) select referral laboratories and consultants who provide opinions as well as interpretations; 2) evaluate and monitor the performance of referral laboratories and consultants who provide opinions as well as interpretations; 3) maintain a list of approved referral laboratories and consultants; 4) maintain a records of referred samples; 5) tracking of referred samples and their results; 6) report results from referral labs; 7) package and transport referred samples; 8) record communication of results from referral laboratories and consultants? |  |  |  |  | ISO15189:2012 Clause 4.5 and 5.8 and 4.13 Note: The laboratory must have a documented procedure for selecting and evaluating referral laboratories and consultants who provide opinions as well as interpretation for complex testing in any discipline. | 100
|  | External Services and Suppliers How the laboratory will: 1) select external purchases and services; 2) establish its selection criteria, including acceptance and rejection criteria; 3) approve and maintain its approved suppliers list; 4) define the requirements of its purchase supplies and services; 5) review and monitor the performance of its approved suppliers; 6) establish frequency of reviews? |  |  |  |  | ISO15189:2012 Clause 4.6 and 5.3 Note: The laboratory must have a documented procedure for the selection and purchasing of external services, equipment, reagents and consumable supplies that affect the quality of its service. | 100
|  | Purchasing and Inventory Control How the laboratory will: 1) request, order and receive supplies; 2) establish acceptance/rejection criteria for purchased items; 3) store purchased supplies; 4) control their inventory; 5) monitor and handle expired consumables? |  |  |  |  | ISO15189 :2012 Clause 4.6 and 5.3.2 Note: The laboratory shall have a documented procedure for the reception, storage, acceptance testing and inventory management of reagents and consumables. | 100
|  | Advisory Services How the laboratory will: 1) advise on the choice of examinations it offers; 2) communicate its advisory services to its users; 3) advise on clinical indications and limitations of examination procedures; 4) advise on the frequency of examination; 5 ) provide individual clinical case advice; 6) advise on interpretation of results; 7) promote the effective utilization of laboratory services; 8) provide consultation on scientific and logistic matters; 9) advise on the required type of sample and volume for testing? |  |  |  |  | ISO15189:2012 Clause 4.7 Note: The laboratory must have a system in place for providing advise to its users. | 100
|  | Resolution of Complaints and Feedback How the laboratory will: 1) manage complaints received from clinicians, patients, laboratory staff or other parties; 2) collect, receive and handle feedback received from clinicians, patients, laboratory staff or other parties; 3) keep records of all complaints, the investigations and actions taken, 4) determine the timeframe for closure and feedback to the complainant; 5) monitor effectiveness of corrective and preventative actions taken on complaints and feedback? |  |  |  |  | ISO15189:2012 Clause 4.8 and 4.10 Note: The laboratory must have a documented procedure for the management of complaints or other feedback received from clinicians, patients, laboratory staff or other parties. Records shall be maintained of all complaints and their investigation and the action taken | 100
|  | Identification and Control of Nonconformities (NC) How the laboratory will: 1) identify types of nonconformities in any aspect of the quality management system from pre, analytic and post analytic; 2) record NCs (how and where); 3) assign who is responsible for resolving the NC; 4) determine time frame for resolving NCs; 5) halt examinations (by an authorized person); 6) ensure the recall of released results of nonconforming or potentially nonconforming examinations; 7) release results after corrective action has been taken? |  |  |  |  | ISO15189:2012 Clause 4.9 Note: Nonconforming examinations or activities occur in many different areas and can be identified in many different ways, including clinician complaints, internal quality control indications, and instrument calibrations, checking of consumable materials, inter-laboratory comparisons, staff comments, reporting and certificate checking, laboratory management reviews, and internal and external audits. | 100
|  | Corrective Action (CA) How the laboratory will: 1) determine the root cause; 2) evaluate the need for CA to ensure that NCs do not recur; 3) assign the person responsible for the CA; 4) determine and implement CA(including person responsible and timeframe); 5) record CA taken; 6) monitor and review the effectiveness of the CA taken? |  |  |  |  | ISO15189:2012 Clause 4.10 Note: Action taken at the time of the nonconformity to mitigate effects is considered “immediate” action. Only action taken to remove the root cause of the problem that is causing the Non Conformities is considered “corrective” action. Any immediate action taken must also be documented. Corrective actions must be appropriate to the effects of the nonconformities encountered. | 100
|  | Preventive Action (PA) How the laboratory will: 1) review laboratory data and information to determine potential nonconformities; 2) determine the root cause(s) of potential nonconformities; 3) evaluate the need for preventive action; 4) record the PA; 5) determine and implement PA (including person responsible and timeframe); 6) monitor and review the effectiveness of implementation of PA? |  |  |  |  | ISO15189 :2012 Clause 4.11 Note: Preventive action is a proactive process for identifying opportunities for improvement rather than a reaction to the identification of problems or complaints (i.e. nonconformities). In addition to review of the operational procedures, preventive action might involve analysis of data, including trend and risk analyses and external quality assessment (proficiency testing). The laboratory shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be appropriate to the effects of the potential problems. | 100
|  | Continual Improvement How the laboratory will: 1) identify improvement activities within the Quality Management System; 2) develop improvement plans; 3) record improvement plans; 4) implement action plans; 5) communicate improvement plans and related goals to staff? |  |  |  |  | ISO15189:2012 Clause 4.1.1.2; 4.12; 4.14.5 Note: Improvement activities must be identified within the pre-examination, examination and post-examination processes. Laboratory management shall ensure that the laboratory participates in continual improvement activities that encompass relevant areas and outcomes of patient care. | 100
|  | Internal Audits How the laboratory will: 1) determine an audit schedule; 2) determine the roles and responsibilities for planning and conducting audits; 3) select the auditors; 4) define the types of audits; 4) define the frequency of audits; 5) define the scope of the internal audit; 6) record audit findings (forms and reports); 7) ensure corrective action is taken for all nonconformities identified with in the allocated time frame; 8) closure of Non Conformities identified during audits? |  |  |  |  | ISO15189:2012 Clause 4.14.5 Note: The cycle for internal auditing should normally be completed in one year. It is not necessary that internal audits cover each year, in depth, all elements of the quality management system. The laboratory may decide to focus on a particular activity without completely neglecting the others. The laboratory shall conduct internal audits at planned intervals to determine whether all activities in the quality management system, including pre-examination, examination, and post-examination. | 100
|  | Risk Management How the laboratory will: 1) evaluate the impact of potential pitfalls on work processes and examination results that affect patient results? (Refer to Question 6.3 of this checklist) |  |  |  |  | ISO15189:2012 Clause 4.14.6 Notes: Risk must be managed at the pre-examination processes, examination processes and post examination processes. The laboratory shall evaluate the impact of work processes and potential failures on examination results as they affect patient safety, and shall modify processes to reduce or eliminate the identified risks and document decisions and actions taken. | 100
|  | Management Review How the laboratory will: 1) define frequency of having management reviews; 2) define the agenda (input); 3) determine the key attendees; 4) record decisions and actions to be taken (output); 5) assign a person responsible and due dates for actions arising; 6) communicate decisions and actions to be taken to the relevant persons including laboratory staff; 7) ensure all actions arising are completed within the defined timeframe? (refer to Question 2.2 of this checklist for the agenda of the meeting) |  |  |  |  | ISO15189:2012 Clause 4.15 Note: Laboratory management shall review the quality management system at planned intervals to ensure its continuing suitability, adequacy and effectiveness and support of patient care. | 100
|  | Personnel Management How the laboratory will: 1) define the structure of the organization (organizational plan); 2) manage personnel (personnel policies); 3) maintain personnel records? (refer to Question 3.5 of the checklist for list of personnel records required) |  |  |  |  | ISO15189:2012 Clause 5.1.1; 5.1.9; 4.13 Note: The laboratory must have a documented procedure for personnel management and maintain records for all personnel to indicate compliance with requirements. | 100
|  | Personnel Training How the laboratory will: 1) perform staff orientation; 2) conduct initial and refresher training; 3) provide a continuous education program; 4) identify required training relevant to job title and responsibilities; 5) keep record of training; 6) evaluate the effectiveness of training? |  |  |  |  | ISO15189:2012 Clause 4.1.1.4 and 5.1.5 Note: Training includes external and internal trainings. The effectiveness of the training programme must be periodically reviewed. | 100
|  | Competency Assessment How the laboratory will: 1) assess the competence of personnel to perform assigned managerial or technical tasks; 2) assess ongoing competency; 3) establish competency criteria; 4) provide feedback to persons assessed; 5) schedule retraining based on the assessment outcome; 6) keep records of competency assessments and outcomes? |  |  |  |  | ISO15189 :2012 Clause 4.1.1.4 and 4.4 and 5.1.6 Note: Competency could be assessed using a combination of some or all of the following methods: direct observation; monitoring and the recording of examination results; review of work records; problem solving skills; blinded samples, review of accumulative IQC and EQA. Competency assessment for professional judgment should be designed as specific and fit for purpose. | 100
|  | Authorization How the laboratory will: 1) document authorization levels for the different tasks and roles; 2) appoint deputies for the key positions where appropriate? |  |  |  |  | ISO15189:2012 Clause 4.1.2 Note: Authorization may be in the form of a Job description, letter of appointment, approved authority matrix etc. | 100
|  | Review of Staff Performance How the laboratory will: 1) plan and perform staff appraisals; 2) establish frequency of monitoring and review of staff performance outcome; 3) keep records of staff performance; 4) train staff who perform staff appraisals? |  |  |  |  | ISO15189:2012 Clause 4.1.2.1 and 5.1.7 Note: In addition to the assessment of technical competence, the laboratory management must ensure that reviews of staff performance consider the needs of the laboratory and of the individual in order to maintain or improve the quality of service given to the users and encourage productive working relationships. Staff performing reviews should receive appropriate training. | 100
|  | Accommodation and Environmental Conditions How the laboratory will: 1) evaluate and determine the sufficiency and adequacy of the space allocated for the performance of and scope of work; 2) ensure the laboratory and office facilities are suitable for the tasks to be undertaken; 3) ensure the storage and disposal facilities meet the applicable requirements; 4) ensure staff have space for staff activities (supply of drinking water, storage space for personal and protective equipment and clothing); 5) monitor, control and record any specific environmental and accommodation requirements? |  |  |  |  | ISO15189:2012 Clause 4.1.1.4 and 5.2; 5.2.6 Note: The laboratory must have space allocated for the performance of its work that is designed to ensure the quality, safety and efficacy of the service provided to the users and the health and safety of laboratory personnel, patients and visitors. The laboratory shall evaluate and determine the sufficiency and adequacy of the space allocated for the performance of the work. Evaluating and determining the sufficiency and adequacy of space may be done during internal audits, risk assessments or at management review meeting, however it must be documented that it was evaluated and found to be adequate. | 100
|  | Laboratory equipment How the laboratory will: 1) select equipment; 2) purchase equipment; 3) manage equipment; 4) maintain equipment records 5) capture the minimum information on equipment label; 6) manage defective equipment; 7) define the equipment maintenance frequency; 8) record the maintenance ; 9) prevent unauthorized use (access control) of equipment; 10) manage obsolete equipment; 11) manage safe handling, transportation, storage and use to avoid deterioration and contamination, 12) track and verify completion of repairs? |  |  |  |  | ISO15189:2012 Clause 4.13; 5.3.1.1; 5.3.1.3 NOTE: For the purposes of this checklist, laboratory equipment includes hardware and software of instruments, measuring systems, and laboratory information systems. The laboratory shall have a documented procedure for the selection, purchasing and management of equipment. | 100
|  | Calibration of Equipment How the laboratory will: 1) define frequency of calibration; 2) handle in house calibrations (pipettes, thermometers, timers etc.); 3) record calibration status (use of stickers and calibration certificates); 4) handle failed calibrations? |  |  |  |  | ISO15189:2012 Clause 5.3.1.4 Notes: The laboratory must have a documented procedure for the calibration of equipment that directly or indirectly affects examination results. Documentation of calibration traceability to a higher order reference material or reference procedure may be provided by an examination system manufacturer. Such documentation is acceptable as long as the manufacturer's examination system and calibration procedures are used without modification. | 100
|  | Pre-examination Processes How the laboratory will provide information for patients and users on: 1) primary sample collection and handling; 2) instructions for pre-collection activities; 3) instructions for collection activities; 4) preparation and storage prior to dispatch to the laboratory; 5) sample and volume requirements; 6) Sample transportation; 7) time limits and special handling; 8) acceptance and rejection criteria; 9) confidentiality; 10) complaints procedure? |  |  |  |  | ISO15189:2012 Clause 5.4; 5.4.1; 5.4.3; 5.4.4.1; 5.4.5; 5.4.6; 5.4.7 Note: The laboratory must have documented procedures and information for pre-examination activities to ensure the validity of the results of examinations. | 100
|  | Validation and Verification of examination procedures / Equipment How the laboratory will: 1) select testing procedures; 2) perform equipment validation; 3) perform method validation; 4) perform equipment verification; 5) perform method verification; 6) define validation /verification protocol specific for each procedure at the time of validation or verification; 7) compare results from the different procedures, equipment, methods being used for the same test either located at the same site or at different sites? |  |  |  |  | ISO15189:2012 Clause 5.5.1.2; 5.6.4 and 5.5.1.3 Note: Validations should be done on a) non-standard methods; b) laboratory designed or developed methods; c) standard methods used outside their intended scope; d) validated methods subsequently modified. "Verification” is performed on methods that are being used without any modifications and is a process of evaluating of whether or not the procedure meets the performance characteristics stated by the manufacturer i.e. the manufacturer validation claims. The performance characteristics are obtained from the manufacture (validation reports) or from package inserts. Comparison of different methods used for same tests is ongoing verification. The frequency and characteristics to be checked in ongoing verification must be clearly defined. Note: All procedures or equipment used as backup must also be validated/verified as relevant. | 100
|  | Measurement of Uncertainty How the laboratory will: 1) determine Measurement of uncertainty on measured quantity values (quantitative tests); 2) define the performance requirements for the measurement uncertainty (e.g Standard Deviation; Clinical decision points)? Refer to Question 5.4 on this checklist |  |  |  |  | ISO15189:2012 Clause 5.5.1.4 Note: Uncertainty of measurement is used to indicate the confidence we have that the reported figure is correct. Uncertainty of measurement may be calculated using the calculated CV of at least 30 sets of internal QC data: CV% x 2 = Uncertainty of measurement (UM). The laboratory shall calculate the UM for all quantitative tests. These shall only be reported to clinicians if they request for them. For well-established methods, it is recommended a minimum of six months internal QC data should be used to calculate UM, updated at least annually where possible. For new methods at least 30 data points for each level of QC across at least two different batches of calibrator and reagents should be used to provide an interim estimate of uncertainty of measurement. | 100
|  | Biological Reference Intervals or Clinical Decision Values How the laboratory will: 1) define the biological reference intervals; 2) document the source of the reference intervals; 3) communicate changes to the users? |  |  |  |  | ISO15189:2012 Clause 5.5.2 Note: The laboratory shall define the biological reference intervals or clinical decision values, document the basis for the reference intervals or decision values and communicate this information to users. | 100
|  | Documentation of examination procedures How the laboratory will: 1) format general and technical Standard Operating Procedures; 2) define the minimum requirements for a SOP? |  |  |  |  | ISO15189:2012 Clause 5.5.3 Note: Working instructions, card files or similar systems that summarize key information are acceptable for use as a quick reference at the workbench, provided that a fully documented procedure is available for reference. Information from product instructions for use may be incorporated into examination procedures by reference in the SOP. The minimum requirements for a technical SOP should be a) purpose of the examination; b) principle and method of the procedure used for examinations; c) type of sample; d) required equipment and reagents; e) environmental and safety controls; f) procedural steps; g) interferences (e.g. lipemia, hemolysis, bilirubinemia, drugs) and cross reactions; h) principle of procedure for calculating results; i) laboratory clinical interpretation; j) potential sources of variation; k) references. | 100
|  | Laboratory Contingency Plan How the laboratory will ensure that there are no interruption to services in the event of: 1) staff shortage; 2) equipment breakdown; 3) prolonged power outages; 4) stock outs of reagents and consumables; 5) fire, natural disasters e.g. severe weather or floods, bomb threat or civil disturbances; 7) LIS failure? |  |  |  |  | ISO15189:2012 Clause 4.1.1.4; 5.2; 5.3.1; 5.10 Notes: the laboratory should maintain sufficient replacement parts to minimize testing downtime (e.g. pipette components, microscope bulbs and fuses, safety caps or buckets for safety centrifuge). Contingency plans should be periodically tested. Where the laboratory uses another laboratory as a backup, the performance of the back-up laboratory shall be regularly reviewed to ensure quality results | 100
|  | Quality control and Quality Assurance How the laboratory will: 1) use IQC and EQA (Interlaboratory comparison); 2) define the frequency of processing IQC; 3) define the acceptable ranges; 4) Evaluate and monitor laboratory performance using EQA and QC data; 5) troubleshoot unacceptable EQA and QC; 6) compare results using different procedures, equipment and sites; 7) notify users of any differences in comparability of results? |  |  |  |  | ISO15189:2012 Clause 4.10; 5.6; 5.6.2.1; 5.6.2.3; 5.6.3.1 Note: The laboratory should choose concentrations of control materials, wherever possible, especially at or near clinical decision values, which ensure the validity of decisions made. Use of independent third party control materials should be considered, either instead of, or in addition to, any control materials supplied by the reagent or instrument manufacturer. EQA should cover the pre-examination process, examination process and post examination process. Where an EQA program is not available, the laboratory can use alternative methods with clearly defined acceptable results e.g. exchange of samples with other labs, testing certified materials, sample previously tested. All procedures or equipment used as backup must also be included in EQA programme. | 100
|  | Reporting and Release of Results How the laboratory will : 1) issue standardized report (define the format and medium); 2) review patient results; 3) communicate patient results including alert, urgent and critical results; 4) ensure release of results to authorized persons; 5) amend reports; 6) issue of amended reports; 7) store patient results; 8) maintain patient results. (Refer to Question 9.3 of this checklist) |  |  |  |  | ISO15189:2012 Clause 5.8.1; 5.9.1 Note: Reports may be issued as a hard copy or electronically, all results issued verbally must be followed by a final report. The results of each examination must be reported accurately, clearly, unambiguously and in accordance with any specific instructions in the examination procedures. The laboratory must define the format and medium of the report (i.e. electronic or paper) and the manner in which it is to be communicated from the laboratory. | 100
|  | Laboratory Information System (LIS) (Computerized or non-computerized) How the laboratory will: 1) select a LIS; 2) verify /validate the LIS; 3) define authorities and responsibilities for the management and use of the information system; 4) ensure patient confidentiality is maintained at all times; 5) maintain the system; 6) back-up data; 7) safeguard against tempering by un-authorized users? |  |  |  |  | ISO15189:2012 Clause 5.10 Note: “information systems” includes the management of data and information contained in both computer and non-computerized systems. Some of the requirements may be more applicable to computer systems than to non-computerized systems. Computerized systems can include those integral to the functioning of laboratory equipment and stand-alone systems using generic software, such as word processing, spreadsheet and database applications that generate, collate, report and archive patient information and reports. | 100
|  | Laboratory Safety Manual How the laboratory will: 1) ensure all safety measures are implemented at the laboratory as applicable to national and international guidelines and regulations? (Refer to section 12 of this checklist for the contents of a safety manual) |  |  |  |  | ISO15190:2013 Clause 4.1.1.4; 5.2 Note: Laboratory management must implement a safe laboratory environment in compliance with good practice and applicable requirements. | 100
|  | Answer to TB question G1.1 | [5 empty] | 100
|  | Answer to TB question S1.1 | [5 empty] | 100
|  | Answer to TB question S1.2 | [5 empty] | 100
|  | Answer to TB question C1.1 | [5 empty] | 100
|  | Answer to TB question C1.2 | [5 empty] | 100
|  | Answer to TB question D1.1 | [5 empty] | 100
|  | Answer to TB question D1.2 | [5 empty] | 100
|  | Answer to TB question X1.1 | [5 empty] | 100
|  | Answer to TB question X1.2 | [5 empty] | 100
|  | Answer to TB question T1.1 | [5 empty] | 100
|  | Answer to TB question T1.2 | [5 empty] | 100
|  | Answer to TB question L1.1 | [5 empty] | 100
|  | Answer to TB question L1.2 | [5 empty] | 100
|  | Answer to TB question P1.1 | [5 empty] | 100
|  | Answer to TB question P1.2 | [5 empty] | 100
|  | Answer to TB question N1.1 | [5 empty] | 100
|  | Answer to TB question N1.2 | [5 empty] | 100 | 5200
1.6 | Policy and SOPs Accessibility Are policies and SOPs easily accessible/ available to all staff and written in a language commonly understood by respective staff? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.2.2.1; 4.3; 5.5 Note: All documentation must be current and approved by an authorized person. The documentation can be in any form or type of medium, providing it is readily accessible and protected from unauthorized changes and undue deterioration
1.7 | Policies and SOPs Communication Is there documented evidence that all relevant policies and SOPs have been communicated to and are understood and implemented by all staff as related to their responsibilities? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.2.2.2; 5.1.5(b) Note: The lab must have a system in place to ensure all staff are aware of the contents of all documents. All laboratory staff shall have access to and be instructed on the use and application of the quality manual and the referenced documents.
1.8 | Document Control Log Are policies and procedures dated to reflect when it was put into effect, its location, when it was reviewed and when it was discontinued? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3 Note: Current authorized editions and their distribution are identified by means of a list (e.g. document register, log or master index).
1.9 | Discontinued Policies and SOPs Are invalid or discontinued policies and procedures clearly marked / identified and removed from use and one copy retained for reference purposes? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3 Note: Obsolete controlled documents are dated and marked as obsolete. At least one copy of an obsolete controlled document is retained for a specified time period or in accordance with applicable specified requirements.
1.10 | Data Files Are test results, technical and quality records, invalid or discontinued policies and procedures archived for a specified time period in accordance with national/international guidelines? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.3; 4.13 Note: Copies or files of results should be archived. The retention period may vary; however, the reported results shall be retrievable for as long as medically relevant or as required by national, regional or local authorities.
1.11 | Archived Results Accessibility Is there an archiving system that allows for easy and timely retrieval of archived records and results? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13 Note: Records can be in any form or type of medium providing they are readily accessible and protected from unauthorized alterations. Archived patient results must be easily, readily and completely retrievable within a timeframe consistent with patient care needs.
| Totals |  | 28 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S2
Section 2 - Management Reviews
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
2.1 | Routine Review of Quality and Technical Records Does the laboratory supervisor routinely perform a documented review of all quality and technical records? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4; 4.2.1 There must be documentation that the laboratory manager/supervisor or a designee reviews the quality programme regularly. The review must ensure that recurrent problems have been addressed, and that new or redesigned activities have been evaluated.
Does the supervisor’s review include the following?
|  | a) Follow-up of action items from previous reviews | [5 empty] | 100
|  | b) Status of corrective actions taken and required preventive actions | [5 empty] | 100
|  | c) Reports from personnel | [5 empty] | 100
|  | d) Environmental monitoring log sheets | [5 empty] | 100
|  | e) Specimen rejection logbook | [5 empty] | 100
|  | f) Equipment calibration and maintenance records | [5 empty] | 100
|  | g) IQC records across all test areas | [5 empty] | 100
|  | h) Outcomes of PTs and other forms of inter-laboratory comparisons | [5 empty] | 100
|  | i) Quality indicators | [5 empty] | 100
|  | j) Customer complaints and feedback | [5 empty] | 100
|  | k) Results of improvement projects | [5 empty] | 100
|  | l) Documentation of review and action planning with staff for resolution and follow-up review | [5 empty] | 100 | 1200
2.2 | Management Review Does the laboratory management annually perform a review of all quality systems at a management review meeting? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4; 4.15.2; 4.15.4 Note: The interval between management reviews should be no greater than 12 months; however, shorter intervals should be adopted when a quality management system is being established.
Does the management review meeting include the following INPUTS?
|  | a) The periodic review of requests, and suitability of procedures and sample requirements | [5 empty] | 100
|  | b) Assessment of user feedback | [5 empty] | 100
|  | c) Staff suggestions | [5 empty] | 100
|  | d) Internal audits | [5 empty] | 100
|  | e) Risk management | [5 empty] | 100
|  | f) Use of quality indicators | [5 empty] | 100
|  | g) Assessments by external organizations | [5 empty] | 100
|  | h) Results of participation in interlaboratory comparison programmes (PT/EQA) | [5 empty] | 100
|  | i) Monitoring and resolution of complaints | [5 empty] | 100
|  | j) Performance of suppliers | [5 empty] | 100
|  | k) Identification and control of nonconformities | [5 empty] | 100
|  | l) Results of continual improvement including, current status of corrective actions and preventive actions | [5 empty] | 100
|  | m) Follow-up actions from previous management reviews | [5 empty] | 100
|  | n) Changes in the volume and scope of work, personnel, and premises that could affect the quality management system | [5 empty] | 100
|  | o) Recommendations for improvement, including technical requirements | [5 empty] | 100
|  | p) Review of quality objectives and the quality policy for appropriateness and continuous improvement | [5 empty] | 100
Does the management review meeting include the following OUTPUTS?
|  | a) Are management review outputs recorded? | [5 empty] | 100
|  | b) Does the output records of the MR meeting capture decisions made, persons responsible for actions to be taken and timeframes? | [5 empty] | 100
|  | c) Does the report address resources required (human, financial, material)? | [5 empty] | 100
|  | d) Does it refer to improvement for the users? | [5 empty] | 100
|  | e) Does it refer to improvement of the effectiveness of the quality system? | [5 empty] | 100
|  | f) Were the quality objectives and the quality policy reviewed for appropriateness and continuous improvement? | [5 empty] | 100
|  | Answer to TB question G2.1 | [5 empty] | 100
|  | Answer to TB question G2.2 | [5 empty] | 100 | 2400
2.3 | Are findings and actions from MR communicated to the relevant staff? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4; 4.15.4 Note: Findings and actions arising from management reviews shall be recorded and reported to laboratory staff.
2.4 | Does lab management ensure actions from MR are completed within defined timeframes? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4; 4.15.4 Note: Laboratory management shall ensure that actions arising from management review are completed within a defined timeframe.
| Totals |  | 14 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S3
Section 3 - Organization & Personnel
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
3.1 | Duty Roster and Daily Routine Does the laboratory have a duty roster that covers normal and after hours? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(c); 4.1.2.1(i) Note: A duty roster designates specific laboratory personnel to specific workstations. Daily routines should be prioritized, organized and coordinated to achieve optimal service delivery for patients.
3.2 | Organizational Chart and External/Internal Reporting Systems Is an organizational chart available that indicates the relationship between the laboratory and its parent organization? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.2.5 Note: An up-to-date organizational chart and/or narrative description should be available detailing the external and internal reporting relationships for laboratory personnel. The organizational chart or narrative should clearly show how the laboratory is linked to the rest of the hospital and laboratory services where applicable.
3.3 | Laboratory Director Is the laboratory directed by a person(s) with the competency and delegated responsibility to perform the following: |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4 Note: a director may be a person(s) with responsibility for, and authority over, a laboratory. The person or persons referred to may be designated collectively as laboratory director. Other settings may not use the term “Lab Director” but in this question, it refers to person/persons that are running the laboratory, however they decide to name them
|  | a) Provide effective leadership, budgeting and planning | [5 empty] | 100
|  | b) Communicate with stakeholders | [5 empty] | 100
|  | c) Ensure adequate competent staff | [5 empty] | 100
|  | d) Ensure the implementation of the QMS | [5 empty] | 100
|  | e) Selection and monitoring of lab supplies | [5 empty] | 100
|  | f) Selection and monitoring of referral labs | [5 empty] | 100
|  | g) Ensure a safe lab environment | [5 empty] | 100
|  | h) Advisory services | [5 empty] | 100
|  | i) Provide professional development programs for laboratory staff | [5 empty] | 100
|  | j) Address complaints, requests or suggestions from staff and/or lab users | [5 empty] | 100
|  | k) Design and implement a contingency plan | [5 empty] | 100 | 1100
3.4 | Quality Management System Oversight Is there a quality officer/manager with delegated responsibility to oversee compliance with the quality management system? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.2.7 There should be a quality manager (however named) with delegated authority to oversee compliance with the requirements of the quality management system. The quality manager must report directly to the level of laboratory management at which decisions are made on laboratory policy and resources.
|  | a) Is there an appointment letter, job description or terms of reference available? | [5 empty] | 100
|  | b) Does the quality manager ensure that processes needed for the quality management system are established, implemented, and maintained? | [5 empty] | 100
|  | c) Does the QM report to management which decisions relating to quality are made? | [5 empty] | 100
|  | d) Does the QM promote awareness of users’ needs and requirements throughout the organization? | [5 empty] | 100
|  | e) Does the QM participate in management reviews? | [5 empty] | 100 | 500
3.5 | Personnel Filing System Are records of personnel maintained and do they include the following? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.1.9 Note: Personnel files must be maintained for all current staff. Wherever (offsite or onsite) and however the records are kept, the records must be easily accessible when required. In some laboratories, not all records may be kept in a single file in one place e.g. training and competency records should be kept in the laboratory, medically related information with the administration.
If files are present, do they document or contain the following:
|  | a) Educational and professional qualifications | [5 empty] | 100
|  | b) Copy of certification or license to practice, when applicable | [5 empty] | 100
|  | c) Previous work experience e.g. CV | [5 empty] | 100
|  | d) Job description | [5 empty] | 100
|  | e) Record of introduction of new staff to the laboratory environment | [5 empty] | 100
|  | f) Training in current job tasks including vendor training received on-site | [5 empty] | 100
|  | g) Records of competency assessments | [5 empty] | 100
|  | h) Records of continuing education | [5 empty] | 100
|  | i) Reviews of staff performance | [5 empty] | 100
|  | j) Reports of accidents and exposure to occupational hazards | [5 empty] | 100
|  | k) Immunization status, as applicable, relevant to assigned duties | [5 empty] | 100
|  | l) Letter of employment or appointment | [5 empty] | 100
|  | m) Employee medical surveillance records | [5 empty] | 100
|  | Answer to TB question G3.1 | [5 empty] | 100 | 1400
3.7 | Laboratory Staff Training Is there a system for training that covers the following? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4(c); 5.1.5 Note: The effectiveness of the training program shall be reviewed regularly. Personnel that are undergoing training shall be supervised at all times.
|  | a) The quality management system | [5 empty] | 100
|  | b) Assigned work processes, procedures and tasks | [5 empty] | 100
|  | c) The applicable laboratory information system | [5 empty] | 100
|  | d) Health and safety, including the prevention or containment of the effects of adverse incidents | [5 empty] | 100
|  | e) Laboratory Ethics | [5 empty] | 100
|  | f) Confidentiality of patient information | [5 empty] | 100
|  | g) Is there supervision for persons undergoing training | [5 empty] | 100
|  | h) Continuous medical education | [5 empty] | 100
|  | i) Review of effectiveness of the training program | [5 empty] | 100
|  | Answer to TB question G3.2 | [5 empty] | 100 | 1000
3.8 | Staff Competency Assessment and Retraining Is there a system for competency assessment that covers the following? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.2.1(h); 5.1.6 Note: Newly hired lab staff must be assessed for competency before performing independent duties. All lab staff must undergo ongoing competency at a frequency defined by the laboratory. Staff assigned to a new section should be assessed before fully assuming independent duties. When deficiencies are noted, retraining and reassessment must be planned and documented. If the employee’s competency remains below standard, further action might include supervisory review of work, re-assignment of duties, or other appropriate actions. Records of competency assessments and resulting actions should be retained in personnel files and/or quality records. Records should show which skills were assessed, how those skills were measured, and who performed the assessment.
|  | a) Are competency assessments performed according defined criteria | [5 empty] | 100
|  | b) New hires | [5 empty] | 100
|  | c) Existing staff | [5 empty] | 100
|  | d) Retraining and re-assessment where needed | [5 empty] | 100 | 400
3.9 | Staff meetings Are staff meetings held regularly? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.2.1(a); (e); 4.1.2.2; 4.1.2.6; 4.4; 4.14.3 Note: The laboratory should hold regular staff meetings to ensure communication within the laboratory. Meetings should have recorded notes to facilitate review of progress over time.
Do meetings include the following items?
|  | a) Follow-up of action items from previous staff meetings | [5 empty] | 100
|  | b) Systemic and/or recurrent problems and issues addressed, including actions to prevent recurrence | [5 empty] | 100
|  | c) Complaints | [5 empty] | 100
|  | d) Communication on reviewed/revised/redundant SOPs | [5 empty] | 100
|  | e) Review of results from prior corrective actions | [5 empty] | 100
|  | f) Discussion and evaluation of improvement topics/projects | [5 empty] | 100
|  | g) Feedback given by staff that have attended hospital meetings, external meetings, training, conferences, workshops, etc. | [5 empty] | 100
|  | h) Relay of reports and updates from lab attendance at meetings with clinicians (the use of lab services and/or attendance at clinical rounds) | [5 empty] | 100
|  | i) Recording and monitoring of meeting notes for progress on issues. | [5 empty] | 100 | 900
| Totals |  | 22 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S4
Section 4 - Client Management and Customer Service
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
4.1 | Advice and Training by Qualified Staff |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4(g); 4.7 Note: Authorized staff should provide advice on sample type, examination choice, frequency and results interpretation.
|  | Do staff members with appropriate professional qualifications provide clients with advice and/or training regarding required types of samples, choice of examinations, repeat frequency, and interpretation of results? | [5 empty] | 100
|  | Answer to TB question G4.1 | [5 empty] | 100
|  | Answer to TB question S4.1 | [5 empty] | 100
|  | Answer to TB question C4.1 | [5 empty] | 100
|  | Answer to TB question D4.1 | [5 empty] | 100
|  | Answer to TB question X4.1 | [5 empty] | 100
|  | Answer to TB question T4.1 | [5 empty] | 100
|  | Answer to TB question L4.1 | [5 empty] | 100
|  | Answer to TB question P4.1 | [5 empty] | 100
|  | Answer to TB question N4.1 | [5 empty] | 100 | 1000
4.2 | Resolution of Complaints Does the laboratory investigate (review) and resolves of customer complaints? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(m); 4.8; 4.15.2(i) Note: The laboratory must have a documented procedure for the management of complaints or other feedback received from clinicians, patients, laboratory staff or other parties. Feedback must be given to the complainant.
4.3 | Laboratory Handbook for Clients - information for users |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.1.1.4(g); 4.5; 5.4.2 Note: The laboratory should provide its clients with a handbook that outlines the laboratory’s hours of operation, available tests, specimen collection instructions, packaging and shipping directions, and expected turnaround times.
|  | Is there a laboratory handbook for laboratory users that includes information on: - location of the laboratory, - services offered, - laboratory operating times, - instructions on completion of request forms, - instruction for preparation of the patient; - sample collection including patient collected samples, - transport, - agreed turnaround times, - acceptance and rejection criteria, - availability of advice on examination and interpretation of results; - lab policy on protection of personal information, - laboratory complaints procedure. | [5 empty] | 100
|  | Answer to TB question G4.2 | [5 empty] | 100 | 200
4.4 | Communication Policy on Delays in Service Is timely, documented notification provided to customers when the laboratory experiences delays or interruptions in testing (due to equipment failure, stock-outs, staff levels, etc.) or finds it necessary to change examination procedures, and when testing resumes? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.2.6; 4.4; 5.8.1 There should be a policy for notifying the requester when an examination is delayed. Such notification shall be documented for both service interruption and resumption as well as related feedback from clinicians. Clinical personnel do not need to be notified of all delays of examination but should be notified in those situations where the delay could compromise patient care.
4.5 | Evaluation Tool and Follow Up Is there a tool for regularly evaluating client satisfaction, staff suggestions and is the feedback received effectively utilized to improve services? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(m); 4.8; 4.14.3; 4.14.4 The laboratory should measure the satisfaction of clients, clinicians and patients regarding its services, either on an ongoing basis or through episodic solicitations.
| Totals |  | 10 |  | 0
Yes
Partial
No
n/a

## SLIPTA_S5
Section 5 - Equipment
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
5.1 | Adherence to Proper Equipment Protocol |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.1.2 Equipment should be properly placed as specified in user manual away from the following but not limited to water, direct sunlight, vibrations, in traffic and with more than 75% of the base of the equipment sitting on the bench top to avoid tip-over.
|  | Is equipment installed and placed as specified in the operator’s manuals and uniquely labeled or marked? | [5 empty] | 100
|  | Answer to TB question G5.2 | [5 empty] | 100 | 200
5.2 | Are equipment operated by trained, competent and authorized personnel? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.1.3 Note: The staff must be trained, and deemed competent to operate equipment.
5.3 | Equipment and Method Validation/Verification and Documentation Are newly introduced equipment and methods validated and verified on-site and are validation and verification reports available? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.1.2; 5.5.1 Newly introduced methods or equipment must be verified onsite to ensure that their introduction yields performance equal to or better than the previous method or equipment. Manufacturers’ validation may be used. Back up equipment must also be included in verification procedures.
|  | a) Are specific verification/validation protocols in place for each equipment and examination procedure? | [5 empty] | 100
|  | b) Is validation performed for all methods laboratory designed or developed methods by the laboratory, standard methods used outside their intended scope, and validated methods that are subsequently modified? | [5 empty] | 100
|  | c) Has validation information been obtained from the manufacturer/method developer as part of the verification? | [5 empty] | 100
|  | d) Have performance characteristics been appropriately selected and evaluated as per intended use? | [5 empty] | 100
|  | e) Were the verification/validation studies appropriate and adequate? | [5 empty] | 100
|  | f) Was the analysis of data appropriate for the selected performance characteristics? | [5 empty] | 100
|  | g) Have the verification/validation results/reports been reviewed and approved by an authorised person? | [5 empty] | 100
|  | Answer to TB question G5.1 | [5 empty] | 100 | 800
5.4 | Mesurement uncertainty of measured quantity tests Does the laboratory have documented estimates of measurement of uncertainty (UM)? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.5.1.4 Note: Measurement of uncertainty should be calculated at different clinical decision levels. Cumulative IQC (minimum 6 months data) may be used to calculate measurement of uncertainty.
|  | a) Has the laboratory calculated the measurement uncertainty for each quantitative measurement procedure? | [5 empty] | 100
|  | b) Has the laboratory defined the performance requirements (factors that affect the UM) for the measurement uncertainty of each measurement procedure and regularly review estimates of measurement uncertainty? | [5 empty] | 100
|  | c) Does the lab make its calculated measurement of uncertainty available to its users upon request? | [5 empty] | 100 | 300
5.5 | Equipment Maintenance Records Is current equipment inventory data available on all equipment in the laboratory? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.7 Records shall be maintained for each item of equipment used in the performance of examinations. Such equipment list must include major analysers as well as ancillary equipment like centrifuges, water baths, rotators, fridges, pipettes, timers, printers, computers.
|  | a) Name of equipment | [5 empty] | 100
|  | b) Manufacturer’s or authorized supplier contact details | [5 empty] | 100
|  | c) Condition received (new, used, reconditioned) | [5 empty] | 100
|  | d) Serial number | [5 empty] | 100
|  | e) Date of receiving | [5 empty] | 100
|  | f) Where equipment is obsolete, date when put “out of service” | [5 empty] | 100
|  | g) Date of entry into service after validation / verification) | [5 empty] | 100
|  | h) Location | [5 empty] | 100 | 800
5.6 | Equipment Maintenance Records Is relevant equipment service information readily available in the laboratory? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5; 5.3.1.7 Maintenance records must be maintained for each item of equipment used in the performance of examinations. These records shall be maintained and shall be readily available for the lifespan of the equipment or for any time period required by national, regional and local authorities.
|  | a) Service contract information or service schedule that has been adhered to | [5 empty] | 100
|  | b) Contact details for service provider | [5 empty] | 100
|  | c) Decontamination records before service , repair or decommissioning | [5 empty] | 100
|  | d) Engineer or service provider preventive maintenance records | [5 empty] | 100
|  | e) Last date of service | [5 empty] | 100
|  | f) Next date of service | [5 empty] | 100 | 600
5.7 | Defective Equipment Waiting for Repair Is defective equipment waiting for repair not used and clearly labelled? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5 Note: label should include the date of malfunction and “not in use” and a signature of approval.
5.8 | Obsolete Equipment Procedures Is non-functioning equipment appropriately labelled and removed from the laboratory or path of workflow following the equipment management policies and procedures? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5 Note: Label should include the date made obsolete and “obsolete” and a signature of approval.
5.9 | Equipment calibration and Metrological traceability Protocol |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.1.4 Note: Documentation of calibration traceability to a higher order reference material or reference procedure may be provided by an examination system manufacturer. Such documentation is acceptable as long as the manufacturer's examination system and calibration procedures are used without modification.
|  | a) Is routine calibration of laboratory equipment (including pipettes, centrifuges, balances and thermometers) scheduled, as indicated on the equipment, and verified? | [5 empty] | 100
|  | b) Is the calibration traceable (e.g. use of reference materials and equipment like certified thermometers, tachometer? | [5 empty] | 100
|  | c) Is there evidence of review of calibration certificates/results by the laboratory before acceptance back into use? | [5 empty] | 100
|  | d) Is certified reference materials, examination and calibration by another procedure, use of mutual consent standards or methods used for in house calibrations? | [5 empty] | 100 | 400
5.10 | Equipment Preventive Maintenance Is routine user preventive maintenance performed on all equipment and recorded according to manufacturer’s minimum requirements? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5 Note: Preventative maintenance by operators must be done on all equipment used in examinations including centrifuges, autoclaves, microscopes, and safety cabinets.
5.11 | Equipment Service Maintenance |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5 Note: All equipment must be serviced at specified intervals by a qualified service engineer either through service contracts or otherwise. Service schedule must at minimum meet manufacturer’s requirements.
|  | Is equipment routinely serviced according to schedule as per the minimum manufacturer recommendations by qualified and competent personnel and is this information documented in appropriate logs? | [5 empty] | 100
|  | Answer to TB question G5.3 | [5 empty] | 100 | 200
5.12 | Equipment Malfunction - Response and Documentation Is equipment malfunction resolved by corrective action(s) and the associated root cause analysis? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.9; 4.10, 4.13; 5.3.1.5 Note: All equipment malfunctions must be investigated and documented as per the non-conforming procedure. In the event that the user cannot resolve the problem, a repair order must be initiated.
5.13 | Equipment Repair Monitoring and Documentation - Are repair orders monitored to determine if the service is completed? - Does the laboratory verify and document that it is in proper working order before being put back into service? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.3.1.5; 5.6 Note: After a repair all levels of QC must or other performance checks must be processed to verify that the equipment is in proper working condition. Copies of the QC or performance checks results should be attached to the repair records as evidence.
5.14 | Equipment Failure - Contingency Plan Is there a functional back-up system that prevents interruption of lab services? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4 (n); 5.3.1 Note: Interruption to services is considered when a laboratory cannot release results to their users. Testing services should not be subject to interruption due to equipment malfunctions. Contingency plans must be in place, in the event of equipment failure, for the completion of testing. In the event of a testing disruption, planning may include the use of a back-up instrument, the use of a different testing method, the referral of samples to another laboratory.
5.15 | Manufacturer's Operator Manual Are the equipment manufacturer’s operator manuals readily available to testing staff, and where possible, available in the language understood by staff? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.1.3 Operator manuals must be readily available for reference by testing staff.
5.16 | Laboratory Testing Services Has the laboratory provided uninterrupted testing services, with no disruptions due to equipment failure in the last year (or since the last audit)? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(a);(n); 4.1.2.1(i) Note: Interruption to services is considered when a laboratory cannot release results to their users. Testing services should not be subject to interruption due to equipment malfunctions. Contingency plans must be in place, in the event of equipment failure, for the completion of testing. In the event of a testing disruption, planning may include the use of a back-up instrument, the use of a different testing method, the referral of samples to another laboratory
| Totals |  | 35 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S6
Section 6 - Evaluation and Audits
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
6.1 | Internal audits Are internal audits conducted at intervals as defined in the quality manual and do these audits address areas important to patient care? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 4.14.5 Note: The cycle for internal auditing should normally be completed in one year. The laboratory must conduct internal audits at planned intervals to determine whether all activities in the quality management system, including pre-examination, examination, and post-examination.
|  | a) Is there an audit plan/schedule that ensures all activities of the QMS are audited? | [5 empty] | 100
|  | b) Are audits being carried with minimal conflict of interest e.g. where possible, carried out by persons who are not involved in lab activities in the section being audited? | [5 empty] | 100
|  | c) Are the personnel conducting the internal audits trained and competent in auditing? | [5 empty] | 100
|  | d) Is root cause analysis performed for nonconformities/noted deficiencies? | [5 empty] | 100
|  | e) Are internal audit findings documented and presented to the laboratory management and relevant staff for review? | [5 empty] | 100
|  | Answer to TB question G6.1 | [5 empty] | 100 | 600
6.2 | Audit Recommendations and Action Plan & Follow Up |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.10; 4.13; 4.14.5 Note: For actions that are not implemented as per the due dates there should be a motivation and an approval of extension.
|  | a) Are internal audits reports generated? | [5 empty] | 100
|  | b) Are recommendations for corrective/preventive actions made based on audit findings? | [5 empty] | 100
|  | c) Is an action plan developed with clear timelines, assigned personnel & documented follow-up within the timeframe defined by the laboratory? | [5 empty] | 100 | 300
6.3 | Risk Management Are assessment of potential pitfalls performed for all laboratory processes including pre examination, examination and post examination? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 4.14.6 The Laboratory shall assess all steps in for all its processes (pre-analytical, analytical and post analytical) for areas of potential pitfalls e.g. pre-analytical step of sample collection, potential pitfalls could be; wrong sample collected, sample collected in wrong container, sample collected at wrong time. Post analytical could be; result sent to wrong patient, results sent outside of TAT. The Lab must assess all steps, list potential pitfalls and document action taken to prevent these from occurring. Note: Risks should be graded and acted upon as per their grading.
|  | a) Documented assessment of potential pitfalls for all processes | [5 empty] | 100
|  | b) Documented actions taken to reduce or eliminate identified potential pitfalls | [5 empty] | 100 | 200
| Totals |  | 15 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S7
Section 7 - Purchasing & Inventory
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
7.1 | Inventory and Budgeting System Is there a system for accurately forecasting needs for supplies and reagents? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.2.1(i); 5.3.2.1; 5.3.2.4 The laboratory must have a systematic way of determining its supply and testing needs through inventory control and budgeting systems that take into consideration past patterns, present trends and future plans.
7.2 | Specifications for supplies and consumables |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.6 Note: Specification could be in the form of catalogue number; item number, manufacturer name etc.
|  | Does the laboratory provide specification for their supplies and consumables that are required when placing a requisition? | [5 empty] | 100
|  | Answer to TB question G7.1 | [5 empty] | 100 | 200
7.3 | Service Supplier Performance Review Does the lab monitor the performance of the suppliers to ensure that the stated criteria are met? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.6 All suppliers of services used by the laboratory must be reviewed for their performance. Those that perform well must be identified and listed as approved suppliers. Results of these reviews must be documented.
7.4 | Inventory Control Does the lab maintain records for each reagent and consumable that contributes to the performance of examinations. These records shall include but not be limited to the following: |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.3.2.7; 5.3.2.4 Note: All incoming orders should be inspected for condition and completeness of the original requests, receipted and documented appropriately; the date received in the laboratory and the expiry date for the product should be clearly indicated.
|  | a) Identity of the reagent or consumable | [5 empty] | 100
|  | b) Batch code or lot number | [5 empty] | 100
|  | c) Manufacturer or supplier name and contact information | [5 empty] | 100
|  | d) Date of receiving, the expiry date, date of entering into service and, where applicable, the date the material was taken out of service | [5 empty] | 100
|  | e) Manufacturer's instruction/package insert | [5 empty] | 100
|  | f) Records of inspection of reagents and consumables when received (e.g. acceptable or damaged) | [5 empty] | 100 | 600
7.5 | Budgetary Projections Are budgetary projections based on personnel, test, facility and equipment needs as well as quality assurance procedures and materials? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(a) Note: Budgetary projections will ensure that there are no disruptions to services provided
7.6 | Management Review of Supply Requests Does management review/approve the finalized supply requests? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.2.3; 5.3.2.7 Note: Due to the fact that labs have different purchasing approval systems, there should be a system in place that the lab reviews final approval of their original request.
7.7 | Laboratory Inventory System |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.2 Note: The laboratory inventory system should reliably inform staff of the minimum amount of stock to be kept in order to avoid interruption of service due to stock-outs and the maximum amount to be kept by the laboratory to prevent expiry of reagents.
|  | a) Are inventory records complete and accurate, with minimum and maximum stock levels denoted and monitored? | [5 empty] | 100
|  | b) Is the consumption rate of all reagents and consumables monitored? | [5 empty] | 100
|  | c) Are stock counts routinely performed? | [5 empty] | 100 | 300
7.8 | Storage Area Are storage areas set up and monitored appropriately? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.2.2 Note: Storage of supplies and consumables must be as per the manufacturer’s specifications.
|  | a) Is the storage area well-organized and free of clutter? | [5 empty] | 100
|  | b) Are there designated places for all inventory items for easy access? | [5 empty] | 100
|  | c) Is adequate cold storage available? | [5 empty] | 100
|  | d) Are storage areas monitored as per prescribed storage conditions? | [5 empty] | 100
|  | e) Is the ambient temperature monitored routinely? | [5 empty] | 100
|  | f) Is storage in direct sunlight avoided? | [5 empty] | 100
|  | g) Is the storage area adequately ventilated? | [5 empty] | 100
|  | h) Is the storage area clean and free of dust and pests? | [5 empty] | 100
|  | i) Are storage areas access-controlled? | [5 empty] | 100
|  | Answer to TB question G7.2 | [5 empty] | 100 | 1000
7.9 | Inventory Organization and Watage Minimization Is First-Expiry-First-Out (FEFO) practised? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.2.2 and USAID Deliver Project, Logistics Handbook, Task Order 1 To minimize wastage from product expiry, inventory should be organized in line with the First-Expiry-First-Out (FEFO) principle. Place products that will expire first in front of products with a later expiry date and issue stock accordingly to ensure products in use are not past their expiry date. Remember that the order in which products are received is not necessarily the order in which they will expire.
7.10 | Product Expiration |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.2.3 All reagents and test kits in use, as well as those in stock, should be within the manufacturer-assigned expiry dates. Expired stock should not be entered into use and should be documented before disposal.
|  | Are all reagents/test kits in use (and in stock) currently within the manufacturer-assigned expiration dates or within stability? | [5 empty] | 100
|  | Answer to TB question S7.1 | [5 empty] | 100
|  | Answer to TB question C7.1 | [5 empty] | 100
|  | Answer to TB question D7.1 | [5 empty] | 100
|  | Answer to TB question X7.1 | [5 empty] | 100
|  | Answer to TB question T7.1 | [5 empty] | 100
|  | Answer to TB question L7.1 | [5 empty] | 100
|  | Answer to TB question P7.1 | [5 empty] | 100
|  | Answer to TB question N7.1 | [5 empty] | 100 | 900
7.12 | Disposal of Expired Products Are expired products labeled and disposed of properly? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.2.7 Expired products should be disposed of properly and records maintained. If safe disposal is not available at the laboratory, the manufacturer/supplier should take back the expired stock at the time of their next delivery.
7.13 | Laboratory Testing Services Has the laboratory provided uninterrupted testing services, with no disruptions due to stock-outs in the last year or since last audit? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.1.1.4(a);(n); 4.1.2.1(i); 5.5 Note: Interruption to services is considered when a laboratory cannot release results to their users. Testing services should not be subject to interruption due to stock-outs. Laboratories should pursue all options for borrowing stock from another laboratory or referring samples to another testing facility while the stock-out is being addressed.
| Totals |  | 24 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S8
Section 8 - Process Control
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
8.1 | Information for patients and users Are guidelines for patient identification, specimen collection (including client safety), labeling and transport readily available to persons responsible for primary sample collection? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.4.1 Note: The laboratory shall have documented procedures and information for pre-examination activities to ensure the validity of the results of examinations and must make these available to those who collect samples.
8.2 | Does the laboratory adequately collect information needed for examination performance? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.4; 5.4.3 Note: Each request accepted by the laboratory for examination(s) shall be considered an agreement. The request may be in the form of a hard copy or electronically.
|  | a) Are all test requests accompanied by an acceptable and approved test requisition form (and a transmittal sheet/checklist/manifest where applicable)? | [5 empty] | 100
|  | b) Does the request form have patient ID including gender, date of birth, location of patient and unique identifier? | [5 empty] | 100
|  | c) Name, signature or initials of authorized requester | [5 empty] | 100
|  | d) Type of sample and examination requested | [5 empty] | 100
|  | e) Clinically relevant information | [5 empty] | 100
|  | f) Date of sample collection (And time of collection where relevant – where time has an impact on the result) | [5 empty] | 100
|  | g) Date and time of sample receipt | [5 empty] | 100
|  | h) Written consent for invasive procedures with increased risk of complications | [5 empty] | 100
|  | Answer to TB question G8.1 | [5 empty] | 100
|  | Answer to TB question G8.2 | [5 empty] | 100 | 1000
8.3 | Are adequate sample receiving procedures in place? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.4; 5.4.6 Note: The review of service agreements occurs on sample reception. All portions of the primary sample must be unequivocally traceable to the original primary sample.
|  | a) Patient Unique Identifier? | [5 empty] | 100
|  | b) Are received specimens evaluated according to acceptance/rejection criteria? | [5 empty] | 100
|  | c) Are specimens logged appropriately upon receipt in the laboratory (including date, time, and name of receiving officer)? | [5 empty] | 100
|  | d) Are procedures in place to process “urgent” specimens and verbal requests? | [5 empty] | 100
|  | e) When samples are split, can the portions be traced back to the primary sample? | [5 empty] | 100
|  | f) If not a 24-hour lab, is there a documented method for handling specimens received after hours? | [5 empty] | 100
|  | g) Are specimens delivered to the correct workstations in a timely manner? | [5 empty] | 100 | 700
8.4 | Pre-examination Handling, Preparation and Storage Where testing does not occur immediately upon arrival in the laboratory, are specimens stored appropriately prior to testing? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.4.7 Note: Specimens should be stored under the appropriate conditions to maintain the stability of the specimen.
8.5 | Sample Transportation Are specimens, either received or referred, packaged appropriately according to local and/or international regulations and transported within acceptable timeframes and temperature intervals? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.4.4.3; 5.4.5 Note: All samples should be transported to the laboratory in a manner that is safe to the patients, the public and the environment. The laboratory must ensure that the samples were received within a temperature interval specified for sample collection.
8.6 | Does the laboratory select and evaluate referral Labs and Consultants? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 4.5 Note: The laboratory must have system in place to ensure that the referral laboratories are competent to perform the services required. Evaluations in the form of checking their accreditation status, using a questionnaire, performing audits, use of blinded samples etc.
|  | a) Are there documented reviews and evaluations of referral laboratories and consultants as defined by the laboratory? | [5 empty] | 100
|  | b) Is there a register of referral Laboratories and consultants? | [5 empty] | 100
|  | c) Are referred specimens tracked properly using a logbook, tracking form or electronically? | [5 empty] | 100 | 300
8.7 | Documentation of Examination Procedures Are examination procedures documented in a language commonly understood by all staff and available in appropriate locations? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.5.3 Note: examination procedures are for the laboratory staff to use therefore it should be in the language that is commonly understood by the staff; the lab may translate the documents into other languages which must be document controlled.
8.8 | Reagents Acceptance Testing |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.3.2.3 Note: This may be accomplished by a comparison study or examining quality control samples and verifying that results are acceptable.
|  | Is each new reagent preparation, lot number, new shipment of reagents, or consumable verified before use and documented? | [5 empty] | 100
|  | Answer to TB question S8.1 | [5 empty] | 100
|  | Answer to TB question C8.1 | [5 empty] | 100
|  | Answer to TB question D8.1 | [5 empty] | 100
|  | Answer to TB question X8.1 | [5 empty] | 100
|  | Answer to TB question L8.1 | [5 empty] | 100 | 600
8.9 | Quality Control |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.6.2 The laboratory shall design internal quality control systems that verify the attainment of the intended quality of results.
|  | Is internal quality control performed, documented and verified for all tests/procedures before releasing patient results? | [5 empty] | 100
|  | Answer to TB question P8.1 | [5 empty] | 100
|  | Answer to TB question P8.2 | [5 empty] | 100
|  | Answer to TB question P8.3 | [5 empty] | 100
|  | Answer to TB question N8.1 | [5 empty] | 100 | 500
8.10 | Quality Control Data Are QC results monitored and reviewed (biases, shifts, trends and Levy-Jennings charts)? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.6.2.3 Note: The lab must document and implement a system it would use to evaluate patient results since the last successful quality control; the evaluation could be done by re-examining selected samples of various batches, re-examining samples as per the stability of the Quality Control etc.
|  | a) Is there documentation of timely corrective action taken when quality control results exceed the acceptable range or reviews identify nonconformities? | [5 empty] | 100
|  | b) Does the Lab evaluate the results from the patient samples that were examined after the last successful quality control event | [5 empty] | 100
|  | Answer to TB question G8.3 | [5 empty] | 100
|  | Answer to TB question G8.4 | [5 empty] | 100
|  | Answer to TB question S8.2 | [5 empty] | 100
|  | Answer to TB question C8.2 | [5 empty] | 100
|  | Answer to TB question C8.3 | [5 empty] | 100
|  | Answer to TB question C8.4 | [5 empty] | 100
|  | Answer to TB question C8.5 | [5 empty] | 100
|  | Answer to TB question C8.6 | [5 empty] | 100
|  | Answer to TB question C8.7 | [5 empty] | 100
|  | Answer to TB question D8.2 | [5 empty] | 100
|  | Answer to TB question D8.3 | [5 empty] | 100
|  | Answer to TB question D8.4 | [5 empty] | 100
|  | Answer to TB question X8.2 | [5 empty] | 100
|  | Answer to TB question T8.1 | [5 empty] | 100
|  | Answer to TB question T8.2 | [5 empty] | 100
|  | Answer to TB question T8.3 | [5 empty] | 100
|  | Answer to TB question L8.2 | [5 empty] | 100
|  | Answer to TB question P8.4 | [5 empty] | 100
|  | Answer to TB question P8.5 | [5 empty] | 100
|  | Answer to TB question N8.2 | [5 empty] | 100 | 2200
8.11 | Comparability of Examination Results Does the laboratory compare results of the same test performed with different procedures and equipment? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.6.4 Note: The lab should document and implement a system to ensure there is comparability of results, this could be done by the use of EQA performance; using blinded samples, parallel testing.
|  | a) Where there is more than one procedure for the same measure, does the laboratory compare results from the different procedures, equipment or methods? | [5 empty] | 100
|  | b) Does the lab discuss, document and act upon (including notifying users) problems or deficiencies from these comparison studies? | [5 empty] | 100 | 200
8.12 | Are environmental conditions checked and reviewed accurately? Are the following environmental conditions checked daily? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.4.1 Note: The laboratory shall monitor, control and record environmental conditions, as required by relevant specifications or where they may influence the quality of the results and/or the health of staff.
|  | a) Room temperature | [5 empty] | 100
|  | b) Freezer(s) | [5 empty] | 100
|  | c) Refrigerator(s) | [5 empty] | 100
|  | d) Incubator(s) | [5 empty] | 100
|  | e) Water bath(s) | [5 empty] | 100 | 500
8.13 | Have acceptable ranges been defined for all temperature-dependent equipment with procedures and documentation of action taken in response to out-of-range temperatures? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.2(c) Note: Acceptable ranges should take into consideration manufacturers’ recommendations and requirements.
8.14 | Does the laboratory participate in inter-laboratory comparison program or alternative assessment systems for all tests? |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.6.3 Note: The laboratory should handle, analyze, review and report results for proficiency testing in a manner similar to regular patient testing. Investigation and correction of problems identified by unacceptable proficiency testing should be documented. Acceptable results showing bias or trends suggest that a problem should also be investigated.
|  | a) Do samples come from providers who are accredited or approved? | [5 empty] | 100
|  | b) Are specimens handled and tested the same way as patient specimens? | [5 empty] | 100
|  | c) Is the performance of the laboratory in the PT program reviewed and discussed with relevant staff? | [5 empty] | 100
|  | d) Is cause analysis performed for unacceptable results? | [5 empty] | 100
|  | e) Is corrective action documented for unacceptable results? | [5 empty] | 100
|  | Answer to TB question G8.5 | [5 empty] | 100 | 600
| Totals |  | 32 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S9
Section 9 - Information Management
Question no. | Question |  | Avaliable Score | Answer | Score | Comments | Standard
9.1 | Test Result Reporting System Are test results legible, technically verified by an authorized person and confirmed against patient identity? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.8.1 Results must be written in ink and written clearly with no mistakes in transcription. Cancellation must follow Good Lab Practices. The persons performing the test must indicate verification of the results. There must be a signature or identification of the person authorizing the release of the report.
9.1 | Testing Personnel Are testing personnel identified on the result report or other records (manual or electronic)? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13 ; 5.5.1.1; 5.8.1 Note: The person who performed the procedure must be identified on the report (hard copy or electronic) purposes of traceability.
9.2 | Report Content Does the laboratory report contain at least the following: |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.8.2; 5.8.3; 5.9.3 Note: When the reporting system cannot capture amendments, changes or alterations, a record of such shall be kept.
|  | a) Test requested | [5 empty] | 100
|  | b) Identification of the laboratory | [5 empty] | 100
|  | c) Identification of all examinations performed by a referral laboratory | [5 empty] | 100
|  | c) Patient identification and location | [5 empty] | 100
|  | d) Name of the requester | [5 empty] | 100
|  | e) Date of primary sample collection (and time, if relevant to patient care) | [5 empty] | 100
|  | f) type of primary sample | [5 empty] | 100
|  | g) Is the result reported in SI units where applicable? | [5 empty] | 100
|  | h) Biological reference intervals where applicable | [5 empty] | 100
|  | i) Is there space for interpretation or comments of results, when applicable? | [5 empty] | 100
|  | j) Identification of the person(s) reviewing and authorizing the report | [5 empty] | 100
|  | k) Date and time of the report | [5 empty] | 100
|  | l) Page number to total number of pages (e.g. "Page 1 of 5", "Page 2 of 5", etc.) | [5 empty] | 100
|  | m) When issuing revised reports, is it clearly identified as a revision and includes reference to the date and patient's identity in the original report and the user made aware of the revision? | [5 empty] | 100
|  | n) Does the revised record show the time and date of the change and the name of the person responsible for the change? | [5 empty] | 100
|  | o) Does the original report entry remain in the record when revisions are made? | [5 empty] | 100 | 1600
9.3 | Analytic System/Method Tracing |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13(g) Note: There must be traceability of specimen results to a specific analytical system or method. Proficiency testing specimens would also fall under specimen results.
|  | When more than one instrument is in use for the same test, are test results traceable to the equipment used for testing? | [5 empty] | 100
|  | Answer to TB question G9.1 | [5 empty] | 100 | 200
9.4 | Archived data Labeling and Storage Are archived results (paper or data-storage media) properly labeled and stored in a secure location accessible only to authorized personnel? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.13; 5.10.3 All patient data, paper, tapes, disks should be properly labeled and stored securely in places accessible only to authorized personnel. ISO 15189: 5.8.3 Annex B 6.4.
9.5 | Authorities and Responsibilities Has the laboratory defined and implemented authorities and responsibilities for the management and use of the laboratory information system– paper based and electronic, including maintenance and modifications that may affect patient care? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.9; 5.10.2; 5.10.3 Note:"information systems" includes the management of data and information contained in both computer and non-computerized systems. Some of the requirements may be more applicable to computer systems than to non-computerized systems. Computerized systems can include those integral to the functioning of laboratory equipment and standalone systems using generic software, such as word processing, spreadsheet and database applications that generate, collate, report and archive patient information and reports.
Is the following in place and implemented?
|  | a) Controlled access to patient data and information | [5 empty] | 100
|  | b) Controlled access to enter patient data and examination results | [5 empty] | 100
|  | c) Controlled access to changing patient data or examination results | [5 empty] | 100
|  | d) Controlled access to the release of examination results and reports | [5 empty] | 100
|  | e) Verification that results that have been transmitted electronically or reproduced external to the laboratory (computers, fax machines, email and websites and personal web devices) are correct. | [5 empty] | 100 | 500
9.6 | Information Management System Does the laboratory have evidence of how the Laboratory Information Management System (LIMS) was selected? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.1.1 Note: The laboratory must have a documented procedure and records for the selection, purchasing and management of equipment.
9.7 | Test Result Are test results validated, interpreted and released by appropriately authorized personnel? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.1; 5.8; 5.10.3; 5.9.1 Note: There must be a signature or identification of the person authorizing the release of the report.
9.8 | Verification of Electronic Laboratory Information System |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.13; 5.10.3 Note: The lab must perform verification of system after upgrades and to ensure previously stored patient results have not been affected.
If not applicaple select N/A for all sub questions
|  | a) Has the system been verified before implementation that include the verification reports to check functioning and inter-phasing by the laboratory? | [5 empty] | 100
|  | b) Are records of the validation by the supplier available and approved for use? | [5 empty] | 100
|  | c) Are reports of ongoing system checks available for correct transmissions, calculations and storage of results and records. | [5 empty] | 100 | 300
9.9 | Is the Laboratory Information System properly maintained to ensure continued functioning: |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.10.3 Note: If the LIS is maintained offsite, records of maintenance must be readily available .The lab should include the LIS as part of their internal audit.
If not applicaple select N/A for all sub questions
|  | a) Documented regular service by authorized and trained personnel | [5 empty] | 100
|  | b) Documented system failures with documented appropriate root cause analysis, corrective actions and preventive actions | [5 empty] | 100
|  | c) System operated in an environment recommended by the supplier for optimal functioning | [5 empty] | 100 | 300
| Totals |  | 21 |  | 0
Yes
Partial
No
N/A

## SLIPTA_S10
Section 10 - Identification of Non Conformities, Corrective and Preventive Actions
Question no. | Question |  | Avaliable Score | Answer | Score | Comments | Standard
10.1 | Are all identified nonconforming activities/ work identified and documented adequately |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.9 Note: nonconformities should be identified and managed in any aspect of the quality management system, including pre-examination, examination or post-examination processes. Nonconforming examinations or activities occur in many different areas and can be identified in many different ways, including clinician complaints, internal quality control indications, and instrument calibrations, checking of consumable materials, inter-laboratory comparisons, staff comments, reporting and certificate checking, laboratory management reviews, and internal and external audits.
|  | a) Indicating details of what happened, when, reporting person | [5 empty] | 100
|  | b) Immediate actions being taken | [5 empty] | 100
|  | c) Determination of the extent of the non- conformity? | [5 empty] | 100
|  | d) Are examinations halted and results withheld or recalled where the non-conformity compromises patient results? | [5 empty] | 100
|  | e) Informing the requester where the non-conformity has an effect on the management of the patient | [5 empty] | 100
|  | f) Authorization of resumption of testing documented (where testing has been halted) | [5 empty] | 100
|  | Answer to TB question G10.1 | [5 empty] | 100 | 700
10.2 | Root cause analysis |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.10(b) Note: Root cause analysis is a process of identifying and removing the underlying factor of the non-conformance.
|  | Is documented root cause analysis performed on nonconformities before corrective actions are implemented? | [5 empty] | 100
|  | Answer to TB question G10.2 | [5 empty] | 100 | 200
10.3 | Corrective action |  | 3 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.10; 4.13; 4.14.5 Note: Documenting corrective action allows the lab to review its effectiveness and to perform trend analysis for continual improvement.
|  | Is corrective action performed and documented for nonconformities? | [5 empty] | 100
|  | Answer to TB question X8.3 | [5 empty] | 100 | 200
10.4 | Are implemented corrective actions monitored and reviewed for their effectiveness before closure/clearance? |  | 3 |  | 0 |  | ISO15189:2012 Clause 4.10(f) Note: Implemented corrective action does not imply effectiveness; therefore the lab has to monitor to ensure that the NC has not recurred.
10.5 | Preventive Actions Are documented preventive actions implemented and monitored for their effectiveness? |  | 5 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 4.11; 4.12; Note: Preventive action should be an ongoing process involving analysis of laboratory data, including trend and risk analyses and external quality assessment (proficiency testing).
|  | a) Reviewing of laboratory data and information to determine potential non conformities | [5 empty] | 100
|  | b) Determining root causes for potential non conformities | [5 empty] | 100
|  | c) Implementing and documenting preventive actions | [5 empty] | 100
|  | d) Reviewing and documenting effectiveness of preventive actions | [5 empty] | 100 | 400
| Totals |  | 19 |  | 0
Yes
Partial
No

## SLIPTA_S11
Section 11 - Occurrence/Incident Management & Process Improvement
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
11.1 | Are graphical tools (charts and graphs) used to communicate quality findings and identify trends? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.12; 4.13; 4.14 Use of graphical displays of quality data communicates more effectively than tables of numbers. Examples of graphical tools commonly used for this purpose include Pareto charts, cause-and-effect diagrams, frequency histograms, trend graphs, and flow charts.
11.2 | Quality Management System Improvement Measures Does the laboratory identify and undertake continual quality improvement projects? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.12; 4.15 Note: The lab should use its management review activities to continually improve its quality management system by comparing its actual performance to its intentions stated in the quality policy and objectives.
11.3 | Communication System on Laboratory Operations Does the laboratory communicate with upper management regularly regarding needs for continual improvement? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.15.2 (o) Note: The laboratory staff should give input for management meetings.
11.4 | Quality indicators |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause4.12; 4.14.7 Note: The lab should select QI in line with meeting its objectives from pre-analytic, analytic and post-analytic phases critical to patient outcomes.
|  | Are quality indicators (TAT, rejected specimens, stock-outs, etc.) selected, tracked and reviewed? | [5 empty] | 100
|  | Answer to TB question G11.1 | [5 empty] | 100
|  | Answer to TB question G11.2 | [5 empty] | 100
|  | Answer to TB question G11.3 | [5 empty] | 100
|  | Answer to TB question S11.1 | [5 empty] | 100
|  | Answer to TB question C11.1 | [5 empty] | 100
|  | Answer to TB question D11.1 | [5 empty] | 100
|  | Answer to TB question X11.1 | [5 empty] | 100
|  | Answer to TB question T11.1 | [5 empty] | 100
|  | Answer to TB question L11.1 | [5 empty] | 100
|  | Answer to TB question P11.1 | [5 empty] | 100
|  | Answer to TB question N11.1 | [5 empty] | 100 | 1200
11.5 | Is the outcome of the review of quality indicators used to improve lab performance? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.14.7; 4.15.2(f) Note: The lab should review the QI to ensure its continued appropriateness.
11.6 | Are the actions taken checked and monitored to determine the effectiveness of improved quality of lab performance? |  | 2 |  | 0 |  | ISO15189:2012 Clause 4.14.7 Note: the lab should create an action plan to monitor the QI stating the objectives, methodology, interpretation, limits, action plan and duration of measurement for each QI.
| Totals |  | 12 |  | 0
Yes
Partial
No

## SLIPTA_S12
Section 12 - Facilities and Safety
Question no. | Question |  | Available Score | Answer | Score | Comments | Standard
12.1 | Is there documented evidence that the laboratory has evaluated the adequacy of the size and overall layout of the laboratory and organized the space so that workstations are positioned for optimal workflow? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.1 Note: Documentation could be in the form of a floor plan, results from internal audits, etc.
12.2 | Are the patient care and testing areas of the laboratory distinctly separate from one another? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.1 Client service areas (i.e. waiting room, phlebotomy room) should be distinctly separate from the testing areas of the laboratory. Client access should not compromise “clean” areas of the laboratory. For biosafety reasons, microbiology and TB testing should be segregated in a separate room(s) from the general laboratory testing.
12.3 | Is each individual workstation maintained free of clutter and set up for efficient operation? |  | 2 | Answer all sub questions | 0 |  | ISO15190 Clause 6.3.5
Are the following criteria met?
|  | a) Does the equipment placement/layout facilitate optimum workflow? | [5 empty] | 100
|  | b) Are all needed supplies present and easily accessible? | [5 empty] | 100
|  | c) Are the chairs/stools at the workstations appropriate for bench height and the testing operations being performed? | [5 empty] | 100 | 300
12.4 | Is the physical work environment appropriate for testing? |  | 2 | Answer all sub questions | 0 |  | ISO15189:2012 Clause 5.2 Note: The laboratory space should be sufficient to ensure the quality of work, safety of personnel and the ability of staff to carry out their tasks without compromising the quality of the examinations. The laboratory should be clean and well organized, free of clutter, well ventilated, adequately lit and within acceptable temperature ranges.
|  | a) Free of clutter? |  |  |  |  | ISO 15190:2003 Clause 13.0 | 100
|  | b) Adequately ventilated? |  |  |  |  | ISO 15190:2003 Clause 6.3.3 | 100
|  | c) Adequately lit? |  |  |  |  | ISO 15190:2003 Clause 6.3.1 | 100
|  | d) Climate-controlled for optimum equipment function? |  |  |  |  | ISO 15190:2003 Clause 6.3.2 | 100
|  | e) Are filters checked, cleaned and/or replaced at regular intervals, where air-conditioning is installed? | [5 empty] | 100
|  | f) Are wires and cables properly located and protected from traffic? | [5 empty] | 100
|  | g) Is there a functioning back-up power supply (generator)? | [5 empty] | 100
|  | h) Is critical equipment supported by uninterrupted power source (UPS) systems? | [5 empty] | 100
|  | i) Is equipment placed appropriately (away from water hazards, out of traffic areas)? | [5 empty] | 100
|  | j) Are appropriate provisions made for adequate water supply, including deionized or distilled water, if needed? | [5 empty] | 100
|  | k) Is clerical work completed outside the testing area? | [5 empty] | 100
|  | l) Is major safety signage posted and enforced, including NO EATING, SMOKING, DRINKING? | [5 empty] | 100
|  | Answer to TB question G12.2 | [5 empty] | 100 | 1300
12.5 | Laboratory Access Is the laboratory properly secured from unauthorized access with appropriate signage? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.2 Note: Access control should take into consideration safety, confidentiality, and quality.
12.6 | Laboratory Storage Areas Is laboratory-dedicated cold and room temperature storage free of staff food items, and are patient samples stored separately from reagents and blood products in the laboratory refrigerators and freezers? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2; 5.2.4 Note: there should be effective separation to prevent contamination.
12.7 | Cleaning & disinfection Is the work area clean and free of leakage and spills, and are disinfection procedures conducted and documented? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.6 Note: The work area should be cleaned regularly. An appropriate disinfectant should be used. At a minimum, all bench tops and working surfaces should be disinfected at the beginning and end of every shift. All spills should be contained immediately and the work surfaces disinfected.
12.8 | Biosafety Cabinet |  | 2 | Answer all sub questions | 0 |  | ISO 15189:2012 Clause 5.2.1; 5.2.2 Note: A biosafety cabinet should be used to prevent aerosol exposure to contagious specimens or organisms. For proper functioning and full protection, biosafety cabinets require periodic maintenance and should be serviced accordingly. Biosafety cabinet should be recertified according to national protocol or manufacturer requirements.
|  | Where a Biosafety cabinet is required to perform work, is it certified and appropriate? | [5 empty] | 100
|  | Answer to TB question G12.3 | [5 empty] | 100 | 200
12.9 | Laboratory Safety Manual Is a laboratory safety manual available, accessible and up-to-date? |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 7.4 A safety manual should be readily available in work areas as required reading for all employees. The manual should be specific to the laboratory's needs; it should be reviewed and updated at least annually by laboratory management.
Does the safety manual include guidelines on the following topics?
|  | a) Blood and body fluid precautions | [5 empty] | 100
|  | b) Hazardous waste disposal | [5 empty] | 100
|  | c) Hazardous chemicals/materials | [5 empty] | 100
|  | d) MSDS sheets | [5 empty] | 100
|  | e) Personal protective equipment | [5 empty] | 100
|  | f) Vaccination | [5 empty] | 100
|  | g) Post-exposure prophylaxis | [5 empty] | 100
|  | h) Fire safety | [5 empty] | 100
|  | i) Electrical safety | [5 empty] | 100 | 900
12.10 | Waste disposal Is sufficient waste disposal capacity available and is waste separated into infectious and non-infectious waste, with infectious waste autoclaved/incinerated? |  | 2 |  | 0 |  | ISO 15190: 22 Waste should be separated according to biohazard risk, with infectious and non-infectious waste disposed of in separate containers. Infectious waste should be discarded into containers that do not leak and are clearly marked with a biohazard symbol. Sharp instruments and needles should be discarded in puncture resistant containers. Both infectious waste and sharps containers should be autoclaved before being discarded to decontaminate potentially infectious material. To prevent injury from exposed waste, infectious waste should be incinerated, burnt in a pit or buried.
12.11 | Hazardous chemicals Are hazardous chemicals/materials properly handled? |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 17.1 and 17.3 All hazardous chemicals must be labeled with the chemical’s name and with hazard markings clearly indicated. Flammable chemicals must be stored out of sunlight and below their flashpoint, preferably in a still cabinet in a well-ventilated area. Flammable and corrosive agents should be separated from one another. Distinct care should always be taken to handle hazardous chemicals safety in the workplace.
|  | a) Are hazardous chemicals properly labelled? | [5 empty] | 100
|  | b) Are hazardous chemicals properly stored to ensure safety and prevent theft? | [5 empty] | 100
|  | c) Are hazardous chemicals properly utilized according to MSDS? | [5 empty] | 100
|  | d) Are hazardous chemicals properly disposed according to national guidelines or MSDS? | [5 empty] | 100 | 400
12.12 | Handling of sharps Are ‘sharps’ handled and disposed of properly in ‘sharps’ containers that are appropriately utilized? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.2.3 All syringes, needles, lancets or other bloodletting devices capable of transmitting infection must be used only once and discarded in puncture resistant containers that are not overfilled. Sharps containers should be clearly marked to warn handlers of the potential hazard and should be located in areas where sharps are commonly used.
12.13 | Fire safety Is fire safety included as part of the laboratory’s overall safety programme? |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 19.7 and 9.3 Electrical cords and plugs, power-strips and receptacles should be maintained in good condition and utilized appropriately. Overcrowding should be avoided and cords should be kept out of walkway areas. An approved fire extinguisher should be easily accessible within the laboratory and be routinely inspected and documented for readiness. Fire extinguishers should be kept in their assigned place and not hidden or blocked; the pin and seal should be intact, nozzles should be free of blockage, pressure gauges should show adequate pressure, and there should be no visible signs of damage. A fire alarm should be installed in the laboratory and tested regularly for readiness; all staff should participate in periodic fire drills.
|  | a) Are all electrical cords, plugs and receptacles used appropriately and in good repair? | [5 empty] | 100
|  | b) Is an appropriate fire extinguisher available, properly placed, in working condition and routinely inspected? | [5 empty] | 100
|  | c) Is an operational fire warning system in place? | [5 empty] | 100
|  | d) Are periodic fire drills conducted at defined intervals? | [5 empty] | 100 | 400
12.14 | Safety Audits Are safety inspections or audits conducted regularly and documented? |  | 3 | Answer all sub questions | 0 |  | ISO15190 Clause 7.3.1 and 7.3.2 Note: The safety programme shall be audited and reviewed at least annually (by appropriately trained personnel.
|  | a) Is there an audit plan/schedule that ensures all activities of the lab are checked for safety compliance? | [5 empty] | 100
|  | b) Are inspections/audits being carried out by authorized persons? | [5 empty] | 100
|  | c) Are the personnel conducting the internal audits trained in safety? | [5 empty] | 100
|  | d) Is root cause analysis performed and action taken for nonconformities/noted deficiencies? | [5 empty] | 100
|  | e) Are safety findings documented and presented to the laboratory management and relevant staff for review? | [5 empty] | 100 | 500
12.15 | Safety Equipment Is standard safety equipment available and in use in the laboratory? |  | 2 | Answer all sub questions | 0 |  | ISO15190 Clause 5.1 Note: It is the responsibility of laboratory management to ensure that the laboratory is equipped with standard safety equipment. The list above is a partial list of necessary items. Biosafety cabinets should be in place and in use as required. All centrifuges should have covers. Hand-washing stations should be designated and equipped and eyewash stations (or an acceptable alternative method of eye cleansing) should be available and operable. Spill kits and first aid kits should be kept in a designated place and checked regularly for readiness.
|  | a) Biosafety cabinet(s) |  |  |  |  | ISO 15190:2003 Clause 16 | 100
|  | b) Covers, safety caps, safety buckets on centrifuge(s) | [5 empty] | 100
|  | c) Hand-washing station |  |  |  |  | ISO 15190:2003 Clause 12.7 | 100
|  | d) Eyewash station/bottle(s) and emergency showers where applicable |  |  |  |  | ISO 15190:2003 Clause 12.10 | 100
|  | e) Spill kit(s) | [5 empty] | 100
|  | f) First aid kit(s) |  |  |  |  | ISO 15190:2003 Clause 12.9 | 100 | 600
12.16 | Personal Protective Equipment |  | 2 | Answer all sub questions | 0 |  | ISO 15190: 12 Management is responsible for providing appropriate personal protective equipment (gloves, lab coats, eye protection, etc) in useable condition. Laboratory staff must utilize PPE at all times while in the laboratory. Protective clothing should not be worn outside the laboratory. Gloves should be replaced immediately when torn or contaminated and not washed for reuse.
|  | Is personal protective equipment (PPE) easily accessible at the workstation and utilized appropriately and consistently? | [5 empty] | 100
|  | Answer to TB question G12.1 | [5 empty] | 100 | 200
12.17 | Staff Vaccinations Are laboratory personnel offered appropriate vaccination and employee medical surveillance? |  | 2 |  | 0 |  | ISO 15190: 11.3 Laboratory staff should be offered appropriate vaccinations—particularly Hepatitis B. Staff may decline to receive the vaccination, but they must then sign a declination form to be held in the staff member’s personnel file.
12.18 | Post Exposure Prophylaxis Are post-exposure prophylaxis policies and procedures posted and implemented after possible and known exposures? |  | 2 |  | 0 |  | ISO 15190: 9 The laboratory must have a procedure for follow-up of possible and known percutaneous, mucus membrane or abraded skin exposure to HIV, HBV or HCV. The procedure should include clinical and serological evaluation and appropriate prophylaxis.
12.19 | Are adverse incidents or injuries from equipment, reagents, occupational injuries, medical screening or illnesses, documented and investigated? |  | 2 |  | 0 |  | ISO15189:2012 Clause 5.3.1.6; 5.3.2.6; ISO15190 Clause 9 Note: All occupational injuries or illnesses should be thoroughly investigated and documented in the safety log or occurrence log, depending on the laboratory. Corrective actions taken by the laboratory in response to an accident or injury must also be documented.
12.20 | Biosafety Training Are drivers/couriers and cleaners working with the laboratory trained in biosafety practices relevant to their job tasks? |  | 2 |  | 0 |  | ISO 15190: 10 All occupational injuries or illnesses should be thoroughly investigated and documented in the safety log or occurrence log, depending on the laboratory. Corrective actions taken by the laboratory in response to an accident or injury must also be documented.
12.21 | Laboratory Safety Officer Is a trained safety officer designated to implement and monitor the safety programme in the laboratory, including the training of other staff? |  | 2 |  | 0 |  | ISO 15190: 7, 10 A safety officer should be designated to work with the laboratory manager to implement the safety programme, monitor the ongoing safety conditions and needs of the laboratory, coordinate safety training, and serve as a resource for other staff. This officer should receive safety training.
| Totals |  | 43 |  | 0
Yes
Partial
No
N/A