## 1 - Instrument and tools
Item Number | Method | Family | Sub Family | Item Name | GDF SPECIFICATION | Detailled Specification (Y/N) | Request for sample (Y/N) | Link to | Importance | Unit / package GDF | Quantity to equip 1 lab | Supplier Specifications | Primary Packing Type | Number of units per primary packing type | Secondary Packing Type | Number of units per secondary packing type | Minimum Shelf Life in days | Gross weight in kg | Depth (in cm) | Width (in cm) | Height (in cm) | Volume in M3 | Special requirement for delivery | Special requirement for storage | Manufacturer Name | Model or Item Number (Manufacturer) | Country of Origin | Custom Tariff code | Voltage / Amp / Frequency | Please mention quality standard under which the item is produced (if any)
106045 | MOLECULAR BIOLOGY | Instrument | Hoods, Cabinets, Workstations | PCR-Workstation, UV light | PCR-Workstation, UV light, Glass with UV/Air Cleaner e.g. KISKER - L020-GC | Yes |  |  | Always needed | 1 | 2
106074 | MOLECULAR BIOLOGY | Instrument | Centrifuges | Microfuge / minispin for 0.2-2ml tubes for short-rapid spin down of reagents | Resistant to bleach |  |  |  | Always needed | 1 | 1
106080 | BIOSAFETY | Instrument | Laboratory, Miscellaneous | UV lamps, 30 W, 90 cm, incl. holder | Meant only for countries of the former Soviet Union |  |  | 106076 | Optional | 1 | 5
106081 | LIQUID CULTURE | Instrument | Laboratory, Miscellaneous | UPS - 3000 VA - 2100W | UPS with battery pack for 20 minutes at 2100W | Yes |  | 106013 | Always needed | 1 | 2
106491 | BIOSAFETY | Instrument | Laboratory, Miscellaneous | Battery Pack for UPS 3000 VA | Battery pack for 120 minutes at 500W | Yes |  |  | Regular needed | 1 | 1
106101 | CLEANING KITCHEN | Instrument | Disinfectants and Cleaning | Laboratory washer-disinfector | Optional: Laboratory washer-disinfector for laboratory glass including upper and lower baskets |  |  |  | Optional | 1 | 1
106102 | CLEANING KITCHEN | Instrument | Disinfectants and Cleaning | Washer-disinfector 1/2 insert | Washer-disinfector 1/2 insert with 28 small spring hooks |  |  | 106101 | Optional | 1 | 1
106103 | CLEANING KITCHEN | Instrument | Disinfectants and Cleaning | 1/2 insert for baskets | 1/2 insert for baskets |  |  | 106101 | Optional | 1 | 1
106104 | CLEANING KITCHEN | Instrument | Disinfectants and Cleaning | Upper basket/ injector for nozzles | Upper basket/ injector for nozzles |  |  | 106101 | Optional | 1 | 1
106105 | CLEANING KITCHEN | Instrument | Disinfectants and Cleaning | 1/4 insert, for test tubes 160x16 | 1/4 insert, for test tubes 160x16 |  |  | 106101 | Optional | 1 | 2
106107 | CLEANING KITCHEN | Instrument | Disinfectants and Cleaning | Separate water supply for deionised water | Separate water supply for deionised water with conductivity measure R 3/4" with digital display and hose-set for in-and outflow |  |  | 106101 | Optional | 1 | 1
106108 | CLEANING KITCHEN | Instrument | Disinfectants and Cleaning | Aqua purificator cabinet for two de-mineralising cartridges | Aqua purificator cabinet for two de-mineralising cartridges |  |  | 106101 | Optional | 1 | 1
106124 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Ion exchanger cartridge | Ion exchanger cartridge, S/S, flow rate approx. 950 L/h, complete set with tubing set with quick connection coupling and one spare set of tubing with quick connection coupling, pressure resistant; with an exhaust and excess pressure valve of 2.5 -10 bar |  |  |  | Optional | 1 | 2
106125 | MEDIA KITCHEN | Instrument | Phosphate buffer according to Soerensen | Conductivity measuring device | Conductivity measuring device for measurement of quality of laboratory grade water, digital |  |  |  | Optional | 1 | 2
106126 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Ion exchanger resin for laboratory grade water | Ion exchanger resin for laboratory grade water, pack of 5x10 L resin with filter bag for used up resin |  |  |  | Optional | 1 | 5
106127 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Fully automatic water distil | Fully automatic water distil, capacity (according to the need) 4 - (12) litres / hour, distilled water approx. 3 µS | Yes |  |  | Optional | 1 | 1
106128 | MEDIA KITCHEN | Instrument | Weighing & Analytical Balances | Electronic analytical balance | Semimicro balance (electronic analytical balance), EC stamp-approved version, with built-in motorized weight, dynamic temperature compensation, 8 segment display, stability indicator. Accuracy class I. Readability 0.01 mg up to 60 g and 0.1 mg up to approx. 200 g. Tara range (subtractive) 200 g. | Yes |  |  | Optional | 1 | 1
106129 | MEDIA KITCHEN | Instrument | Weighing & Analytical Balances | Electronic precision balance | Electronic precision balance, EC stamped version with 8 segment display, stability indicator, plus optimised display update rate and a built-in motorized, automatic calibration feature. Level indicator. Range 0.5 - 3200 g, readability 0,01 g, tara range (subtractive) 3200 g, accuracy class II | Yes |  |  | Optional | 1 | 1
106130 | MEDIA KITCHEN | Instrument | Meters | Bench top pH meter | Bench top pH meter with temperature compensation, complete set with electrode and 3 calibration buffers, resolution 0.01pH | Yes |  |  | Always needed | 1 | 1
106132 | MEDIA KITCHEN | Instrument | Mixers, Shakers, Stirrers | Magnetic stirrer with heating plate | Magnetic stirrer with heating plate | Yes |  |  | Regular needed | 1 | 1
106133 | MEDIA KITCHEN | Instrument | Mixers, Shakers, Stirrers | Magnetic bars set of 10 | Set of round shaped PTFE coated magnetic stirring bars, set of 10 magnetic bares, 1x 15 mm, 2X 20 mm, 2x 25 mm, 1x 30 mm, 2x 40 mm, 2x 50 mm length |  |  |  | Regular needed | 1 set of 10 | 1
106134 | MEDIA KITCHEN | Instrument | Disinfectants and Cleaning | Sterile filling station of culture media 1000ml | Sterile filling station of culture media, reservoir 1000 ml, (Squibb sedimentation funnel or cylindric dropping funnel, graduation, with tube for pressure exchange), with stop cock made of glass and NSPE stopper, outlet diameter 6mm |  |  |  | Optional | 1 | 4
106135 | MEDIA KITCHEN | Instrument | Disinfectants and Cleaning | Sterile filling station of culture media 250ml | Sterile filling station of culture media, reservoir 250 ml, (Squibb sedimentation funnel or cylindric dropping funnel, graduation, with tube for pressure exchange), with stop cock made of glass and NSPE stopper, outlet diameter 4 mm (several smaller funnels 250 ml to 500 ml for drug containing media) |  |  |  | Optional | 1 | 10
106136 | MEDIA KITCHEN | Instrument | Disinfectants and Cleaning | Automatic filling station | Automatic filling station (usually very sensitive to faults) consisting of a peristaltic pump for sterile dispensing. Four roller pump head, 1 - 350 rpm, accuracy ± 1%, easy load pump head, silicon tubing bore 1.6 mm (1/16'' inch), 15 m, minimum dose 3.5 ± 0.5 ml, maximum dose 15 ± 2 ml, alphanumerical menu in English and XYZ,, storage of minimal 5 dispensing programs, applying standards: IEC 335-1 EN 60529, CE, EN 60204-1, Low voltage directive 73/23/EEC and EN 61010, connection 230 ±10 V AC / 50 Hz |  |  |  | Optional | 1 | 1
106137 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Stative / Stand | Stative (heavy quality for LJ separation funnel) Set (plate+rod) and clamp |  |  |  | Optional | 1 | 3
106138 | MEDIA KITCHEN | Instrument | Incubators | Inspissator for coagulation of egg-based culture media | Inspissator for coagulation of egg-based culture media, capacity ≥ 350 tubes, 230 ± 10 V AC/ 50 Hz |  |  |  | Optional | 1 | 1
106140 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Laboratory Floor standing refrigerator 140L | Floor standing refrigerator, at least energy efficiency class A+ , approx. 150 litre | Yes |  |  | Always needed | 1 | 1
106142 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Deep Freezer 70L / - 50 °C to - 85 °C | Chest freezer, capacity approx. 70 litres, temperature range: - 50 °C to - 85 °C, complete with cassettes and racks (only for the laboratory preserving control strains for the country or for scientific purpose) |  |  |  | Optional | 1 | 1
106205 | LIQUID CULTURE | Instrument | Laboratory, Miscellaneous | UV-light illuminator - Liquid culture - 365 nm | UV-light illuminator at least 200 x 200 mm window |  |  |  | Regular needed | 1 | 1
106211 | MOLECULAR BIOLOGY | Instrument | Laboratory, Miscellaneous | Electrophoresis chamber | Electrophoresis, horizontal mini gel system for separation of nucleic acids fragments |  |  |  | Optional | 1 | 1
106212 | MOLECULAR BIOLOGY | Instrument | Laboratory, Miscellaneous | Power supply for electrophoresis | Power supply for electrophoresis, 0 - 300 V DC |  |  | 106211 | Optional | 1 | 1
106213 | MOLECULAR BIOLOGY | Instrument | Laboratory, Miscellaneous | UV-light illuminator - Ethidium bromide - 312 nm | UV-light illuminator at least 200 x 200 mm window - wave lenght 312 nm |  |  |  | Optional | 1 | 1
106219 | MOLECULAR BIOLOGY | Instrument | Mixers, Shakers, Stirrers | Minishaker / Vortex | Minishaker or Vortex with Plates for 0.2ml PCR Tubes or Cryo vial 1.5ml or 2ml | Yes |  |  | Always needed | 1 | 3
106221 | PROCESSING | Instrument | Laboratory, Miscellaneous | Safety gas burner | Safety gas burner with glass chimney, footswitch for temporary ignition and power supply | Yes |  |  | Optional | 1 | 2
106222 | PROCESSING | Instrument | Laboratory, Miscellaneous | Safety gas tubing | Safety gas tubing with bending preventing safety clamps on both ends, inner diameter 9mm, length 750 mm (or shorter, depending on the size of BSC) |  |  | 106222 | Optional | 1 | 2
106223 | SOLID CULTURE | Instrument | Laboratory, Miscellaneous | Electric micro incinerator for loops | Electric micro incinerator for loops |  |  |  | Optional | 1 | 2
106411 | MOLECULAR BIOLOGY | Instrument | Hoods, Cabinets, Workstations | UV Crosslinker | e.g. Spectrolinker XL1000 (254nm, 8watts, 230V power cord) |  |  |  | Optional | 1 | 1
106424 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Laboratory Freezer 140 L |  | Yes |  |  | Always needed | 1 | 2
106426 | MEDIA KITCHEN | Instrument | Mixers, Shakers, Stirrers | Blender 2x Speed |  |  |  |  | Optional | 1 | 1
106427 | MEDIA KITCHEN | Instrument | Mixers, Shakers, Stirrers | Glass receptable for blender 1l |  |  |  |  | Optional | 1 | 1
106443 | BIOSAFETY | Instrument | Laboratory, Miscellaneous | UPS - 700 VA - 500 W | UPS with battery pack for 20 minutes at 700W | Yes |  | 106046 | Always needed | 1 | 3
106492 | BIOSAFETY | Instrument | Laboratory, Miscellaneous | Battery Pack for UPS 700VA | Battery pack for 120 minutes at 500W | Yes |  |  | Regular needed | 1 | 1
106451 | LIQUID CULTURE | Instrument | Laboratory, Miscellaneous | Densitometer | Densitometer with 1 adaptator for MGIT tubes (15-18 mm diameter - 120 mm length) + 1 set of Mc Farland standards (0.5, 1.0, 2.0, 3.0,4.0) |  |  |  | Regular needed | 1set | 1
106475 | LIQUID CULTURE | Instrument | Laboratory, Miscellaneous | Slide Warmer | Stretchning and Dying Table - Temperature range up to 100°C |  |  |  | Optional | 1 | 1
106494 | BIOSAFETY | Instrument | Meters | calibrated thermometer | Accuracy +/- 0.2°C - Range : -25°C to 100°C - Resolution 0.1°C |  |  |  | Optional | 1 | 1
106051 | MOLECULAR BIOLOGY | Tools | Tubes | Desktop Deep freeze storage box with lid for 0.2ml PCR tubes | Deep freeze storage box with lid for 0.2ml PCR tubes, autoclavable PP. Suitable for storing 0.2 ml single vessels, 8-strips, and 12-strips, and 96-well PCR plates. |  |  | 106385 | Always needed | 1 | 5
106052 | MOLECULAR BIOLOGY | Tools | Tubes | Desktop Deep freeze storage box with lid for 1.5 ml cryovials - Cardboard cryobox | Deep freeze storage box with lid for 1.5 ml cryovials,Cardboard, with 9x9 inlet |  |  | 106350 | Always needed | 1 | 10
106053 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 0.5 - 10 µl - LPA | 1-channel pipette, variable from 1 - 10 µl, - accuracy < 0.025 to 0.100, precision < 0.012 to 0.040 | Yes | Yes |  | Always needed | 1 | 4
106054 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 2,0 - 20 µl - LPA | 1-channel pipette, variable from 2,0 - 20 µ, - accuracy < 0.10 to 0.20, precision < 0.03 to 0.06 | Yes | Yes |  | Always needed | 1 | 4
106055 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 10 - 100 µl - LPA | 1-channel pipette, variable from 10 - 100 µl | Yes | Yes |  | Optional | 1 | 4
106056 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-chanel pipette, variable from 20 -200 ul - LPA | 1-chanel pipette, variable from 20 -200 ul - accuracy < 0.50 to 1.60, precision < 0.20 to 0.30 | Yes | Yes |  | Always needed | 1 | 4
106057 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 200 - 1000 µl - LPA | 1-channel pipette, variable from 100 - 1000 µl - accuracy < 3.00 to 8.00, precision < 0.60 to 1.50 | Yes | Yes |  | Always needed | 1 | 4
106058 | MOLECULAR BIOLOGY | Tools | Pipetting | Racks for 1 channel pipette - 6 positions | Racks suitable for pipettes offered, 6 positions, rotating stand, disinfectable with bleach and alcoohol at least |  |  |  | Always needed | 1 | 4
106068 | MOLECULAR BIOLOGY | Tools | Tubes | Stand (96) for 0.2 ml pcr tubes / PCR Rack | Stand (96) for 0.2 ml, plastic resistant to bleach and alcoohol at least |  | Yes | PCR Tubes | Always needed | set of 10 | 3
106072 | LIQUID CULTURE | Tools | Pipetting | 1-chanel pipette, variable from 20 -200 ul - LC | 1-chanel pipette, variable from 20 -200 ul - accuracy < 0.50 to 1.60, precision < 0.20 to 0.30 | Yes | Yes |  | Regular needed | 1 | 3
106075 | LIQUID CULTURE | Tools | Pipetting | 1-channel pipette, variable from 200 - 1000 µl - LC | 1-channel pipette, variable from 100 - 1000 µl - accuracy < 3.00 to 8.00, precision < 0.60 to 1.50 | Yes | Yes |  | Regular needed | 1 | 4
106082 | BIOSAFETY | Tools | Disinfectants and Cleaning | Swivel stool/laboratory chair | Swivel stool (laboratory chair) for BSC, vertical adjustable 480 - 540 mm, seat and back made of PU foam, disinfectable, 5 wheels, no armrests |  |  | 106076 | Regular needed | 1 | 2
106083 | BIOSAFETY | Tools | Disinfectants and Cleaning | Mobile container | Mobile container (width 600 mm, depth approx 550 mm, height approx. 750 mm) on castors, set of five roller castors, two lockable, carcasse closed on five sides, disinfectable, worktop about 30 mm thick, preferably PP, high chemical resistance, temperature resistance to 90 °C, light and colourfast, edges of the worktop lipped with impact resistant PP, four drawers, smooth action, blocked for total extraction - Please provide drawings |  |  |  | Regular needed | 1 | 2
106084 | BIOSAFETY | Tools | Pipetting | Safety-Pipetting Ball | Safety-Pipetting Ball, standard |  |  |  | Optional | 1 | 10
106085 | LIQUID CULTURE | Tools | Pipetting | Electricity net independent automatic pipetting aid | Electricity net independent automatic pipetting aid with accu, rechargeable device, for usual pipettes and pasteur pipettes, volumina between 0.1 ml and 100 ml.Pipetting velocities in/out at least 3 levels. Battery capacity for 8 h work, low battery indication, 200 g weight |  |  |  | Optional | 1 | 2
106086 | BIOSAFETY | Tools | Pipetting | Set of mechanical pipetting aids | Set of mechanical pipetting aids: up to 0.2 ml; up to 2 ml and up to 10 ml, thumb wheel, operating for standard pipettes with quick release |  |  |  | Optional | 1 | 2
106087 | BIOSAFETY | Tools | Sample Collection and Media Transport | Transport box for cultures with BioPack insert | Transport box for cultures with BioPack insert, IATA approved, approx 1.5 litre volume, height inside ≥ 145mm, used for the exchange of strains between |  |  |  | Optional | 1 | 6
106088 | BIOSAFETY | Tools | Meters | Electronic maxima-minima-thermometer | Electronic maxima-minima-thermometer, two channels, one for room temperature and the second for temperature inside the equipment (incubator/refrigerator), switchable Celsius/Fahrenheit, thin cable with sensor, battery, measuring range for room temperature -10°C/+50°C (-14°F/+!122°F) and for equipment from -50°C/+60°C (-58°F/+140°F), reset function, number depending on the sum of incubators refrigerators, freezers and deep freezers |  | Yes |  | Always needed | 1 | 10
106089 | BIOSAFETY | Tools | Disinfectants and Cleaning | Plastic-foil welder 30 cm | Plastic-foil welder 30 cm - Simple robust plastic welder for the welding of PE, PP, Soft-PVC and Polyamide-film/bags. |  |  |  | Regular needed | 1 | 1
106090 | BIOSAFETY | Tools | Disinfectants and Cleaning | Dispenser for disinfectant/soap | Dispenser 'Eurospender 2000' for 1000ml bottles for wall mounting (2 per sink, one for liquid soap and one for disinfectant) |  |  |  | Regular needed | 1 | 8
106091 | BIOSAFETY | Tools | Disinfectants and Cleaning | Paper towel box | Wall mounted, refillable box for a capacity of at least 150 paper towels |  |  |  | Always needed | 1 | 4
106092 | BIOSAFETY | Tools | Disinfectants and Cleaning | Goggles | Goggles, according to EN 166 and EN 170, frame adjustable (at least three positions, non-skid, integrated side and top protection, total weight <35 g , UV Protection |  |  |  | Regular needed | 1 | 5
106094 | BIOSAFETY | Tools | Disinfectants and Cleaning | First Aid Kit for scientific laboratories | First Aid Kit for scientific laboratories, content according DIN 13169 or equivalent, with additional eye rinsing bottle, wall mounted, lockable cupboard, door with sealing gasket |  |  |  | Regular needed | 1 | 1
106095 | BIOSAFETY | Tools | Disinfectants and Cleaning | Stainless steel discard funnel | Stainless steel discard funnel, short wide stem with air drain d = 150 mm on top |  |  |  | Regular needed | 1 | 6
106349 | BIOSAFETY | Tools | Laboratory, Miscellaneous | Tri-pod stand for waste bags of ca. 2 litres | Stands, 3 foot, for small plastic bags (size ca. 2 litre) |  | Yes |  | Always needed | 1 | 5
106445 | BIOSAFETY | Tools | Laboratory, Miscellaneous | Stand for 30L plastic bags | Stable, should stay in a fix position especially when bag is filled |  | Yes | 106445 / 106348 | Always needed | 1 | 3
106096 | SOLID CULTURE | Tools | Plastic & Glassware | Glass beads, massive glass, 3 mm diameter | Glass beads, massive glass, 3 mm diameter |  |  |  | Optional | 1 | 1
106097 | SOLID CULTURE | Tools | Plastic & Glassware | Glass beads, massive glass, 5 mm diameter | Glass beads, massive glass, 5 mn diameter |  |  |  | Optional | 1 | 2
106098 | SOLID CULTURE | Tools | Tubes | Caped glass tubes for preparing suspension | Thick walled Pyrex tubes with ISO winding and screw cap with tight gasket, both autoclavable, 16x100 mm; package of 100; and for dilutuion of suspension, thick walled borosilicate 3.3 test tubes 12 mm diameter and 100 mm length may be used |  |  |  | Optional | 100 | 5
106109 | CLEANING KITCHEN | Tools | Pipetting | Pipette box aluminium/stainless steel with lid for 20 - 25 cm long pipettes | Pipette box aluminium/stainless steel with lid, square, silicon insert, desinfectable and autoclavable/sterilizable up to +205 °C , suitable for 20 - 25 cm long pipettes, for working in a BSC |  |  |  | Optional | 1 | 10
106110 | CLEANING KITCHEN | Tools | Pipetting | Pipette box aluminium/stainless steel with lid for 35 cm long pipettes | Pipette box aluminium/stainless steel with lid, square, silicon insert, desinfectable and autoclavable/sterilizable up to +205 °C , suitable for 35 cm long pipettes |  |  |  | Optional | 1 | 10
106111 | CLEANING KITCHEN | Tools | Pipetting | Inlay for pipette box | Inlay for pipette box (e.g. Teflon ) |  |  |  | Optional | 1 | 20
106112 | CLEANING KITCHEN | Tools | Pipetting | Pipette washer | Pipette washer, automatic siphon system, consisting of one cylindric washing container with siphon outlet, connecting tube, two pipette baskets as inserts, total height approx.750 mm, diameter approx. 170 mm |  |  |  | Optional | 1 | 1
106114 | LIQUID CULTURE | Tools | Plastic & Glassware | Petri-dishes glass for reuse | 94 x 16 mm |  |  |  | Always needed | 72 | 50
106143 | MEDIA KITCHEN | Tools | Plastic & Glassware | Balloon HDPE 30 L | Balloon HDPE (Spare) winding with discharge cock, 30 L (additional storage tank) |  |  |  | Optional | 1 | 1
106144 | MEDIA KITCHEN | Tools | Plastic & Glassware | Balloon HDPE 10 L | Balloon HDPE (Spare) winding with discharge cock, 10 L (for distributing distilled water to other parts of the laboratory) |  |  |  | Optional | 1 | 4
106145 | MEDIA KITCHEN | Tools | Plastic & Glassware | Stop cock PP | Stop cock PP, HDPE, winding for HDPE balloon |  |  |  | Optional | 1 | 10
106146 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring pipettes 1 ml - Glass | Measuring pipettes 1 ml, graduated in brown 0,01 gdt., conformity certified, pack of 10 pieces |  |  |  | Optional | 10 | 3
106147 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring pipettes 2 ml - Glass | Measuring pipettes 2 ml, graduated in brown 0,02 gdt., conformity certified, pack of 10 pieces |  |  |  | Optional | 10 | 3
106148 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring pipettes 5 ml - Glass | Measuring pipettes 5 ml, graduated in brown 0,05 gdt., conformity certified, pack of 10 pieces |  |  |  | Optional | 10 | 3
106149 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring pipettes 10 ml - Glass | Measuring pipettes 10 ml, graduated in brown 0,1 gdt., conformity certified, pack of 10 pieces |  |  |  | Optional | 10 | 3
106150 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Desiccator | Desiccator made of Duran, corpus: total height approx. 150 mm, inner diameter approx. 175 mm, corresponding lid with tread GL 32, cock with screw cap for lid, insert according to inner diameter of porcelain |  |  |  | Optional | 1 | 2
106151 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Membrane vacuum pump for desiccator | Membrane vacuum pump for desiccator, suction > 15 litre per hour, pressure (endpoint) to be reached < 20 mbar |  |  | 106151 | Optional | 1 | 1
106152 | MEDIA KITCHEN | Tools | Plastic & Glassware | Woulff's flask | Woulff's flask consisting of vacuum and pressure resistant Duran glass flask,1000 ml with thread GL 45 with screw cap, top part/insert according to Drechsel, adjustable in height, two olives, 8 mm outer diameter with screw capes GL 14, silicone ring |  |  | 106151 | Optional | 1 | 1
106153 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Vacuum tube, 3 m | Vacuum tube, 3 m |  |  | 106151 | Optional | 1 | 1
106154 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Silica gel for desiccator, 0.5 kg | Silica gel for desiccator, 0.5 kg |  |  | 106151 | Optional | 1 kg | 1
106155 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Bunsen burner | Teclu - all gas burner with fine tuning |  |  |  | Optional | 1 | 2
106156 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Tripod | Tripod stand with fire mat |  |  |  | Optional | 1 | 1
106157 | MEDIA KITCHEN | Tools | Plastic & Glassware | Swan neck bottles - Set | Swan neck bottles (PE flask, narrow neck with screw cap and water dispenser) set of one each 250, 500, 1000ml |  |  |  | Optional | 1 | 5
106158 | MEDIA KITCHEN | Tools | Plastic & Glassware | Bowl | Bowl, plastic aprox. 500 x 350 mm |  |  |  | Optional | 1 | 5
106159 | MEDIA KITCHEN | Tools | Plastic & Glassware | Erlenmeyer flasks - 2000ml | Erlenmeyer flasks, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, wide neck, graduated, autoclavable, 2000 ml, |  |  |  | Optional | 10 | 2
106160 | MEDIA KITCHEN | Tools | Plastic & Glassware | Erlenmeyer flasks - 1000ml | Erlenmeyer flasks, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, wide neck, graduated, autoclavable, 1000 ml, pack. unit 10 |  |  |  | Optional | 10 | 1
106161 | MEDIA KITCHEN | Tools | Plastic & Glassware | Erlenmeyer flasks - 500ml | Erlenmeyer flasks, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, wide neck, autoclavable, 500 ml, pack. unit 10 |  |  |  | Optional | 10 | 1
106162 | MEDIA KITCHEN | Tools | Plastic & Glassware | Erlenmeyer flasks - 250ml | Erlenmeyer flasks, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, wide neck, graduated, autoclavable, 250 ml, pack. unit 10 |  |  |  | Optional | 10 | 1
106163 | MEDIA KITCHEN | Tools | Plastic & Glassware | Erlenmeyer flasks - 100ml | Erlenmeyer flasks, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, wide neck, graduated, 100 ml, pack. unit 10 |  |  |  | Optional | 10 | 2
106164 | MEDIA KITCHEN | Tools | Plastic & Glassware | Erlenmeyer flasks - 50ml | Erlenmeyer flasks, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, wide neck, graduated, autoclavable, 50 ml, pack. unit 10 |  |  |  | Optional | 10 | 2
106165 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 1000ml - 0.5ml graduated | Measuring cylinder, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, class A, autoclavable, blue graduated with ring at main points, according to DIN 12680-2, with spout, adjusted to IN, 1000 ml, 0.5 ml graduated |  |  |  | Optional | 1 | 10
106166 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 500ml - 2.5ml graduated | Measuring cylinder, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, class A, autoclavable, blue graduated with ring at main points, according to DIN 12680-2, with spout, adjusted to IN, 500 ml, 2.5 ml graduated |  |  |  | Optional | 1 | 10
106167 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 250ml - 1ml graduated | Measuring cylinder, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, class A, autoclavable, blue graduated with ring at main points, according to DIN 12680-2, with spout, adjusted to IN, 250 ml, 1.0 ml graduated |  |  |  | Optional | 1 | 10
106168 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 100ml - 0.5ml graduated | Measuring cylinder, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, class A, autoclavable, blue graduated with ring at main points, according to DIN 12680-2, with spout, adjusted to IN, 100 ml, 0.5 ml graduated |  |  |  | Optional | 1 | 20
106169 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 50ml - 0.5ml graduated | Measuring cylinder, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, class A, autoclavable, blue graduated with ring at main points, according to DIN 12680-2, with spout, adjusted to IN, 50 ml, 0.5 ml graduated |  |  |  | Optional | 1 | 10
106170 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 25ml - 0.25ml graduated | Measuring cylinder, Duran glass/borosilicate 3.3 according to DIN/ISO 3585, class A, autoclavable, blue graduated with ring at main points, according to DIN 12680-2, with spout, adjusted to IN, 25 ml, 0.25 ml graduated |  |  |  | Optional | 1 | 10
106171 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 1000ml - 10ml graduated | Measuring cylinder, class A; TPX, autoclavable, DIN 12681/ISO 6706, 1000 ml, graduated 10.0 ml |  |  |  | Optional | 1 | 5
106172 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 500ml - 5ml graduated | Measuring cylinder, class A; TPX, autoclavable, DIN 12681/ISO 6706, 500 ml, graduated 5.0 ml |  |  |  | Optional | 1 | 5
106173 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 250ml - 2ml graduated | Measuring cylinder, class A; TPX, autoclavable, DIN 12681/ISO 6706, 250 ml, graduated 2.0 ml |  |  |  | Optional | 1 | 5
106174 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 100ml - 1ml graduated | Measuring cylinder, class A; TPX, autoclavable, DIN 12681/ISO 6706, 100 ml, graduated 1.0 ml |  |  |  | Optional | 1 | 10
106175 | MEDIA KITCHEN | Tools | Plastic & Glassware | Measuring cylinders - 50ml - 1ml graduated | Measuring cylinder, class A; TPX, autoclavable, DIN 12681/ISO 6706, 50 ml, graduated 1.0 ml |  |  |  | Optional | 1 | 5
106176 | MEDIA KITCHEN | Tools | Plastic & Glassware | Thick-walled glass flask - 2500ml | Thick-walled glass flask for preparation of culture medium; with flanged rim; 2500 ml |  |  |  | Optional | 1 | 6
106177 | MEDIA KITCHEN | Tools | Plastic & Glassware | Thick-walled glass flask - 2000ml | Thick walled glass bottles, Duran, with thread (DIN GL 45 winding), graduated, transparent, autoclavable, according to ISO 4796, with ring and cap of PP, 2000 ml |  |  |  | Regular needed | 1 | 3
106178 | MEDIA KITCHEN | Tools | Plastic & Glassware | Thick-walled glass flask - 1000ml | Thick walled glass bottles, Duran, with thread (DIN GL 45 winding), graduated, transparent, autoclavable, according to ISO 4796, with ring and cap of PP, 1000 ml |  |  |  | Regular needed | 1 | 10
106179 | MEDIA KITCHEN | Tools | Plastic & Glassware | Thick-walled glass flask - 500ml | Thick walled glass bottles with thread (DIN GL 45 winding), graduated, transparent, autoclavable, according to ISO 4796, with ring and cap of PP, 500 ml; |  |  |  | Optional | 1 | 5
106182 | MEDIA KITCHEN | Tools | Plastic & Glassware | Screw caps with weld in PFTE membrane, 0.2 µm pore size, autoclavable to 140 °C | screw caps with weld in PFTE membrane, 0.2 µm pore size, autoclavable to 140 °C |  |  |  | Optional | 1 | 10
106183 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beaker, low form with graduation and spout, 2000 ml | Laboratory beaker, low form with graduation and spout, 2000 ml |  |  |  | Optional | 1 | 3
106184 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beaker, low form with graduation and spout, 1000 ml | Laboratory beaker, low form with graduation and spout, 1000 ml |  |  |  | Optional | 1 | 3
106185 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beaker, low form with graduation and spout, 600 ml | Laboratory beaker, low form with graduation and spout, 600 ml |  |  |  | Optional | 1 | 10
106186 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beaker, low form with graduation and spout, 250 ml | Laboratory beaker, low form with graduation and spout, 250 ml |  |  |  | Optional | 1 | 10
106187 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beaker, low form with graduation and spout, 100 ml | Laboratory beaker, low form with graduation and spout, 100 ml |  |  |  | Optional | 1 | 10
106188 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beakers; 100 ml (ETFE) | Laboratory beakers; 100 ml; ETFE; heat-stable up to150 °C, with spout and graduation |  |  |  | Optional | 1 | 5
106189 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beakers; 250 ml (ETFE) | Laboratory beakers; 250 ml; ETFE; heat-stable up to150 °C; low form; with spout and graduation |  |  |  | Optional | 1 | 5
106190 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beakers; 500 ml (ETFE) | Laboratory beakers; 500 ml; ETFE; heat-stable up to150 °C; low form; with spout and graduation |  |  |  | Optional | 1 | 5
106191 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beakers; 1000 ml (ETFE) | Laboratory beakers; 1000 ml; ETFE ;heat-stable up to150 °C; low form; with spout and graduation |  |  |  | Optional | 1 | 3
106192 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory beakers; 2000 ml (PP) | Laboratory beakers; 2000 ml; ETFE ;heat-stable up to150 °C; low form; with spout and graduation |  |  |  | Optional | 1 | 3
106193 | MEDIA KITCHEN | Tools | Plastic & Glassware | Volumetric flasks 1000 ml | Volumetric flasks 1000 ml, Duran with NS cone and PE stopper, class A, autoclavable, adjusted to IN, error margin ± 0.5 ml, per piece |  |  |  | Optional | 1 | 3
106194 | MEDIA KITCHEN | Tools | Plastic & Glassware | Volumetric flasks 500 ml | Volumetric flasks 500 ml, Duran with NS cone and PE stopper, class A, autoclavable, adjusted to IN, error margin ± 0.25 ml, per piece |  |  |  | Optional | 1 | 10
106195 | MEDIA KITCHEN | Tools | Plastic & Glassware | Volumetric flasks 250 ml | Volumetric flasks 250 ml, Duran with NS cone and PE stopper, class A, autoclavable, adjusted to IN, error margin ± 0.15 ml, per piece |  |  |  | Optional | 1 | 10
106196 | MEDIA KITCHEN | Tools | Plastic & Glassware | Volumetric flasks 200 ml | Volumetric flasks 200 ml, Duran with NS cone and PE stopper, class A, autoclavable, adjusted to IN, error margin ± 0.15 ml, per piece |  |  |  | Optional | 1 | 15
106197 | MEDIA KITCHEN | Tools | Plastic & Glassware | Volumetric flasks 100 ml | Volumetric flasks 100 ml, Duran with NS cone and PE stopper, class A, autoclavable, adjusted to IN, error margin ± 0.1 ml, per piece |  |  |  | Optional | 1 | 3
106198 | MEDIA KITCHEN | Tools | Plastic & Glassware | Volumetric flasks 50 ml | Volumetric flasks 50 ml, Duran glass/borosilicate, with NS cone and stopper, class A, autoclavable, adjusted to IN, error margin ± 0.08 ml, per piece |  |  |  | Optional | 1 | 15
106199 | MEDIA KITCHEN | Tools | Plastic & Glassware | Volumetric flasks 25 ml | Volumetric flasks 25 ml, Duran glass/borosilicate, with NS cone and stopper, class A, autoclavable, adjusted to IN, error margin ± 0.06 ml, per piece |  |  |  | Optional | 1 | 5
106200 | MEDIA KITCHEN | Tools | Plastic & Glassware | Volumetric flasks 20 ml | Volumetric flasks 20 ml, Duran glass/borosilicate, with NS cone and stopper class A, autoclavable, adjusted to IN, error margin ± 0.06 ml, per piece |  |  |  | Optional | 1 | 25
106201 | MEDIA KITCHEN | Tools | Plastic & Glassware | Laboratory glassware drying rack | Laboratory glassware drying rack, wall mounted; polystyrene; min. 24 pegs for cylinders and conical flasks |  |  |  | Optional | 1 | 1
106202 | MEDIA KITCHEN | Tools | Plastic & Glassware | Funnel, soda lime glass, plain, short stem, d = 75 mm | Funnel, soda lime glass, plain, short stem, d = 75 mm (or plastic, if not used for dye solutions) |  |  |  | Optional | 1 | 5
106203 | MEDIA KITCHEN | Tools | Plastic & Glassware | Funnel, soda lime glass, plain, short stem, d = 125 mm | Funnel, soda lime glass, plain, short stem, d = 125 mm (or plastic, if not used for dye solutions) |  |  |  | Optional | 1 | 5
106204 | MEDIA KITCHEN | Tools | Plastic & Glassware | Powder funnel , diameter 100mm | Powder funnel , diameter 100mm |  |  |  | Optional | 1 | 5
106206 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Strainer, stainless steel, 1 grip, 2 eyes, diameter 19 cm | Strainer, stainless steel, 1 grip, 2 eyes, diameter 20 cm |  |  |  | Optional | 1 | 1
106207 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Set of spatula | Set of spatula: stainless steel 18/8 electrolytically polished, length approx. 150 mm, blade approx. 40 mm in length and width ranging from 3 mm to 9 mm, 6 pieces in most common sizes of double and micro spatulas, 3 of them with spoon on one side |  | Yes |  | Optional | 1 | 6
106208 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Forceps; pointed; straight; 18/8 stainless steel, polished | Forceps; pointed; straight; 18/8 stainless steel, polished |  |  |  | Optional | 1 | 5
106216 | MOLECULAR BIOLOGY | Tools | Laboratory, Miscellaneous | Floating rack | Floating rack with 20 places for reaction containers, diameter 10 mm, 2ml, for ultrasonic bath |  |  |  | Optional | 5 | 2
106217 | MOLECULAR BIOLOGY | Tools | Tubes | Racks with lid (plastic preferred), suitable for cryovial | Racks with lid (plastic preferred), suitable for cryovial (2ml, round outer winding), Operating range -196 °C to +121 °C, autoclavable at 121 °C (20 min). |  | Yes | 106344 /106381 / 106382 | Always needed | 1 | 50
106230 | LIQUID CULTURE | Tools | Centrifuges | Test tube rack 18/8 stainless steel 16 mn | diameter 16 mm, 3x12 positions, sturdy |  | Yes | 106338 | Regular needed | 1 | 5
106231 | PROCESSING | Tools | Centrifuges | Test tube rack 18/8 stainless steel 35 mn | diameter 35 mm, 2x4 positions |  | Yes |  | Optional | 1 | 5
106232 | SOLID CULTURE | Tools | Laboratory, Miscellaneous | Loop Holder | Loop holder, needle holder, upper part aluminium or stainless steel, with plastic handle, length 240mm |  |  |  | Optional | 1 | 12
106233 | SOLID CULTURE | Tools | Laboratory, Miscellaneous | Racks for loop holder | Racks for loop holder, smooth, disinfectable, proper stand, 3 to 4 positions |  |  | 106232 | Optional | 1 | 4
106234 | SOLID CULTURE | Tools | Laboratory, Miscellaneous | Inoculation loops 4mn | Inoculation loops, Platinum/Iridium wire, length 5 cm, wire diameter 0.6 - 0.7 mm, loop diameter 4mm, closed (or closed loops, 3 - 4 mm diameter, made of wire with 0.8 - 1 mm diameter, nickel-chromium alloy 89 20, heat resistant to 1200 °C) |  |  | 106232 | Optional | 1 | 8
106235 | SOLID CULTURE | Tools | Laboratory, Miscellaneous | Inoculation loops 3mn | Inoculation loops, Platinum/Iridium wire, length 5 cm, wire diameter 0.6 - 0.7 mm, loop diameter 3mm, closed, 10 µl for DST |  |  | 106232 | Optional | 1 | 8
106237 | PROCESSING | Tools | Tubes | Universal test tube rack 16mm | Universal test tube rack, PP, temperature resistant form -15 °C to +130 °C, autoclavable, for 16 mm diameter tubes, dimensions (W x D x H) 250 mm x 105 mm x 70 mm, 60 positions. (not swimming) |  | Yes |  | Regular needed | 1 | 12
106238 | PROCESSING | Tools | Tubes | Universal test tube rack 18mm | Test tube rack PP, diameter 18mm, 4x12, sturdy |  | Yes |  | Optional | 1 | 15
106239 | PROCESSING | Tools | Tubes | Universal test tube rack 28mm - | Test tube rack, 21 holes, with alphanumeric markings, Ø 30mm |  |  |  | Optional | 1 | 15
106240 | PROCESSING | Tools | Meters | Alarm Clock | Alarm clock / digital count down timer, 4 digits, programmable for max. 99 minutes, alarm at O, memory function |  |  |  | Always needed | 1 | 3
106241 | LIQUID CULTURE | Tools | Pipetting | Measuring pipette, short to work inside BSC, sterile, 1 ml | Measuring pipette, class A, short for work in BSC, approx. 25 cm long, 1 ml, 0.01 subdivision, with Coton |  |  |  | Optional | 1 | 200
106242 | PROCESSING | Tools | Sample Collection and Media Transport | Transport boxes for specimen | Transport boxes for specimen from collection points, styrofoam covered with PP, including cooling elements |  | Yes |  | Optional | 1 | 10
106243 | PROCESSING | Tools | Sample Collection and Media Transport | Cool boxes for sputum transport | Cool boxes for sputum transport by car, inner volume approx. 27 litre. Inner dimensions (W x D x H) ≥ 335 mm x ≥ 230 mm x ≥ 380 mm Inside and outside durable, disinfectable surface (e.g. LDPE), isolated with PU foam, Peltier elements, 12 V plug for cigarette lighter in a car cable length ≥ 2.5 m, possibility to change the polarity for cooling or heating alternatively. Cooling capacity 20 °C below ambient and heating 40°C above ambient. This transportation box is used in areas with extreme temperature deviations over the year and long duration of transport. |  | Yes |  | Optional | 1 | 3
106244 | PROCESSING | Tools | Sample Collection and Media Transport | sturdy containers | sturdy containers with thight closing, leak prove covers are needed as second container for packing. 1 L or 2L containers (depending on the samle size to be transported from the remote laboratory), minimum height 140 mm, made from PP with gasket tightened screw cap or from stainless steel with tightly fixed cover (e.g. for coffee) |  |  |  | Optional | 1 | 10
106245 | PROCESSING | Tools | Sample Collection and Media Transport | material for packing specimen | Sufficient soft, sorbent material for packing specimen container into containers transport boxes of position 106244, in kg |  |  | 106244 | Optional | 1 | 10
106246 | MOLECULAR BIOLOGY | Tools | Tubes | Racks, desinfectable / autoclavable for 1.5/2 ml tubes | Racks, desinfectable / autoclavable for tubes 1.5 - 2 ml, 10 mm diameter, at least 20 positions |  | Yes | 106381 / 106382 | Always needed | 1 | 10
106249 | SOLID CULTURE | Tools | Laboratory, Miscellaneous | Hand lens | Hand lens |  |  |  | Optional | 1 | 1
106250 | WASTE MANAGEMENT | Tools | Laboratory, Miscellaneous | Stainless steel bucket including lid | Stainless steel bucket including lid, capacity approx. 10 L, height 250 mm, upper diameter 280 mm, lower diameter 185 mm |  |  |  | Regular needed | 1 | 5
106251 | WASTE MANAGEMENT | Tools | Laboratory, Miscellaneous | Discard bin 2 Liters | Discard bin stainless steel or thick walled glass with screw cap |  |  |  | Optional | 1 | 6
106252 | WASTE MANAGEMENT | Tools | Laboratory, Miscellaneous | Funnel, stainless steel 18/8 | Funnel, stainless steel 18/8, with groove for air outlet in the ramp, diameter of ramp approx. 15 mm, top diameter approx. 120 mm, total height approx. 165 mm |  |  |  | Optional | 1 | 6
106253 | WASTE MANAGEMENT | Tools | Laboratory, Miscellaneous | Safety box for liquid | Safety box for liquid, PP, approx 3 L, |  |  |  | Optional | 30 | 5
106255 | WASTE MANAGEMENT | Tools | Laboratory, Miscellaneous | Waste containers for solid | Waste containers |  |  |  | Optional | 6 | 3
106256 | WASTE MANAGEMENT | Tools | Laboratory, Miscellaneous | Waste containers for liquid | Waste containers |  |  |  | Optional | 1 | 6
106270 | MOLECULAR BIOLOGY | Tools | Pipetting | Racks for 1 channel pipette - 4 positions | Racks suitable for pipettes offered, 4 positions, disinfectable with bleach and alcoohol at least |  |  |  | Regular needed | 1 | 4
106409 | WASTE MANAGEMENT | Tools | Laboratory, Miscellaneous | Waste containers for sharps | Waste containers |  |  |  | Optional | 10 | 1
106410 | MOLECULAR BIOLOGY | Tools | Tubes | Racks for cryovials for One hand operation |  |  | Yes | 106383 | Always needed | 1 | 5
106413 | MOLECULAR BIOLOGY | Tools | Tubes | Rack for 50 ml PP-tubes | Rack for PP-tubes for centrifuge 50 ml, stable, |  | Yes | 106340 / 106341 | Always needed | 1 | 10
106432 | SOLID CULTURE | Tools | Tubes | Slanted racks | Made of PC, autoclavable, for simple and speedy preparation of angular agar with two different inclination angles. ( 5° and 20°) With 40 alphanumerically marked holes (4x10). Dimensions: L 285 x W 115 mm. (e.g. carl roth item N°E100.1) |  | Yes | 106338 | Regular needed | 1 | 5
106433 | MOLECULAR BIOLOGY | Tools | Tubes | Rack for 15 ml PP-tubes -universal test tube rack 17mm | Universal test tube rack, PP, temperature resistant form -15 °C to +130 °C, autoclavable, for 17 mm diameter tubes, dimensions (W x D x H) 250 mm x 105 mm x 70 mm, 60 positions. (not swimming) or comparable |  | Yes | 106342 | Regular needed | 1 | 12
106442 | BIOSAFETY | Tools | Meters | Deep-freeze thermometer | With bimetallic indicator. Range -50 to +50 °C. Graduation 1 °C. Round plastic housing, e.g. Carl roth H717.1 or similar |  | Yes |  | Always needed | 1 | 3
106446 | MOLECULAR BIOLOGY | Tools | Laboratory, Miscellaneous | Reagent Quick Chiller | e.g. Quick Chill™ unit or similar, polycarbonate; non-toxic insulating solution, with handle. For use in melecular biology applications such as ethanol precipitation of DNA/RNA samples. For repeatedly use for up to 45 minutes, holds 12 x 1.5 or 0.5 ml microcentrifuge tubes (16 inserts for 0.5 ml included). Stackable design. |  |  |  | Optional | 1 | 1
106453 | MOLECULAR BIOLOGY | Tools | Laboratory, Miscellaneous | Thermometer for Oven | Minimum temp : 40°C |  | Yes |  | Regular needed | 1 | 2
106454 | MOLECULAR BIOLOGY | Tools | Tubes | Rack 15 ml / 50 ml - 4-Way Microtube Rack | Rack four different usable sides, high chemical resistant, high temperature resistant and autoclavable |  | Yes |  | Regular needed | 1 | 5
106455 | MOLECULAR BIOLOGY | Tools | Tubes | Rack Heat resistant for Cryo vial 1.5ml - Modular Heating Blocks | Rack Heat resistant for Cryo vial 1.5ml - Modular Heating Blocks * close contact of tubes to block walls allows for maximum heat retention * Each block should have a thermometer well for measuring block temperature |  | Yes | 106350 | Regular needed | 1 | 2
106456 | MOLECULAR BIOLOGY | Tools | Pipetting | Racks for 1 channel pipette - 3 positions | Racks suitable for pipettes offered, 3 positions, disinfectable with bleach and alcoohol at least |  |  |  | Always needed | 1 | 2
106495 | MOLECULAR BIOLOGY | Tools | Tubes | Desktop Deep freeze storage box with lid for 1.5 ml cryovials - Plastic | Deep freeze storage box with lid for 1.5 ml cryovials, autoclavable, transparent lid with fields for marking content (e.g. Nalgene Cryoboxes - Part Code: 5026-1010) |  |  |  | Optional | 1 | 10
106496 | LIQUID CULTURE | Tools | Tubes | Slanted racks for BD LJ Tubes | Made of PC, autoclavable, for simple and speedy preparation of angular agar with two different inclination angles. ( 5° and 20°) With 40 alphanumerically marked holes (4x10). - Tube diameter between 20 and 22mm Compatible with 221413 - L-J Medium Slants (10/sp) |  | Yes |  | Optional | 1 | 5
106493 | BIOSAFETY | Instrument | Laboratory, Miscellaneous | Solar Panel system | Solar Panel system with battery pack. This system should provide power for a desktop computer and a GeneXpert® System IV (about 500w/h) |  |  |  | Optional | 1 | 1

## 2 - Consumables and Chemicals
Item Number | Method | Family | Sub Family | Item Name | GDF SPECIFICATION | Detailled Specification (Y/N) | Request for sample (Y/N) | Link to | Importance | Unit / package GDF | Quantity of Pack - Historical data - 2011 | Quantity of Pack - Historical data - 2012 | Quantity of Pack - Forecasted data - 2013 | Quantity of Pack - Forecasted data - 2014 | Supplier Specifications | Primary Packing Type | Number of units per primary packing type | Secondary Packing Type | Number of units per secondary packing type | Minimum Shelf Life in days | Gross weight in kg | Depth (in cm) | Width (in cm) | Height (in cm) | Volume in M3 | Special requirement for delivery | Special requirement for storage | Manufacturer Name | Model or Item Number (Manufacturer) | Country of Origin | Custom Tariff code | Please mention quality standard under which the item is produced (if any)
106259 | SOLID CULTURE | Chemicals | Loewenstein-Jensen-Medium | Ready mixed TB | Ready mixed TB - Medium according to Loewenstein-Jensen (LJ), with correct composition of ingredients from a certified company, 500 g plastic bottles e.g. Merck |  |  |  | Always needed | 500 g | 9 | 44 | 81 | 87
106261 | SOLID CULTURE | Chemicals | Loewenstein-Jensen-Medium | Glycerol ≥ 99% | Glycerol ≥ 99% Ph.Eur., C3H8O3, MW: 92.1, 1 L plastic bottle |  |  |  | Regular needed | 1 L | 3 | 41 | 38 | 39
106262 | SOLID CULTURE | Chemicals | Loewenstein-Jensen-Medium | Sodium pyruvate, 50g; bottle |  |  |  |  | Optional | 25g | 3 | 23 | same as 2012 | same as 2012
106263 | SOLID CULTURE | Chemicals | IUTM base | Potassium dihydrogen phosphate | Potassium dihydrogen phosphate, ≥ 99% p.a. anhydrous, KH2PO4, MW: 136,09, 1 kg plastic bottle |  |  |  | Always needed | 1 kg | 20 | 93 | 102 | 105
106264 | SOLID CULTURE | Chemicals | IUTM base | Magnesium sulfate-heptahydrate | Magnesium sulfate-heptahydrate, ≥ 99% p.a., MgSO4 • 7H2O, MW: 246,48. 500g plastic bottle |  |  |  | Regular needed | 500 g | 1 | 24 | 16 | 15
106265 | SOLID CULTURE | Chemicals | IUTM base | L-Asparagine-momohydrate | L-Asparagine-momohydrate, C4H8N2O3 • H2O, MW: 150.13, ≥ 99%, plastic bottle, 250g |  |  |  | Regular needed | 100g | 6 | 45 | 44 | 45
106266 | SOLID CULTURE | Chemicals | IUTM base | tri-Magnesium di-citrate nonahydrate | tri-Magnesium di-citrate nonahydrate, Mg3(C6H5O7)2 • 9 H2O, MW: 613,25, ≥ 95%, p.a., 500g plastic bottle |  |  |  | Regular needed | 100g | 3 | 27 | 16 | 15
106267 | SOLID CULTURE | Chemicals | IUTM base | Malachite green oxalate | Malachite green oxalate, p. A.; 100 g bottle |  |  |  | Regular needed | 100 g | 0 | 19 | 16 | 15
106268 | SOLID CULTURE | Chemicals | IUTM base | Glycerol | Glycerol |  |  |  | Optional | 1kg | 2 | 2 | same as 2012 | same as 2012
106269 | SOLID CULTURE | Chemicals | Ogawa-Medium | Sodium-L-Glutamate-monohydrate | Sodium-L-Glutamate-monohydrate, C5H8NNaO4 • H2O, MW: 187.13, purum, ≥ 98%, 1 kg plastic bottle |  |  |  | Optional | 1 kg | 0 | 6 | same as 2012 | same as 2012
106275 | SOLID CULTURE | Chemicals | Modified Ogawa-Medium | Magnesium sulphate-heptahydrate | Magnesium sulphate-heptahydrate |  |  |  | Optional | 1 | 0 | 1 | same as 2012 | same as 2012
106278 | SOLID CULTURE | Chemicals | Phosphate buffer according to Soerensen, 1/15 M, pH 6.8 or adjustable between pH 5 to 8.2 | di-Sodium hydrogen phosphate, anhydrous | di-Sodium hydrogen phosphate anhydrous (offered also with crystal water 2, 7, 12) Na2HPO4, MW: 141.96, p.a. for buffer, 1 kg plastic bottle |  |  |  | Always needed | 1 kg | 21 | 82 | 91 | 93
106279 | SOLID CULTURE | Chemicals | Phosphate buffer according to Soerensen, 1/15 M, pH 6.8 or adjustable between pH 5 to 8.2 | di-Sodium hydrogen phosphate dodecahydrate | Na2HPO4 12 H2O , MW: 358,14, p.a. for buffer, 1 kg plastic bottle |  |  |  | Optional | 1 kg | 3 | 11 | same as 2012 | same as 2012
106281 | PROCESSING | Chemicals | Petroff Method | Sodium hydroxide | Sodium hydroxide, NaOH, MW 40.00, purum ≥ 98%, pellets, 1kg plastic bottle, tight screw cap (hygroscopic and CO2 binding from the air) |  |  |  | Always needed | 1 | 20 | 63 | 40 | 40
106282 | PROCESSING | Chemicals | Sodium phosphate Method | tri-Sodium phosphate dodecahydrate | tri-Sodium phosphate dodecahydrate, Na3PO4 • 12H2O, MW: 380.12, (1 kg, 2.5 kg, 25 kg) 2.5 kg plastic bottle |  |  |  | Optional | 2.5 kg | 5 | 4 | same as 2012 | same as 2012
106283 | PROCESSING | Chemicals | NALC Method | N-acetyl-L-cysteine (NALC) | N-acetyl-L-cysteine (NALC), C5H9NO3S, MW: 163.19,puriss. ≥ 99%, 250g plastic bottle |  |  |  | Always needed | 100g | 55 | 114 | 80 | 83
106284 | PROCESSING | Chemicals | NALC Method | tri-Sodium citrate dihydrate | tri-Sodium citrate dihydrate, C6H5Na3O7 • 2H2O, MW: 294.10, plastic bottle 1 kg |  |  |  | Always needed | 1 kg | 23 | 58 | 33 | 33
106286 | PROCESSING | Chemicals | Cethypyridinum chloride for transit | Cetylpyridinium chloride monohydrate | Cetylpyridinium chloride monohydrate, C21H38ClN • H2O, MW: 358.01, melting point 77 ºC, purity ≥ 98% for biochemistry, (100 g /500g) 500g bottle |  |  |  | Optional | 500 g | 1 | 4 | same as 2012 | same as 2012
106287 | PROCESSING | Chemicals | Cethypyridinum chloride for transit | Sodium chloride | Sodium chloride, NaCl, MW: 58.44, Ph.Eur. ≥ 99%, 1kg plastic bottle |  |  |  | Optional | 1 kg | 3 | 4 | same as 2012 | same as 2012
106288 | PROCESSING | Chemicals | Pre-treatment of urine (enrichment before decontamination) | 5-Sulfosalicylic acid dihydrate | 5-Sulfosalicylic acid dihydrate, C6H7NO3S• 2H2O, MW: 254,22, p.a. ≥ 99%, 100g plastic bottle |  |  |  | Optional | 100 g | 1 | 1 | same as 2012 | same as 2012
106292 | IDENTIFICATION | Chemicals | Catalase Test | Tween® 80 | Tween® 80, none ionic detergent, (Polyoxyethlene-80-sorbitan monooleate), Ph.Eur., (250g/ 1 kg) 1kg plastic bottle |  |  |  | Optional | 1 kg | 3 | 4 | same as 2012 | same as 2012
106293 | IDENTIFICATION | Chemicals | Catalase Test | Hydrogen peroxide | Hydrogen peroxide, H2O2, MW: 34.01, 30%, p.a., ISO, stabilised, 250 ml plastic bottle |  |  |  | Optional | 250 ml | 0 | 0 | same as 2012 | same as 2012
106295 | IDENTIFICATION | Chemicals | Nitrate Reduction Test (Nitratase Test) | di-Sodium hydrogen phosphate | di-Sodium hydrogen phosphate |  |  |  | Optional | 1 | 0 | 3 | same as 2012 | same as 2012
106297 | IDENTIFICATION | Chemicals | Nitrate Reduction Test (Nitratase Test) | Sulphanilamide | Sulphanilamide, C6H8N2O2S, MW: 172.21, p.a. ≥ 98%, melting point > 164 °C, 100g plastic bottle |  |  |  | Optional | 100 g | 0 | 0 | same as 2012 | same as 2012
106298 | IDENTIFICATION | Chemicals | Nitrate Reduction Test (Nitratase Test) | N-(1-Naphthyl)-ethylendiamine dihydrochloride in glass bottle | N-(1-Naphthyl)-ethylendiamine dihydrochloride, C12H14N2 • 2HCl, MW: 259.18, p.a.≥ 98%, 25g plastic bottle |  |  |  | Optional | 25 g | 0 | 1 | same as 2012 | same as 2012
106301 | IDENTIFICATION | Chemicals | identification | Cyanogen bromide |  |  |  |  | Optional | 1 | 0 | 0 | same as 2012 | same as 2012
106302 | IDENTIFICATION | Chemicals | identification | Analine / o-toluidine |  |  |  |  | Optional | 1 | 0 | 0 | same as 2012 | same as 2012
106303 | MEDIA KITCHEN | Chemicals | Storage of Mycobacteria in a deep freezer | Skim milk powder | Skim milk powder, 500g (e.g. Oxoid) |  |  |  | Optional | 500 g | 4 | 7 | same as 2012 | same as 2012
106307 | DST | Chemicals | Solvents | Dimethyl sulfoxide | Dimethyl sulfoxide, 99,8%, p. a., 500 ml |  |  |  | Optional | 10 x 25 ml | 2 | 11 | same as 2012 | same as 2012
106308 | DST | Chemicals | First line drugs (FLD) | Isoniazid | 5g Isoniazid (INH) (e.g. Sigma I13377-5G); |  |  |  | Optional | 5 g | 1 | 11 | same as 2012 | same as 2012
106309 | DST | Chemicals | First line drugs (FLD) | Rifampicin | 1g Rifampicin (RMP) (e.g. Sigma R3501-1G) |  |  |  | Optional | 1 g | 1 | 13 | same as 2012 | same as 2012
106310 | DST | Chemicals | First line drugs (FLD) | Ethambutol | 25 g Ethambutol ( e.g. Sigma E4630-25G) |  |  |  | Optional | 25 g | 1 | 11 | same as 2012 | same as 2012
106311 | DST | Chemicals | First line drugs (FLD) | Dihydro-streptomycin | 5g Dihydro-streptomycin (DHSM (e.g. Sigma D7253-5G) |  |  |  | Optional | 5g | 2 | 12 | same as 2012 | same as 2012
106312 | DST | Chemicals | First line drugs (FLD) | Nicotinic acid amide | Nicotinic acid amide (nicotinamid, nicocinamid), pure powder, (may replace PZA at neutral pH in LJ) |  |  |  | Optional | 1 | 1 | 2 | same as 2012 | same as 2012
106313 | DST | Chemicals | Second Line Drugs (SLD) | Ofloxacin | Ofloxacin/Ciprofloxacin, (OFL/CIP), (e.g. Aventis/Bayer or Sigma O8757-1G/17850-5G-F ) |  |  |  | Optional | 1 | 3 | 18 | same as 2012 | same as 2012
106314 | DST | Chemicals | Second Line Drugs (SLD) | Moxifloxacin | Moxifloxacin (MOX), 2g (e.g.Bayer) |  |  |  | Optional | 1g | 4 | 4 | same as 2012 | same as 2012
106315 | DST | Chemicals | Second Line Drugs (SLD) | Protionamid, | Protionamid, (PTH), 2g (e,g, Fatol) |  |  |  | Optional | 2 g | 0 | 1 | same as 2012 | same as 2012
106316 | DST | Chemicals | Second Line Drugs (SLD) | Ethionamide | Ethionamid (ETH), 5g, (e.g. Sigma E6005-5G) |  |  |  | Optional | 5 g | 2 | 7 | same as 2012 | same as 2012
106317 | DST | Chemicals | Second Line Drugs (SLD) | Kanamycin | Kanamycin (KAN), 1g, (e.g. Sigma K1876-1G) |  |  |  | Optional | 1 g | 3 | 24 | same as 2012 | same as 2012
106318 | DST | Chemicals | Second Line Drugs (SLD) | Amikacin | Amikacin (AMI), 5g (e.g. Sigma A3650-5G) |  |  |  | Optional | 5 g | 3 | 9 | same as 2012 | same as 2012
106319 | DST | Chemicals | Second Line Drugs (SLD) | Capreomycin | Capreomycin (CM), 1g (e.g. Sigma C4142-1G) |  |  |  | Optional | 1 g | 3 | 11 | same as 2012 | same as 2012
106320 | DST | Chemicals | Second Line Drugs (SLD) | Para-aminosalicylic acid | para-aminosalicylic acid (PAS), 5g (e.g. Sigma A79604-5G) |  |  |  | Optional | 1 | 1 | 13 | same as 2012 | same as 2012
106321 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Agarose | Agarose, universal use for separation of DNA fragments, range 100 bp to 20 kbp, melting point <90 °C, gel formation < 45 °C, 100 g |  |  |  | Optional | 100 g | 1 | 5 | same as 2012 | same as 2012
106322 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Boric acid | Boric acid, > 99.5%,1 kg for buffer |  |  |  | Optional | 1 kg | 0 | 3 | same as 2012 | same as 2012
106323 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Sodium acetate trihydrate | Sodium acetate trihydrate, analytical grade, 1 kg |  |  |  | Optional | 1 kg | 0 | 3 | same as 2012 | same as 2012
106324 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Ethidium bromide powder | Ethidium bromide, 1 g |  |  |  | Optional | 1 g | 0 | 2 | same as 2012 | same as 2012
106325 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | TRIS | TRIS, > 99%, buffer grade, 1 kg, |  |  |  | Optional | 1 kg | 1 | 3 | same as 2012 | same as 2012
106326 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | EDTA | EDTA, 99%, for biochemistry, 250 g |  |  |  | Optional | 250 g | 1 | 3 | same as 2012 | same as 2012
106327 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Gel load buffer | Gel load buffer, 6x, with Ficoll and bromphenol blue/xylencyan blue, 6 x 1.8 ml fractions (e.g. see Roth) |  |  |  | Optional | 6 | 1 | 3 | same as 2012 | same as 2012
106328 | WASTE MANAGEMENT | Chemicals | Fumigation (BSC before filter change and/or emergency case) | Formalin | Formalin, CH2O, MW: 30.03, technical grade, ~37% stabilised with 10% methanol, 1 L plastic bottle |  |  |  | Regular needed | 1 l | 1 | 25 | same as 2012 | same as 2012
106329 | WASTE MANAGEMENT | Chemicals | Disinfectants | Stable chlorine disinfectant | Stable cholrine disenfectant in tablets to prepare 1% solution free chlorine. Unit per pack to prepare at least 10 liters final solution. |  |  |  | Always needed | 100 | 3 | 9 | 100 | 100
106331 | WASTE MANAGEMENT | Chemicals | Disinfectants | Phenol | Phenolic products e.g. triclosan, chloroxylenol, orthophenylphenol, to be used at concentrations of 5%, 2.5 L bottles |  |  |  | Regular needed | 1kg | 5 | 69 | 417 | 417
106333 | WASTE MANAGEMENT | Chemicals | Disinfectants | Isopropanol | Isopropanol (2-propanol), 5 L bottles |  |  |  | Always needed | 5 L | 23 | 116 | 1134 | 1134
106334 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | DNA ladder | Control for separation of DNA fragments |  |  |  | Optional | 1 | 1 | 3 | same as 2012 | same as 2012
106336 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Ethidium bromide solution, | 500µg/mL water, gel electrophoresis and DNA isolation procedures suitable-5ML |  |  |  | Optional | 5 ml | 1 | 4 | same as 2012 | same as 2012
106468 | LIQUID CULTURE | Chemicals | McFarland Standard | Mc Farland Standards kits | at least stable for 1 year after delivery |  |  |  | Optional | 1 | 5 | 15 | same as 2012 | same as 2012
106113 | LIQUID CULTURE | Consumables | Plastic & Glassware | Petri-dishes plastic | 94 x 16 mm - Autoclavable - Sterile - pack in 10 or 20 pieces per bag |  |  |  | Regular needed | 480 | 17 | 45 | 39 | 39
106181 | LIQUID CULTURE | Consumables | Pipetting | Dispenser-tips Universal 10ml sterile | Dispenser-tips Universal 10ml sterile - DNA- /RNase- /ATP- and endotoxin free - suitable for pipettes offered |  | Yes | 106080 | Always needed | 100 | 43 | 147 | 424 | 420
106209 | LIQUID CULTURE | Consumables | Pipetting | Serological pipette - 5 ml | Serological pipette -5,0 ml, graduation 0.1 ml, packaging single packed |  |  |  | Optional | 200 | 50 | 61 | same as 2012 | same as 2012
106210 | LIQUID CULTURE | Consumables | Pipetting | Serological pipette - 10 ml | Serological pipette - 10 ml, graduation 0.1 ml, packaging single packed |  |  |  | Optional | 250 | 59 | 45 | same as 2012 | same as 2012
106236 | SOLID CULTURE | Consumables | Laboratory, Miscellaneous | Disposable loops 10 µl | Disposable loops 10µl , sterile, length approx. 220 mm, pack of 10 |  |  | 106232 | Always needed | 500 | 52 | 147 | 102 | 104
106254 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Molecular grade water | DNAase, RNAase free water - not chemically treated molecular grade water - 1.7mlx10 or 10ml |  | Yes |  | Always needed | 10x1.7ml | 261 | 320 | 1091 | 1148
106294 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Disinfectant for cleaning instrument | TB Certified concentrated disinfectant e.g. BOMIX |  |  |  | Always needed | 5 L | 16 | 70 | 122 | 122
106304 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Hand pump for Disinfectant for cleaning instrument | Hand pump for cleaning instrument 10ml dosage |  |  |  | Regular needed | 1 | 10 | 33 | same as 2012 | same as 2012
106338 | SOLID CULTURE | Consumables | Tubes | Culture tubes with screw cap | Culture tubes, diameter 16 mm, length 125 mm, wall strength 1mm, thread outside; inner diameter min.13 mm; borosilicate glass with ISO winding, autoclavable screw caps with gasket out of PP or Teflon. Some prefer 18 mm diameter |  | Yes |  | Regular needed | 100 | 430 | 645 | 284 | 294
106339 | SOLID CULTURE | Consumables | Tubes | Universal glass bottle | 28 ml universal glass bottle (Mc Cartney), heavy duty glass, straight neck without shoulder, aluminium screw cap with rubber lining or other sterilisable screw cap with sealing gasket |  |  |  | Optional | 1 | 0 | 6000 | same as 2012 | same as 2012
106340 | MOLECULAR BIOLOGY | Consumables | Tubes | PP-tubes for centrifuge non sterile 50 ml | PP-tubes for centrifuge, 50 ml with screw cap, tight, size 115x30 mm (Falcon or equivalent), 500 pieces/pack - Resistance minimum 3,200 g in a swing out bucket | Yes | Yes |  | Regular needed | 500 | 52 | 36 | 70 | 74
106341 | LIQUID CULTURE | Consumables | Tubes | PP-tubes for centrifuge, sterile, 50 ml | PP-tubes for centrifuge, 50 ml with screw cap, tight, size 115x30 mm (Falcon or equivalent), sterile, packed 25 in racks, 20 racks per pack. | Yes | Yes |  | Always needed | 500 | 196 | 640 | 634 | 655
106342 | LIQUID CULTURE | Consumables | Tubes | PP-tubes for centrifuge, sterile, 15 ml | PP-tubes for centrifuge, 15 ml with screw cap, tight, size 120x17 mm, 500 pieces/pack - Resistance minimum 3,200 g in a swing out bucket |  | Yes |  | Regular needed | 500 | 88 | 135 | same as 2012 | same as 2012
106343 | LIQUID CULTURE | Consumables | Pipetting | Single use plastic Pasteur-pipettes sterile pack by 10 | sterile, single use plastic Pasteur-pipettes volume in a range of 1 to 3 ml, packed by 10, graduated 0.5ml, lenght : minimum 155mm ideal 170mm. Drop size maximum 30ul |  | Yes |  | Regular needed | 500 | 59 | 121 | same as 2012 | same as 2012
106344 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, 2 ml | Cryo-vial, sterile with standard screw cap, 2 ml round with outer winding, PP, graduated, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). pack of 1000 |  | Yes |  | Regular needed | 1000 | 15 | 47 | same as 2012 | same as 2012
106345 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size S | powder-free disposable latex gloves, size S, >0,16 mm thickness, length >240 mm, 1000 per pack |  | Yes |  | Always needed | 1000 | 40 | 377 | 417 | 417
106346 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size M | powder-free disposable latex gloves, size M, >0,16 mm thickness,length >240 mm, 1000 per pack |  | Yes |  | Always needed | 1000 | 60 | 259 | 417 | 417
106347 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size L | powder-free disposable latex gloves, size L, >0,16 mm thickness, length >240 mm, 1000 per pack |  | Yes |  | Always needed | 1000 | 37 | 190 | 417 | 417
106348 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Transparent polypropylene bag | Transparent polypropylene bag for waste bins, at least 40 µm thickness, up to 134 °C steam sterilisable, approx 24 L, approx. 540 x 550 mm, in accordance with countiy's norms for environmental and hygiene assignments, 1000 pieces per pack |  | Yes |  | Always needed | 1000 | 12 | 45 | 60 | 60
106350 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, 1.5 ml | Cryo-vial, sterile with standard screw cap, 1.5 ml, PP with a medical-grade silicone O-ring seal, style conical, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). It can be centrifuged up to 17 000 x g - e.g. SIGMA V4131 |  | Yes |  | Always needed | 1000 | 38 | 109 | 274 | 286
106351 | LIQUID CULTURE | Consumables | Laboratory, Miscellaneous | Single use syringes, sterile | Sterile, single use syringes according to ISO 7886-1 with Luer connection, PP, graduated with resistant colour, safe piston lock, volume 20 ml, individually wrapped, 100 pieces per pack |  |  |  | Always needed | 50 | 17 | 15 | same as 2012 | same as 2012
106352 | LIQUID CULTURE | Consumables | Laboratory, Miscellaneous | Syringe filter for single use, sterile | Syringe filter for single use, entrance Luer connection, hydrophilic (e.g. PVDF filter in PE housing), pore size ≤ 0.22 µm, sterile and individually wrapped, at least 25 mm diameter, pressure up to 5 bar, 100 pieces per pack |  |  |  | Always needed | 50 | 6 | 22 | same as 2012 | same as 2012
106353 | MEDIA KITCHEN | Consumables | Laboratory, Miscellaneous | Filter paper round 185 mm | Filter paper round, weight approx. 75 g/m², retention > 4µm, diameter 185 mm, packs of 100 e.g. Whatman |  |  |  | Optional | 100 | 23 | 15 | same as 2012 | same as 2012
106354 | MEDIA KITCHEN | Consumables | Laboratory, Miscellaneous | Filter paper round 150 mm | Filter paper, round, weight approx. 75 g/m², retention >4 µm, diameter 150 mm, packs of 100 e.g. Whatman |  |  |  | Optional | 100 | 17 | 9 | same as 2012 | same as 2012
106355 | PROCESSING | Consumables | Laboratory, Miscellaneous | Marker pen | Marker pen, resistant against water and alcohol, also against autoclaving, writing on nearly all surfaces, even when cold and wet, fast drying, black, line 0.8 - 1 mm broad, black |  | Yes |  | Regular needed | 1 | 110 | 496 | 465 | 489
106356 | PROCESSING | Consumables | Laboratory, Miscellaneous | Adhesive labels | Adhesive labels on rolls or sheets (e.g. 50 x 19 mm, 24 per sheet and 200 sheets per box = 4800 label) |  | Yes |  | Regular needed | 1 | 11 | 89 | same as 2012 | same as 2012
106357 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Cryo-tags | Cryo-tags sized to fit for use on cryo-tubes and standard micro-tubes (approx.32 mm x 12 mm), made from smear resistant material, but should accept all marking devices, rolles of 1000 |  | Yes |  | Always needed | 1 | 19 | 113 | 301 | 312
106358 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Sterile indicator tape, hot air oven | Sterile indicator tape (roles), approx.19mm broad and 55m long, DIN EN 867-5, for hot air sterilize (green/brown) |  | Yes |  | Always needed | 1 | 22 | 149 | 488 | 488
106359 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Sterile indicator tape, autoclave | Sterile indicator tape (roles), approx.19mm broad and 55m long, DIN EN 867-5, for autoclave (beige/dark brown) |  | Yes |  | Always needed | 1 | 29 | 200 | 488 | 488
106360 | CLEANING KITCHEN | Consumables | Laboratory, Miscellaneous | Aluminium foil | Aluminium foil, 30 µm thickness, 600 mm width, 100 m long |  | Yes |  | Regular needed | 1 | 6 | 65 | 122 | 122
106361 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Parafilm with dispenser | Parafilm, 100 mm width, 58 m length, for use between - 40 °C to +50 °C, with dispenser |  |  |  | Optional | 1 | 7 | 20 | same as 2012 | same as 2012
106362 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Plastic-foil | Plastic-foil rolls, 30 cm |  |  |  | Regular needed | 1 | 16 | 50 | 122 | 122
106363 | CLEANING KITCHEN | Consumables | Laboratory, Miscellaneous | Tube brush | Tube brush, total length min.280 mm, diameter 15-16 mm, brush length 80 mm - too fit witt culture tubes. Soft at the end to not brake the tube |  | Yes |  | Regular needed | 5 | 50 | 47 | 102 | 104
106364 | CLEANING KITCHEN | Consumables | Laboratory, Miscellaneous | Brushes | Brushes for glassware, total length approx.480 mm, diameter 60 mm, brush length 120 mm |  |  |  | Optional | 1 | 21 | 32 | same as 2012 | same as 2012
106366 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size L - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent |  | Yes |  | Always needed | 100 | 17 | 46 | 60 | 60
106367 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size M - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent |  | Yes |  | Always needed | 100 | 67 | 122 | 113 | 113
106368 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size S - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent |  | Yes |  | Always needed | 100 | 51 | 77 | 60 | 60
106370 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Disinfectant for floors | Disinfectant for floors, tuberculozid, locally made or commercially available as concentrate |  |  |  | Always needed | 1 | 39 | 101 | 122 | 122
106371 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Disinfectant for BSCs surface | Disinfectant for BSCs (Biosafety Cabinets) and surfaces, (alcohol containing disinfectant, certified for TB bacilli, disinfection within 1 minute) e.g. 2-propanol (iso-propanol) , good beddabling, evaporates completely and residuelessly; |  |  |  | Always needed | 5 L | 46 | 184 | 488 | 488
106372 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Spray head and 1L bottles | Spray head with 1 liter bottle |  |  |  | Always needed | 1 set | 110 | 146 | 122 | 122
106373 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Liquid soap | Liquid soap, pH neutral, bottle of 1 litre |  |  |  | Always needed | 1 L | 114 | 289 | 732 | 732
106374 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Disinfectant for hands | Disinfectant for hands 1 litre bottle, alcoholic disinfectant locally made or commercially available |  |  |  | Always needed | 1 L | 78 | 317 | 732 | 732
106375 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Single-use paper towels | Single-use paper towels, approx. 230 x 300 mm folded to take as single pieces off a wall mounted box. 150 towels/pack, 30 packs/ carton |  | Yes |  | Always needed | 4500 | 63 | 276 | 401 | 401
106376 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Cotton wool | Cotton wool, white adsorbent, as kg, purchase locally |  |  |  | Always needed | 1 | 7 | 56 | 244 | 244
106377 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Tissue pulp | Tissue pulp, absorbent sheets, approx. 550 x 350 mm |  |  |  | Always needed | 1 | 67 | 370 | 1307 | 1307
106378 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Washing Powder | detergent/washing Powder for laboratory dish washer will be needed (10 kg trading unit) |  |  |  | Optional | 10 kg | 0 | 2 | same as 2012 | same as 2012
106379 | CLEANING KITCHEN | Consumables | Disinfectants and Cleaning | Salt for dish washer | Salt for dish washer, 10 kg/pack |  |  |  | Optional | 12kg | 0 | 1 | same as 2012 | same as 2012
106380 | CLEANING KITCHEN | Consumables | Disinfectants and Cleaning | Rinse aid for dish washer | Rinse aid for dish washer, 1 litre |  |  |  | Optional | 1 L | 0 | 1 | same as 2012 | same as 2012
106383 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, for one hand operation | Cryo-vial, sterile with Tethered Screw Caps, 2 ml round with outer winding PP, graduated, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). pack of 1000 |  | Yes |  | Always needed | 1000 | 11 | 44 | 27 | 26
106384 | MOLECULAR BIOLOGY | Consumables | Tubes | PP-tubes for centrifuge, non sterile, 15 ml | PP-tubes for centrifuge, 15 ml with screw cap, tight, size 120x17 mm, 500 pieces/pack |  | Yes |  | Regular needed | 500 | 8 | 25 | 92 | 96
106385 | MOLECULAR BIOLOGY | Consumables | Tubes | PCR tubes | PCR tubes, the tubes should be extremely thin to ensures speedy and even heat transfer of every sample, pyrogen free, DNase- /RNase-free and autoclavable, 0,2 ml with attached caps, fits all standard thermocyclers with heated lids especially the one provided. |  | Yes |  | Always needed | 1000 | 53 | 97 | 149 | 155
106386 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 0.1 - 10 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 0.1 - 10 µl, |  | Yes | 106053 | Always needed | 480 | 94 | 288 | 567 | 594
106387 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 1.0 - 20 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 1.0 - 20 µl, |  | Yes | 106054 | Regular needed | 480 | 117 | 97 | same as 2012 | same as 2012
106388 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 10 - 100 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered, 10 - 100 µl, |  | Yes | 106055 | Regular needed | 960 | 48 | 34 | same as 2012 | same as 2012
106389 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 20 - 200 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 20 - 200 µl, |  | Yes | 106056 | Always needed | 960 | 81 | 252 | 475 | 496
106390 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS 100 - 1000 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 100 - 1000 µl, |  | Yes | 106057 | Always needed | 960 | 59 | 191 | 157 | 164
106394 | LIQUID CULTURE | Consumables | Pipetting | Single use plastic Pasteur-pipettes sterile individually packed | sterile, single use plastic Pasteur-pipettes volume in a range of 1 to 3 ml, individually packed, graduated 0.5ml, lenght : minimum 155mm ideal 170mm. Drop size maximum 30ul |  | Yes |  | Always needed | 500 | 101 | 509 | 513 | 531
106395 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 5 ml - bag of 10 | Tips for Multipette/Combi 100 pieces/pack - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered |  | Yes | 106115 | Regular needed | 100 | 3 | 7 | same as 2012 | same as 2012
106396 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 5 ml, individually packed | Tips for Multipette/Combi 100 pieces/pack - sterile, DNA- /RNase- /ATP- and endotoxin- free -suitable for pipettes offered |  | Yes | 106115 | Optional | 100 | 6 | 34 | same as 2012 | same as 2012
106397 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 12.5ml- bag of 10 | Tips for Multipette/Combi 12.5 ml - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered |  | Yes | 106115 | Regular needed | 100 | 13 | 11 | 0 | 0
106398 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 12.5 ml - individually packed | Tips 12.5 ml - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered |  | Yes | 106115 | Optional | 100 | 13 | 11 | same as 2012 | same as 2012
106399 | BIOSAFETY | Consumables | Plastic & Glassware | Plastic bags made from PP | small disposable plastic bags made from PP, at least 40 µm thick for tripod (approx. 200 mm width and 300 mm length), liquid tight, for not infections material (PCR, material from test kits), 100 pieces per pack (or roles with 100 pieces, perforated for teare off) |  | Yes | 106349 | Always needed | 100 | 98 | 777 | 1322 | 1322
106400 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Forceps, individually wrap, sterile | Forceps, preferably plastic for handling strips, individually wrap, sterile |  | Yes |  | Always needed | 100 | 19 | 86 | 209 | 217
106401 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Parafilm sealing film | Parafilm, 100 mm width, 58 m length, for use between - 40 °C to +50 °C |  |  |  | Always needed | 1 | 8 | 12 | same as 2012 | same as 2012
106402 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | plastic bottles | plastic bottles for distilled water, alcohol |  |  |  | Regular needed | 1 | 19 | 83 | same as 2012 | same as 2012
106403 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 6-7 - None disposable | nitril gloves, size 6-7, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) |  | Yes |  | Regular needed | 12 | 38 | 47 | 50 | 50
106404 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 7-8 -- None disposable | nitril gloves, size 7-8, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) |  | Yes |  | Regular needed | 12 | 31 | 42 | 50 | 50
106405 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 8-9 - - None disposable | nitril gloves, size 8-9, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) |  | Yes |  | Regular needed | 12 | 20 | 18 | 20 | 20
106407 | SOLID CULTURE | Consumables | Microbiological tests and test systems | Brain Heart Infusion agar | 500 gr |  |  |  | Regular needed | 500 g | 11 | 61 | 61 | 61
106412 | MOLECULAR BIOLOGY | Consumables | Pipetting | Long 1 ml tips with filter | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 100 - 1000 µl, Lenght : 118 to 120 mm |  |  | 106075 or 106057 | Always needed | 100 | 39 | 322 | 787 | 828
106415 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Filter paper - sheets | Packing of 100 - 580x580 mm - Filtration time in acc. with DIN 53137: 35 to 70 s - Weight: 87 g / m2 |  |  |  | Always needed | 100 | 15 | 315 | 1134 | 1134
106419 | LIQUID CULTURE | Consumables | Pipetting | TIPS, PP, 20-200 μl, sterile, autoclavable with Filter | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 20 - 200 µl, |  | Yes | 106072 | Always needed | 960 | 42 | 115 | 133 | 137
106420 | LIQUID CULTURE | Consumables | Pipetting | TIPS, PP, 100-1000 μl, sterile, autoclavable with Filter | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 100 - 1000 µl, |  | Yes | 106075 | Always needed | 960 | 51 | 129 | 133 | 137
106421 | MOLECULAR BIOLOGY | Consumables | Tubes | Standard reaction tubes 1.5 ml | Standard reaction tubes 1.5 ml, conical with screw top (twist top vials) with O-ring (seal ring) with Uniform lid thickness, PP, autoclavable at 121 °C (20 min.), High clarity, Frosted marking area, subdivisions for approx. volume determination Can withstand centrifugation up to 20 000 g | Yes | Yes |  | Regular needed | 500 | 25 | 56 | same as 2012 | same as 2012
106422 | MOLECULAR BIOLOGY | Consumables | Tubes | Standard reaction tubes 2 ml | Standard reaction tubes 2 ml, conical with screw top (twist top vials) with O-ring (seal ring) with Uniform lid thickness, PP, autoclavable at 121 °C (20 min.), High clarity, Frosted marking area, subdivisions for approx. volume determination Can withstand centrifugation up to 20 000 g | Yes | Yes |  | Regular needed | 500 | 19 | 29 | same as 2012 | same as 2012
106434 | MOLECULAR BIOLOGY | Consumables | Pipetting | Disposable pasteur pipettes, graduated, non sterile, 155 mm, 3 ml | Unsterile, plastic Pasteur-pipettes, Length 155 mm, precision-formed, unbreakable, non toxic and with controlled, reproducible drop size (e.g : Carl roth EA65.1 or similar) |  | Yes |  | Always needed | 500 | 26 | 178 | 728 | 760
106435 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size S | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent |  | Yes |  | Always needed | 20 | 92 | 168 | 332 | 332
106436 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size M | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent |  | Yes |  | Always needed | 20 | 145 | 242 | 332 | 332
106437 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size L | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent |  | Yes |  | Always needed | 20 | 95 | 170 | 332 | 332
106440 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Shoe cover | Plastic, water resistant |  | Yes |  | Always needed | 1000 | 16 | 160 | 334 | 334
106441 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Hair Cover |  |  | Yes |  | Always needed | 2000 | 12 | 49 | 52 | 52
106444 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Plastic bags - 30L | Red color, autoclavable with Bio Hazardous Logo |  | Yes |  | Always needed | 200 | 69 | 155 | 29 | 29
106497 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 6-7 - Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. |  | Yes |  | Always needed | 100 |  | 200 | 200 | 200
106498 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 7-8 -Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. |  | Yes |  | Always needed | 100 |  | 200 | 200 | 200
106499 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 8-9 - Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. |  | Yes |  | Always needed | 100 |  | 200 | 200 | 200
106500 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - S | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106501 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - M | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106502 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - L | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106503 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - S | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106504 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - M | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106505 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - L | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106506 | PROCESSING | Consumables | Laboratory, Miscellaneous | Marker pen for PCR Tubes | Permanent - to use with PCR Tubes |  | Yes |  | Always needed | 1 |  | 400 | 400 | 400
106507 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Smock stick starter kit | To control the airflow of a Bio Saftey Cabinet (e.g. Dräger Smock stick starter kit - http://www.skcinc.com/prod/800-216.asp) |  |  |  | Optional | 1 |  | 100 | 75 | 25
106508 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Refill Smock stick pack - 10 pcs | To control the airflow of a Bio Saftey Cabinet - corresponding with starter kit |  |  |  | Optional | 1 |  | 100 | 75 | 100
106509 | PROCESSING | Consumables | Laboratory, Miscellaneous | Special pH indicator strips | Range from pH 6.5 to pH 10; Reading in small increments, especially around pH 8 the incremental reading should be 0.2 pH units, according to a provided colour scale; Colour not “bleeding” from the strip |  | Yes |  | Optional | 100 |  | 100 | 75 | 25

## 3 - Consumables Chemicals India
Item Number | Method | Family | Sub Family | Item Name | GDF SPECIFICATION | Detailled Specification (Y/N) | Request for sample (Y/N) | Link to | Importance | Unit / package GDF | Quantity of Pack - Historical data - 2011 - INDIA | Quantity of Pack - Historical data - 2012 - INDIA | Quantity of Pack - Forecasted data - 2013 - INDIA | Quantity of Pack - Forecasted data - 2014 - INDIA | Supplier Specifications | Primary Packing Type | Number of units per primary packing type | Secondary Packing Type | Number of units per secondary packing type | Minimum Shelf Life in days | Gross weight in kg | Depth (in cm) | Width (in cm) | Height (in cm) | Volume in M3 | Special requirement for delivery | Special requirement for storage | Manufacturer Name | Model or Item Number (Manufacturer) | Country of Origin | Custom Tariff code | Please mention quality standard under which the item is produced (if any) | Manufactured in India or Imported
106259 | SOLID CULTURE | Chemicals | Loewenstein-Jensen-Medium | Ready mixed TB | Ready mixed TB - Medium according to Loewenstein-Jensen (LJ), with correct composition of ingredients from a certified company, 500 g plastic bottles e.g. Merck |  |  |  | Always needed | 500 g | 0 | 111 | 93 | 93
106261 | SOLID CULTURE | Chemicals | Loewenstein-Jensen-Medium | Glycerol ≥ 99% | Glycerol ≥ 99% Ph.Eur., C3H8O3, MW: 92.1, 1 L plastic bottle |  |  |  | Regular needed | 1 L | 0 | 40 | 15 | 15
106262 | SOLID CULTURE | Chemicals | Loewenstein-Jensen-Medium | Sodium pyruvate, 50g; bottle |  |  |  |  | Optional | 25g
106263 | SOLID CULTURE | Chemicals | IUTM base | Potassium dihydrogen phosphate | Potassium dihydrogen phosphate, ≥ 99% p.a. anhydrous, KH2PO4, MW: 136,09, 1 kg plastic bottle |  |  |  | Always needed | 1 kg | 6 | 132 | 46 | 46
106264 | SOLID CULTURE | Chemicals | IUTM base | Magnesium sulfate-heptahydrate | Magnesium sulfate-heptahydrate, ≥ 99% p.a., MgSO4 • 7H2O, MW: 246,48. 500g plastic bottle |  |  |  | Regular needed | 500 g
106265 | SOLID CULTURE | Chemicals | IUTM base | L-Asparagine-momohydrate | L-Asparagine-momohydrate, C4H8N2O3 • H2O, MW: 150.13, ≥ 99%, plastic bottle, 250g |  |  |  | Regular needed | 100g
106266 | SOLID CULTURE | Chemicals | IUTM base | tri-Magnesium di-citrate nonahydrate | tri-Magnesium di-citrate nonahydrate, Mg3(C6H5O7)2 • 9 H2O, MW: 613,25, ≥ 95%, p.a., 500g plastic bottle |  |  |  | Regular needed | 100g
106267 | SOLID CULTURE | Chemicals | IUTM base | Malachite green oxalate | Malachite green oxalate, p. A.; 100 g bottle |  |  |  | Regular needed | 100 g
106268 | SOLID CULTURE | Chemicals | IUTM base | Glycerol | Glycerol |  |  |  | Optional | 1kg
106269 | SOLID CULTURE | Chemicals | Ogawa-Medium | Sodium-L-Glutamate-monohydrate | Sodium-L-Glutamate-monohydrate, C5H8NNaO4 • H2O, MW: 187.13, purum, ≥ 98%, 1 kg plastic bottle |  |  |  | Optional | 1 kg
106275 | SOLID CULTURE | Chemicals | Modified Ogawa-Medium | Magnesium sulphate-heptahydrate | Magnesium sulphate-heptahydrate |  |  |  | Optional | 1
106278 | SOLID CULTURE | Chemicals | Phosphate buffer according to Soerensen, 1/15 M, pH 6.8 or adjustable between pH 5 to 8.2 | di-Sodium hydrogen phosphate, anhydrous | di-Sodium hydrogen phosphate anhydrous (offered also with crystal water 2, 7, 12) Na2HPO4, MW: 141.96, p.a. for buffer, 1 kg plastic bottle |  |  |  | Always needed | 1 kg | 6 | 134 | 48 | 48
106279 | SOLID CULTURE | Chemicals | Phosphate buffer according to Soerensen | di-Sodium hydrogen phosphate dodecahydrate | Na2HPO4 12 H2O , MW: 358,14, p.a. for buffer, 1 kg plastic bottle |  |  |  | Optional | 1 kg
106281 | PROCESSING | Chemicals | Petroff Method | Sodium hydroxide | Sodium hydroxide, NaOH, MW 40.00, purum ≥ 98%, pellets, 1kg plastic bottle, tight screw cap (hygroscopic and CO2 binding from the air) |  |  |  | Always needed | 1 | 24 | 40 | 17 | 17
106282 | PROCESSING | Chemicals | Sodium phosphate Method | tri-Sodium phosphate dodecahydrate | tri-Sodium phosphate dodecahydrate, Na3PO4 • 12H2O, MW: 380.12, (1 kg, 2.5 kg, 25 kg) 2.5 kg plastic bottle |  |  |  | Optional | 2.5 kg
106283 | PROCESSING | Chemicals | NALC Method | N-acetyl-L-cysteine (NALC) | N-acetyl-L-cysteine (NALC), C5H9NO3S, MW: 163.19,puriss. ≥ 99%, 250g plastic bottle |  |  |  | Always needed | 100g | 85 | 185 | 42 | 42
106284 | PROCESSING | Chemicals | NALC Method | tri-Sodium citrate dihydrate | tri-Sodium citrate dihydrate, C6H5Na3O7 • 2H2O, MW: 294.10, plastic bottle 1 kg |  |  |  | Always needed | 1 kg | 24 | 80 | 13 | 13
106286 | PROCESSING | Chemicals | Cethypyridinum chloride for transit | Cetylpyridinium chloride monohydrate | Cetylpyridinium chloride monohydrate, C21H38ClN • H2O, MW: 358.01, melting point 77 ºC, purity ≥ 98% for biochemistry, (100 g /500g) 500g bottle |  |  |  | Optional | 500 g
106287 | PROCESSING | Chemicals | Cethypyridinum chloride for transit | Sodium chloride | Sodium chloride, NaCl, MW: 58.44, Ph.Eur. ≥ 99%, 1kg plastic bottle |  |  |  | Optional | 1 kg
106288 | PROCESSING | Chemicals | Pre-treatment of urine (enrichment before decontamination) | 5-Sulfosalicylic acid dihydrate | 5-Sulfosalicylic acid dihydrate, C6H7NO3S• 2H2O, MW: 254,22, p.a. ≥ 99%, 100g plastic bottle |  |  |  | Optional | 100 g
106292 | IDENTIFICATION | Chemicals | Catalase Test | Tween® 80 | Tween® 80, none ionic detergent, (Polyoxyethlene-80-sorbitan monooleate), Ph.Eur., (250g/ 1 kg) 1kg plastic bottle |  |  |  | Optional | 1 kg
106293 | IDENTIFICATION | Chemicals | Catalase Test | Hydrogen peroxide | Hydrogen peroxide, H2O2, MW: 34.01, 30%, p.a., ISO, stabilised, 250 ml plastic bottle |  |  |  | Optional | 250 ml
106295 | IDENTIFICATION | Chemicals | Nitrate Reduction Test (Nitratase Test) | di-Sodium hydrogen phosphate | di-Sodium hydrogen phosphate |  |  |  | Optional | 1
106297 | IDENTIFICATION | Chemicals | Nitrate Reduction Test (Nitratase Test) | Sulphanilamide | Sulphanilamide, C6H8N2O2S, MW: 172.21, p.a. ≥ 98%, melting point > 164 °C, 100g plastic bottle |  |  |  | Optional | 100 g
106298 | IDENTIFICATION | Chemicals | Nitrate Reduction Test (Nitratase Test) | N-(1-Naphthyl)-ethylendiamine dihydrochloride in glass bottle | N-(1-Naphthyl)-ethylendiamine dihydrochloride, C12H14N2 • 2HCl, MW: 259.18, p.a.≥ 98%, 25g plastic bottle |  |  |  | Optional | 25 g
106301 | IDENTIFICATION | Chemicals | identification | Cyanogen bromide |  |  |  |  | Optional | 1
106302 | IDENTIFICATION | Chemicals | identification | Analine / o-toluidine |  |  |  |  | Optional | 1
106303 | MEDIA KITCHEN | Chemicals | Storage of Mycobacteria in a deep freezer | Skim milk powder | Skim milk powder, 500g (e.g. Oxoid) |  |  |  | Optional | 500 g
106307 | DST | Chemicals | Solvents | Dimethyl sulfoxide | Dimethyl sulfoxide, 99,8%, p. a., 500 ml |  |  |  | Optional | 10 x 25 ml
106308 | DST | Chemicals | First line drugs (FLD) | Isoniazid | 5g Isoniazid (INH) (e.g. Sigma I13377-5G); |  |  |  | Optional | 5 g
106309 | DST | Chemicals | First line drugs (FLD) | Rifampicin | 1g Rifampicin (RMP) (e.g. Sigma R3501-1G) |  |  |  | Optional | 1 g
106310 | DST | Chemicals | First line drugs (FLD) | Ethambutol | 25 g Ethambutol ( e.g. Sigma E4630-25G) |  |  |  | Optional | 25 g
106311 | DST | Chemicals | First line drugs (FLD) | Dihydro-streptomycin | 5g Dihydro-streptomycin (DHSM (e.g. Sigma D7253-5G) |  |  |  | Optional | 5g
106312 | DST | Chemicals | First line drugs (FLD) | Nicotinic acid amide | Nicotinic acid amide (nicotinamid, nicocinamid), pure powder, (may replace PZA at neutral pH in LJ) |  |  |  | Optional | 1
106313 | DST | Chemicals | Second Line Drugs (SLD) | Ofloxacin | Ofloxacin/Ciprofloxacin, (OFL/CIP), (e.g. Aventis/Bayer or Sigma O8757-1G/17850-5G-F ) |  |  |  | Optional | 1
106314 | DST | Chemicals | Second Line Drugs (SLD) | Moxifloxacin | Moxifloxacin (MOX), 2g (e.g.Bayer) |  |  |  | Optional | 1g
106315 | DST | Chemicals | Second Line Drugs (SLD) | Protionamid, | Protionamid, (PTH), 2g (e,g, Fatol) |  |  |  | Optional | 2 g
106316 | DST | Chemicals | Second Line Drugs (SLD) | Ethionamide | Ethionamid (ETH), 5g, (e.g. Sigma E6005-5G) |  |  |  | Optional | 5 g
106317 | DST | Chemicals | Second Line Drugs (SLD) | Kanamycin | Kanamycin (KAN), 1g, (e.g. Sigma K1876-1G) |  |  |  | Optional | 1 g
106318 | DST | Chemicals | Second Line Drugs (SLD) | Amikacin | Amikacin (AMI), 5g (e.g. Sigma A3650-5G) |  |  |  | Optional | 5 g
106319 | DST | Chemicals | Second Line Drugs (SLD) | Capreomycin | Capreomycin (CM), 1g (e.g. Sigma C4142-1G) |  |  |  | Optional | 1 g
106320 | DST | Chemicals | Second Line Drugs (SLD) | Para-aminosalicylic acid | para-aminosalicylic acid (PAS), 5g (e.g. Sigma A79604-5G) |  |  |  | Optional | 1
106321 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Agarose | Agarose, universal use for separation of DNA fragments, range 100 bp to 20 kbp, melting point <90 °C, gel formation < 45 °C, 100 g |  |  |  | Optional | 100 g
106322 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Boric acid | Boric acid, > 99.5%,1 kg for buffer |  |  |  | Optional | 1 kg
106323 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Sodium acetate trihydrate | Sodium acetate trihydrate, analytical grade, 1 kg |  |  |  | Optional | 1 kg
106324 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Ethidium bromide powder | Ethidium bromide, 1 g |  |  |  | Optional | 1 g
106325 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | TRIS | TRIS, > 99%, buffer grade, 1 kg, |  |  |  | Optional | 1 kg
106326 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | EDTA | EDTA, 99%, for biochemistry, 250 g |  |  |  | Optional | 250 g
106327 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Gel load buffer | Gel load buffer, 6x, with Ficoll and bromphenol blue/xylencyan blue, 6 x 1.8 ml fractions (e.g. see Roth) |  |  |  | Optional | 6
106328 | WASTE MANAGEMENT | Chemicals | Fumigation (BSC before filter change and/or emergency case) | Formalin | Formalin, CH2O, MW: 30.03, technical grade, ~37% stabilised with 10% methanol, 1 L plastic bottle |  |  |  | Regular needed | 1 l | 120 | 30
106329 | WASTE MANAGEMENT | Chemicals | Disinfectants | Stable chlorine disinfectant | Stable cholrine disenfectant in tablets to prepare 1% solution free chlorine. Unit per pack to prepare at least 10 liters final solution. |  |  |  | Always needed | 100 | 240 | 20
106331 | WASTE MANAGEMENT | Chemicals | Disinfectants | Phenol | Phenolic products e.g. triclosan, chloroxylenol, orthophenylphenol, to be used at concentrations of 5%, 2.5 L bottles |  |  |  | Regular needed | 1kg |  |  | 264 | 264
106333 | WASTE MANAGEMENT | Chemicals | Disinfectants | Isopropanol | Isopropanol (2-propanol), 5 L bottles |  |  |  | Always needed | 5 L | 40 | 10 | 736 | 736
106334 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | DNA ladder | Control for separation of DNA fragments |  |  |  | Optional | 1
106336 | MOLECULAR BIOLOGY | Chemicals | Gel Electrophoresis | Ethidium bromide solution, | 500µg/mL water, gel electrophoresis and DNA isolation procedures suitable-5ML |  |  |  | Optional | 5 ml
106468 | LIQUID CULTURE | Chemicals | McFarland Standard | Mc Farland Standards kits | at least stable for 1 year after delivery |  |  |  | Optional | 1
106113 | LIQUID CULTURE | Consumables | Plastic & Glassware | Petri-dishes plastic | 94 x 16 mm - Autoclavable - Sterile - pack in 10 or 20 pieces per bag |  |  |  | Regular needed | 480 | 0 | 47 | 16 | 16
106181 | LIQUID CULTURE | Consumables | Pipetting | Dispenser-tips Universal 10ml sterile | Dispenser-tips Universal 10ml sterile - DNA- /RNase- /ATP- and endotoxin free - suitable for pipettes offered |  | Yes | 106080 | Always needed | 100 | 96 | 230 | 339 | 339
106209 | LIQUID CULTURE | Consumables | Pipetting | Serological pipette - 5 ml | Serological pipette -5,0 ml, graduation 0.1 ml, packaging single packed |  |  |  | Optional | 200
106210 | LIQUID CULTURE | Consumables | Pipetting | Serological pipette - 10 ml | Serological pipette - 10 ml, graduation 0.1 ml, packaging single packed |  |  |  | Optional | 250
106236 | SOLID CULTURE | Consumables | Laboratory, Miscellaneous | Disposable loops 10 µl | Disposable loops 10µl , sterile, length approx. 220 mm, pack of 10 |  |  | 106232 | Always needed | 500 | 0 | 64 | 56 | 56
106254 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Molecular grade water | DNAase, RNAase free water - not chemically treated molecular grade water - 1.7mlx10 or 10ml |  | Yes |  | Always needed | 10x1.7ml | 165 | 352 | 1045 | 924
106294 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Disinfectant for cleaning instrument | TB Certified concentrated disinfectant e.g. BOMIX |  |  |  | Always needed | 5 L | 0 | 80 | 80 | 80
106304 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Hand pump for Disinfectant for cleaning instrument | Hand pump for cleaning instrument 10ml dosage |  |  |  | Regular needed | 1
106338 | SOLID CULTURE | Consumables | Tubes | Culture tubes with screw cap | Culture tubes, diameter 16 mm, length 125 mm, wall strength 1mm, thread outside; inner diameter min.13 mm; borosilicate glass with ISO winding, autoclavable screw caps with gasket out of PP or Teflon. Some prefer 18 mm diameter |  | Yes |  | Regular needed | 100 | 0 | 18 | 167 | 167
106339 | SOLID CULTURE | Consumables | Tubes | Universal glass bottle | 28 ml universal glass bottle (Mc Cartney), heavy duty glass, straight neck without shoulder, aluminium screw cap with rubber lining or other sterilisable screw cap with sealing gasket |  |  |  | Optional | 1
106340 | MOLECULAR BIOLOGY | Consumables | Tubes | PP-tubes for centrifuge non sterile 50 ml | PP-tubes for centrifuge, 50 ml with screw cap, tight, size 115x30 mm (Falcon or equivalent), 500 pieces/pack - Resistance minimum 3,200 g in a swing out bucket | Yes | Yes |  | Regular needed | 500 | 0 | 45 | 95 | 84
106341 | LIQUID CULTURE | Consumables | Tubes | PP-tubes for centrifuge, sterile, 50 ml | PP-tubes for centrifuge, 50 ml with screw cap, tight, size 115x30 mm (Falcon or equivalent), sterile, packed 25 in racks, 20 racks per pack. | Yes | Yes |  | Always needed | 500 | 641 | 784 | 366 | 367
106342 | LIQUID CULTURE | Consumables | Tubes | PP-tubes for centrifuge, sterile, 15 ml | PP-tubes for centrifuge, 15 ml with screw cap, tight, size 120x17 mm, 500 pieces/pack - Resistance minimum 3,200 g in a swing out bucket |  | Yes |  | Regular needed | 500
106343 | LIQUID CULTURE | Consumables | Pipetting | Single use plastic Pasteur-pipettes sterile pack by 10 | sterile, single use plastic Pasteur-pipettes volume in a range of 1 to 3 ml, packed by 10, graduated 0.5ml, lenght : minimum 155mm ideal 170mm. Drop size maximum 30ul |  | Yes |  | Regular needed | 500
106344 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, 2 ml | Cryo-vial, sterile with standard screw cap, 2 ml round with outer winding, PP, graduated, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). pack of 1000 |  | Yes |  | Regular needed | 1000
106345 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size S | powder-free disposable latex gloves, size S, >0,16 mm thickness, length >240 mm, 1000 per pack |  | Yes |  | Always needed | 1000 | 144 | 356 | 264 | 264
106346 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size M | powder-free disposable latex gloves, size M, >0,16 mm thickness,length >240 mm, 1000 per pack |  | Yes |  | Always needed | 1000 | 216 | 362 | 264 | 264
106347 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size L | powder-free disposable latex gloves, size L, >0,16 mm thickness, length >240 mm, 1000 per pack |  | Yes |  | Always needed | 1000 | 72 | 203 | 264 | 264
106348 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Transparent polypropylene bag | Transparent polypropylene bag for waste bins, at least 40 µm thickness, up to 134 °C steam sterilisable, approx 24 L, approx. 540 x 550 mm, in accordance with countiy's norms for environmental and hygiene assignments, 1000 pieces per pack |  | Yes |  | Always needed | 1000 | 96 | 85 | 32 | 32
106350 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, 1.5 ml | Cryo-vial, sterile with standard screw cap, 1.5 ml, PP with a medical-grade silicone O-ring seal, style conical, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). It can be centrifuged up to 17 000 x g - e.g. SIGMA V4131 |  | Yes |  | Always needed | 1000 | 200 | 816 | 741 | 655
106351 | LIQUID CULTURE | Consumables | Laboratory, Miscellaneous | Single use syringes, sterile | Sterile, single use syringes according to ISO 7886-1 with Luer connection, PP, graduated with resistant colour, safe piston lock, volume 20 ml, individually wrapped, 100 pieces per pack |  |  |  | Always needed | 50 | 144 | 12
106352 | LIQUID CULTURE | Consumables | Laboratory, Miscellaneous | Syringe filter for single use, sterile | Syringe filter for single use, entrance Luer connection, hydrophilic (e.g. PVDF filter in PE housing), pore size ≤ 0.22 µm, sterile and individually wrapped, at least 25 mm diameter, pressure up to 5 bar, 100 pieces per pack |  |  |  | Always needed | 50 | 288 | 24
106353 | MEDIA KITCHEN | Consumables | Laboratory, Miscellaneous | Filter paper round 185 mm | Filter paper round, weight approx. 75 g/m², retention > 4µm, diameter 185 mm, packs of 100 e.g. Whatman |  |  |  | Optional | 100
106354 | MEDIA KITCHEN | Consumables | Laboratory, Miscellaneous | Filter paper round 150 mm | Filter paper, round, weight approx. 75 g/m², retention >4 µm, diameter 150 mm, packs of 100 e.g. Whatman |  |  |  | Optional | 100
106355 | PROCESSING | Consumables | Laboratory, Miscellaneous | Marker pen | Marker pen, resistant against water and alcohol, also against autoclaving, writing on nearly all surfaces, even when cold and wet, fast drying, black, line 0.8 - 1 mm broad, black |  | Yes |  | Regular needed | 1 | 0 | 129 | 1287 | 1138
106356 | PROCESSING | Consumables | Laboratory, Miscellaneous | Adhesive labels | Adhesive labels on rolls or sheets (e.g. 50 x 19 mm, 24 per sheet and 200 sheets per box = 4800 label) |  | Yes |  | Regular needed | 1
106357 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Cryo-tags | Cryo-tags sized to fit for use on cryo-tubes and standard micro-tubes (approx.32 mm x 12 mm), made from smear resistant material, but should accept all marking devices, rolles of 1000 |  | Yes |  | Always needed | 1 | 120 | 864 | 750 | 664
106358 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Sterile indicator tape, hot air oven | Sterile indicator tape (roles), approx.19mm broad and 55m long, DIN EN 867-5, for hot air sterilize (green/brown) |  | Yes |  | Always needed | 1 | 72 | 394 | 320 | 320
106359 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Sterile indicator tape, autoclave | Sterile indicator tape (roles), approx.19mm broad and 55m long, DIN EN 867-5, for autoclave (beige/dark brown) |  | Yes |  | Always needed | 1 | 72 | 394 | 320 | 320
106360 | CLEANING KITCHEN | Consumables | Laboratory, Miscellaneous | Aluminium foil | Aluminium foil, 30 µm thickness, 600 mm width, 100 m long |  | Yes |  | Regular needed | 1 |  |  | 80 | 80
106361 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Parafilm with dispenser | Parafilm, 100 mm width, 58 m length, for use between - 40 °C to +50 °C, with dispenser |  |  |  | Optional | 1
106362 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Plastic-foil | Plastic-foil rolls, 30 cm |  |  |  | Regular needed | 1 |  |  | 80 | 80
106363 | CLEANING KITCHEN | Consumables | Laboratory, Miscellaneous | Tube brush | Tube brush, total length min.280 mm, diameter 15-16 mm, brush length 80 mm - too fit witt culture tubes. Soft at the end to not brake the tube |  | Yes |  | Regular needed | 5 |  |  | 56 | 56
106364 | CLEANING KITCHEN | Consumables | Laboratory, Miscellaneous | Brushes | Brushes for glassware, total length approx.480 mm, diameter 60 mm, brush length 120 mm |  |  |  | Optional | 1
106366 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size L - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent |  | Yes |  | Always needed | 100 | 270 | 57 | 35 | 35
106367 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size M - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent |  | Yes |  | Always needed | 100 | 270 | 129 | 66 | 66
106368 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size S - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent |  | Yes |  | Always needed | 100 | 0 | 74 | 35 | 35
106370 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Disinfectant for floors | Disinfectant for floors, tuberculozid, locally made or commercially available as concentrate |  |  |  | Always needed | 1 | 0 | 117 | 80 | 80
106371 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Disinfectant for BSCs surface | Disinfectant for BSCs (Biosafety Cabinets) and surfaces, (alcohol containing disinfectant, certified for TB bacilli, disinfection within 1 minute) e.g. 2-propanol (iso-propanol) , good beddabling, evaporates completely and residuelessly; |  |  |  | Always needed | 5 L | 0 | 327 | 320 | 320
106372 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Spray head and 1L bottles | Spray head with 1 liter bottle |  |  |  | Always needed | 1 set | 160 | 177 | 80 | 80
106373 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Liquid soap | Liquid soap, pH neutral, bottle of 1 litre |  |  |  | Always needed | 1 L |  |  | 480 | 480
106374 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Disinfectant for hands | Disinfectant for hands 1 litre bottle, alcoholic disinfectant locally made or commercially available |  |  |  | Always needed | 1 L | 0 | 360 | 480 | 480
106375 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Single-use paper towels | Single-use paper towels, approx. 230 x 300 mm folded to take as single pieces off a wall mounted box. 150 towels/pack, 30 packs/ carton |  | Yes |  | Always needed | 4500 | 240 | 363 | 254 | 254
106376 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Cotton wool | Cotton wool, white adsorbent, as kg, purchase locally |  |  |  | Always needed | 1 |  |  | 160 | 160
106377 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Tissue pulp | Tissue pulp, absorbent sheets, approx. 550 x 350 mm |  |  |  | Always needed | 1 |  |  | 845 | 845
106378 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Washing Powder | detergent/washing Powder for laboratory dish washer will be needed (10 kg trading unit) |  |  |  | Optional | 10 kg
106379 | CLEANING KITCHEN | Consumables | Disinfectants and Cleaning | Salt for dish washer | Salt for dish washer, 10 kg/pack |  |  |  | Optional | 12kg
106380 | CLEANING KITCHEN | Consumables | Disinfectants and Cleaning | Rinse aid for dish washer | Rinse aid for dish washer, 1 litre |  |  |  | Optional | 1 L
106383 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, for one hand operation | Cryo-vial, sterile with Tethered Screw Caps, 2 ml round with outer winding PP, graduated, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). pack of 1000 |  | Yes |  | Always needed | 1000 | 24 | 60 | 9 | 9
106384 | MOLECULAR BIOLOGY | Consumables | Tubes | PP-tubes for centrifuge, non sterile, 15 ml | PP-tubes for centrifuge, 15 ml with screw cap, tight, size 120x17 mm, 500 pieces/pack |  | Yes |  | Regular needed | 500 | 0 | 14 | 76 | 68
106385 | MOLECULAR BIOLOGY | Consumables | Tubes | PCR tubes | PCR tubes, the tubes should be extremely thin to ensures speedy and even heat transfer of every sample, pyrogen free, DNase- /RNase-free and autoclavable, 0,2 ml with attached caps, fits all standard thermocyclers with heated lids especially the one provided. |  | Yes |  | Always needed | 1000 | 136 | 283 | 390 | 344
106386 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 0.1 - 10 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 0.1 - 10 µl, |  | Yes | 106053 | Always needed | 480 | 247 | 1012 | 1583 | 1399
106387 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 1.0 - 20 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 1.0 - 20 µl, |  | Yes | 106054 | Regular needed | 480 | 338 | 32
106388 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 10 - 100 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered, 10 - 100 µl, |  | Yes | 106055 | Regular needed | 960
106389 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 20 - 200 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 20 - 200 µl, |  | Yes | 106056 | Always needed | 960 | 236 | 794 | 1189 | 1051
106390 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS 100 - 1000 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 100 - 1000 µl, |  | Yes | 106057 | Always needed | 960 | 349 | 327 | 406 | 359
106394 | LIQUID CULTURE | Consumables | Pipetting | Single use plastic Pasteur-pipettes sterile individually packed | sterile, single use plastic Pasteur-pipettes volume in a range of 1 to 3 ml, individually packed, graduated 0.5ml, lenght : minimum 155mm ideal 170mm. Drop size maximum 30ul |  | Yes |  | Always needed | 500 | 1115 | 330 | 305 | 306
106395 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 5 ml - bag of 10 | Tips for Multipette/Combi 100 pieces/pack - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered |  | Yes | 106115 | Regular needed | 100
106396 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 5 ml, individually packed | Tips for Multipette/Combi 100 pieces/pack - sterile, DNA- /RNase- /ATP- and endotoxin- free -suitable for pipettes offered |  | Yes | 106115 | Optional | 100
106397 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 12.5ml- bag of 10 | Tips for Multipette/Combi 12.5 ml - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered |  | Yes | 106115 | Regular needed | 100
106398 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 12.5 ml - individually packed | Tips 12.5 ml - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered |  | Yes | 106115 | Optional | 100
106399 | BIOSAFETY | Consumables | Plastic & Glassware | Plastic bags made from PP | small disposable plastic bags made from PP, at least 40 µm thick for tripod (approx. 200 mm width and 300 mm length), liquid tight, for not infections material (PCR, material from test kits), 100 pieces per pack (or roles with 100 pieces, perforated for teare off) |  | Yes | 106349 | Always needed | 100 | 240 | 1264 | 846 | 846
106400 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Forceps, individually wrap, sterile | Forceps, preferably plastic for handling strips, individually wrap, sterile |  | Yes |  | Always needed | 100 | 360 | 201 | 190 | 168
106401 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Parafilm sealing film | Parafilm, 100 mm width, 58 m length, for use between - 40 °C to +50 °C |  |  |  | Always needed | 1 | 72 | 6
106402 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | plastic bottles | plastic bottles for distilled water, alcohol |  |  |  | Regular needed | 1
106403 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 6-7 - None disposable | nitril gloves, size 6-7, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) |  | Yes |  | Regular needed | 12 | 600 | 50
106404 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 7-8 -- None disposable | nitril gloves, size 7-8, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) |  | Yes |  | Regular needed | 12 | 565 | 40
106405 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 8-9 - - None disposable | nitril gloves, size 8-9, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) |  | Yes |  | Regular needed | 12
106407 | SOLID CULTURE | Consumables | Microbiological tests and test systems | Brain Heart Infusion agar | 500 gr |  |  |  | Regular needed | 500 g |  |  | 40 | 40
106412 | MOLECULAR BIOLOGY | Consumables | Pipetting | Long 1 ml tips with filter | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 100 - 1000 µl, Lenght : 118 to 120 mm |  |  | 106075 or 106057 | Always needed | 100 | 486 | 712 | 973 | 860
106415 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Filter paper - sheets | Packing of 100 - 580x580 mm - Filtration time in acc. with DIN 53137: 35 to 70 s - Weight: 87 g / m2 |  |  |  | Always needed | 100 | 100 | 1033 | 740 | 740
106419 | LIQUID CULTURE | Consumables | Pipetting | TIPS, PP, 20-200 μl, sterile, autoclavable with Filter | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 20 - 200 µl, |  | Yes | 106072 | Always needed | 960 | 14 | 47 | 51 | 51
106420 | LIQUID CULTURE | Consumables | Pipetting | TIPS, PP, 100-1000 μl, sterile, autoclavable with Filter | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 100 - 1000 µl, |  | Yes | 106075 | Always needed | 960 | 16 | 47 | 51 | 51
106421 | MOLECULAR BIOLOGY | Consumables | Tubes | Standard reaction tubes 1.5 ml | Standard reaction tubes 1.5 ml, conical with screw top (twist top vials) with O-ring (seal ring) with Uniform lid thickness, PP, autoclavable at 121 °C (20 min.), High clarity, Frosted marking area, subdivisions for approx. volume determination Can withstand centrifugation up to 20 000 g | Yes | Yes |  | Regular needed | 500
106422 | MOLECULAR BIOLOGY | Consumables | Tubes | Standard reaction tubes 2 ml | Standard reaction tubes 2 ml, conical with screw top (twist top vials) with O-ring (seal ring) with Uniform lid thickness, PP, autoclavable at 121 °C (20 min.), High clarity, Frosted marking area, subdivisions for approx. volume determination Can withstand centrifugation up to 20 000 g | Yes | Yes |  | Regular needed | 500
106434 | MOLECULAR BIOLOGY | Consumables | Pipetting | Disposable pasteur pipettes, graduated, non sterile, 155 mm, 3 ml | Unsterile, plastic Pasteur-pipettes, Length 155 mm, precision-formed, unbreakable, non toxic and with controlled, reproducible drop size (e.g : Carl roth EA65.1 or similar) |  | Yes |  | Always needed | 500 | 0 | 1099 | 1696 | 1504
106435 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size S | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent |  | Yes |  | Always needed | 20 | 0 | 370 | 208 | 208
106436 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size M | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent |  | Yes |  | Always needed | 20 | 960 | 450 | 208 | 208
106437 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size L | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent |  | Yes |  | Always needed | 20 | 960 | 298 | 208 | 208
106440 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Shoe cover | Plastic, water resistant |  | Yes |  | Always needed | 1000 | 42 | 308 | 212 | 212
106441 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Hair Cover |  |  | Yes |  | Always needed | 2000 | 25 | 73 | 27 | 27
106444 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Plastic bags - 30L | Red color, autoclavable with Bio Hazardous Logo |  | Yes |  | Always needed | 200 | 0 | 77 | 12 | 12
106497 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 6-7 - Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. |  | Yes |  | Always needed | 100 |  | 200 | 200 | 200
106498 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 7-8 -Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. |  | Yes |  | Always needed | 100 |  | 200 | 200 | 200
106499 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 8-9 - Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. |  | Yes |  | Always needed | 100 |  | 200 | 200 | 200
106500 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - S | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106501 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - M | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106502 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - L | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106503 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - S | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106504 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - M | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106505 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - L | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning |  | Yes |  | Regular needed | 1 |  | 200 | 100 | 100
106506 | PROCESSING | Consumables | Laboratory, Miscellaneous | Marker pen for PCR Tubes | Permanent - to use with PCR Tubes |  | Yes |  | Always needed | 1 |  | 400 | 400 | 400
106507 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Smock stick starter kit | To control the airflow of a Bio Saftey Cabinet (e.g. Dräger Smock stick starter kit - http://www.skcinc.com/prod/800-216.asp) |  |  |  | Optional | 1 |  | 100 | 75 | 25
106508 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Refill Smock stick pack - 10 pcs | To control the airflow of a Bio Saftey Cabinet - corresponding with starter kit |  |  |  | Optional | 1 |  | 100 | 75 | 100
106509 | PROCESSING | Consumables | Laboratory, Miscellaneous | Special pH indicator strips | Range from pH 6.5 to pH 10; Reading in small increments, especially around pH 8 the incremental reading should be 0.2 pH units, according to a provided colour scale; Colour not “bleeding” from the strip |  | Yes |  | Optional | 100 |  | 100 | 75 | 25

## 4 - Instrument and tools India
Item Number | Method | Family | Sub Family | Item Name | GDF SPECIFICATION | Detailled Specification (Y/N) | Request for sample (Y/N) | Link to | Importance | Unit / package GDF | Quantity to equip 1 lab | Supplier Specifications | Primary Packing Type | Number of units per primary packing type | Secondary Packing Type | Number of units per secondary packing type | Minimum Shelf Life in days | Gross weight in kg | Depth (in cm) | Width (in cm) | Height (in cm) | Volume in M3 | Special requirement for delivery | Special requirement for storage | Manufacturer Name | Model or Item Number (Manufacturer) | Country of Origin | Custom Tariff code | Voltage / Amp / Frequency | Please mention quality standard under which the item is produced (if any) | Manufactured in India or Imported
106045 | MOLECULAR BIOLOGY | Instrument | Hoods, Cabinets, Workstations | PCR-Workstation, UV light | PCR-Workstation, UV light, Glass with UV/Air Cleaner e.g. KISKER - L020-GC | Yes |  |  | Always needed | 1 | 2
106074 | MOLECULAR BIOLOGY | Instrument | Centrifuges | Microfuge / minispin for 0.2-2ml tubes for short-rapid spin down of reagents | Resistant to bleach |  |  |  | Always needed | 1 | 1
106081 | LIQUID CULTURE | Instrument | Laboratory, Miscellaneous | UPS - 3000 VA - 2100W | UPS with battery pack for 20 minutes at 2100W | Yes |  | 106013 | Always needed | 1 | 2
106128 | MEDIA KITCHEN | Instrument | Weighing & Analytical Balances | Electronic analytical balance | Semimicro balance (electronic analytical balance), EC stamp-approved version, with built-in motorized weight, dynamic temperature compensation, 8 segment display, stability indicator. Accuracy class I. Readability 0.01 mg up to 60 g and 0.1 mg up to approx. 200 g. Tara range (subtractive) 200 g. | Yes |  |  | Optional | 1 | 1
106130 | MEDIA KITCHEN | Instrument | Meters | Bench top pH meter | Bench top pH meter with temperature compensation, complete set with electrode and 3 calibration buffers, resolution 0.01pH | Yes |  |  | Always needed | 1 | 1
106132 | MEDIA KITCHEN | Instrument | Mixers, Shakers, Stirrers | Magnetic stirrer with heating plate | Magnetic stirrer with heating plate | Yes |  |  | Regular needed | 1 | 1
106133 | MEDIA KITCHEN | Instrument | Mixers, Shakers, Stirrers | Magnetic bars set of 10 | Set of round shaped PTFE coated magnetic stirring bars, set of 10 magnetic bares, 1x 15 mm, 2X 20 mm, 2x 25 mm, 1x 30 mm, 2x 40 mm, 2x 50 mm length |  |  |  | Regular needed | 1 set of 10 | 1
106140 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Laboratory Floor standing refrigerator 140L | Floor standing refrigerator, at least energy efficiency class A+ , approx. 150 litre | Yes |  |  | Always needed | 1 | 1
106205 | LIQUID CULTURE | Instrument | Laboratory, Miscellaneous | UV-light illuminator - Liquid culture - 365 nm | UV-light illuminator at least 200 x 200 mm window |  |  |  | Regular needed | 1 | 1
106424 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Laboratory Freezer 140 L |  | Yes |  |  | Always needed | 1 | 2
106443 | MOLECULAR BIOLOGY | Instrument | Laboratory, Miscellaneous | UPS - 700 VA - 500 W | UPS with battery pack for 20 minutes at 700W | Yes |  | 106046 | Always needed | 1 | 3
106451 | LIQUID CULTURE | Instrument | Laboratory, Miscellaneous | Densitometer | Densitometer with 1 adaptator for MGIT tubes (15-18 mm diameter - 120 mm length) + 1 set of Mc Farland standards (0.5, 1.0, 2.0, 3.0,4.0) |  |  |  | Regular needed | 1set | 1
106051 | MOLECULAR BIOLOGY | Tools | Tubes | Desktop Deep freeze storage box with lid for 0.2ml PCR tubes | Deep freeze storage box with lid for 0.2ml PCR tubes, autoclavable PP. Suitable for storing 0.2 ml single vessels, 8-strips, and 12-strips, and 96-well PCR plates. |  |  | 106385 | Always needed | 1 | 5
106052 | MOLECULAR BIOLOGY | Tools | Tubes | Desktop Deep freeze storage box with lid for 1.5 ml cryovials - Cardboard cryobox | Deep freeze storage box with lid for 1.5 ml cryovials,Cardboard, with 9x9 inlet |  |  | 106350 | Always needed | 1 | 10
106053 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 0.5 - 10 µl - LPA | 1-channel pipette, variable from 1 - 10 µl, - accuracy < 0.025 to 0.100, precision < 0.012 to 0.040 | Yes | Yes |  | Always needed | 1 | 4
106056 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-chanel pipette, variable from 20 -200 ul - LPA | 1-chanel pipette, variable from 20 -200 ul - accuracy < 0.50 to 1.60, precision < 0.20 to 0.30 | Yes | Yes |  | Always needed | 1 | 4
106057 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 200 - 1000 µl - LPA | 1-channel pipette, variable from 100 - 1000 µl - accuracy < 3.00 to 8.00, precision < 0.60 to 1.50 | Yes | Yes |  | Always needed | 1 | 4
106068 | MOLECULAR BIOLOGY | Tools | Tubes | Stand (96) for 0.2 ml pcr tubes / PCR Rack | Stand (96) for 0.2 ml, plastic resistant to bleach and alcoohol at least |  | Yes | PCR Tubes | Always needed | set of 10 | 3
106072 | LIQUID CULTURE | Tools | Pipetting | 1-chanel pipette, variable from 20 -200 ul - LC | 1-chanel pipette, variable from 20 -200 ul - accuracy < 0.50 to 1.60, precision < 0.20 to 0.30 | Yes | Yes |  | Regular needed | 1 | 3
106075 | LIQUID CULTURE | Tools | Pipetting | 1-channel pipette, variable from 200 - 1000 µl - LC | 1-channel pipette, variable from 100 - 1000 µl - accuracy < 3.00 to 8.00, precision < 0.60 to 1.50 | Yes | Yes |  | Regular needed | 1 | 4
106089 | BIOSAFETY | Tools | Disinfectants and Cleaning | Plastic-foil welder 30 cm | Plastic-foil welder 30 cm - Simple robust plastic welder for the welding of PE, PP, Soft-PVC and Polyamide-film/bags. |  |  |  | Regular needed | 1 | 1
106091 | BIOSAFETY | Tools | Disinfectants and Cleaning | Paper towel box | Wall mounted, refillable box for a capacity of at least 150 paper towels |  |  |  | Always needed | 1 | 4
106092 | BIOSAFETY | Tools | Disinfectants and Cleaning | Goggles | Goggles, according to EN 166 and EN 170, frame adjustable (at least three positions, non-skid, integrated side and top protection, total weight <35 g , UV Protection |  |  |  | Regular needed | 1 | 5
106217 | MOLECULAR BIOLOGY | Tools | Tubes | Racks with lid (plastic preferred), suitable for cryovial | Racks with lid (plastic preferred), suitable for cryovial (2ml, round outer winding), Operating range -196 °C to +121 °C, autoclavable at 121 °C (20 min). |  | Yes | 106344 /106381 / 106382 | Always needed | 1 | 50
106246 | MOLECULAR BIOLOGY | Tools | Tubes | Racks, desinfectable / autoclavable for 1.5/2 ml tubes | Racks, desinfectable / autoclavable for tubes 1.5 - 2 ml, 10 mm diameter, at least 20 positions |  | Yes | 106381 / 106382 | Always needed | 1 | 10
106270 | MOLECULAR BIOLOGY | Tools | Pipetting | Racks for 1 channel pipette - 4 positions | Racks suitable for pipettes offered, 4 positions, disinfectable with bleach and alcoohol at least |  |  |  | Regular needed | 1 | 4
106410 | MOLECULAR BIOLOGY | Tools | Tubes | Racks for cryovials for One hand operation |  |  | Yes | 106383 | Always needed | 1 | 5
106413 | MOLECULAR BIOLOGY | Tools | Tubes | Rack for 50 ml PP-tubes | Rack for PP-tubes for centrifuge 50 ml, stable, |  | Yes | 106340 / 106341 | Always needed | 1 | 10
106433 | MOLECULAR BIOLOGY | Tools | Tubes | Rack for 15 ml PP-tubes -universal test tube rack 17mm | Universal test tube rack, PP, temperature resistant form -15 °C to +130 °C, autoclavable, for 17 mm diameter tubes, dimensions (W x D x H) 250 mm x 105 mm x 70 mm, 60 positions. (not swimming) or comparable |  | Yes | 106342 | Regular needed | 1 | 12
106442 | BIOSAFETY | Tools | Meters | Deep-freeze thermometer | With bimetallic indicator. Range -50 to +50 °C. Graduation 1 °C. Round plastic housing, e.g. Carl roth H717.1 or similar |  | Yes |  | Always needed | 1 | 3
106453 | MOLECULAR BIOLOGY | Tools | Laboratory, Miscellaneous | Thermometer for Oven | Minimum temp : 40°C |  | Yes |  | Regular needed | 1 | 2
106454 | MOLECULAR BIOLOGY | Tools | Tubes | Rack 15 ml / 50 ml - 4-Way Microtube Rack | Rack four different usable sides, high chemical resistant, high temperature resistant and autoclavable |  | Yes |  | Regular needed | 1 | 5
106455 | MOLECULAR BIOLOGY | Tools | Tubes | Rack Heat resistant for Cryo vial 1.5ml - Modular Heating Blocks | Rack Heat resistant for Cryo vial 1.5ml - Modular Heating Blocks * close contact of tubes to block walls allows for maximum heat retention * Each block should have a thermometer well for measuring block temperature |  | Yes | 106350 | Regular needed | 1 | 2
106491 | BIOSAFETY | Instrument | Laboratory, Miscellaneous | Battery Pack for UPS 3000 VA | Battery pack for 120 minutes at 500W | Yes |  |  | Regular needed | 1 | 1
106492 | BIOSAFETY | Instrument | Laboratory, Miscellaneous | Battery Pack for UPS 700VA | Battery pack for 120 minutes at 500W | Yes |  |  | Regular needed | 1 | 1

## 5 - Standardization Items
Item Number | Method | Family | Sub Family | Item Name | GDF SPECIFICATION | Unit / package GDF | Link to | Importance | Standardization | Quantity to equip 1 lab | India Local Procurement | Supplier Specifications | Primary Packing Type | Number of units per primary packing type | Secondary Packing Type | Number of units per secondary packing type | Gross weight in kg | Depth (in cm) | Width (in cm) | Height (in cm) | Volume in M3 | Special requirement for delivery | Special requirement for storage | Country of Origin | Custom Tariff code | Manufacturer Name | Item Number (Manufacturer) | Internet Link
106220 | PROCESSING | Instrument | Mixers, Shakers, Stirrers | Overhead stirrer + Universal adapter | Overhead stirrer, capacity for at least ten tubes of 50 ml volume | 1 |  | Optional | YES | 2 | [15 empty] | Heidolph Reax 2 Overhead Shakers | 541-21001-00 | http://www.heidolph-instruments.com/products/shakers-mixers/overhead-shakers/
106122 | PROCESSING | Instrument | Mixers, Shakers, Stirrers | Compact Orbital shaker (for BSC) |  | 1 |  | Regular needed | YES | 1 | [15 empty] | IKA | IKA MS3 Basic | http://www.ika.com/ika/datasheet/datasheet.php?iProduct=3617000&iLang=1&iCountry=756&filename=DataSheet_MS_3_basic.pdf
106247 | SOLID CULTURE | Instrument | Incubators | Precision Incubator | Precision Incubator, calibrated 37 °C, temperature range 30 - 70 °C, capacity ≥ 700 litre Inside and Oustide Stainless steel including 4 Perforated stainless steel | 1 |  | Regular needed | YES | 2 | [15 empty] | MEMMERT | INE 800 | http://www.memmert.com/fileadmin/template/memmert/pdfdata/INE%20800_EN.pdf
106248 | SOLID CULTURE | Instrument | Incubators | Perforated stainless steel | Perforated stainless steel shelf for Precision Incubator - 1 piece | 1 |  | Optional | YES | 2 | [15 empty] | MEMMERT | 0
106048 | MOLECULAR BIOLOGY | Instrument | Centrifuges | Microliter Centrifuge | Centrifuge for standard 1.5 mL reaction tubes, at least 15000 g, areosol tight rotors | 1 |  | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | HETTICH | MIKRO 200 | http://www.hettichlab.com/appc/content_manager/page.php?ID=160032&dbc=2d7mcrm18lltrf7bktb0rd2lj2
106257 | MOLECULAR BIOLOGY | Instrument | Laboratory, Miscellaneous | Universal oven 32 Liters | natural convection, w x h x d: 400 x 320 x 250 mm, 32 l, from +30 °C (however, at least 5 °C above ambient) up to +220 °C | 1 |  | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | MEMMERT | UNB200 | http://www.memmert.com/fileadmin/templates/memmert/pdfdata/UNB%20200_EN.pdf
106046 | MOLECULAR BIOLOGY | Instrument | Laboratory, Miscellaneous | Thermocycler, 96 positions for 0.2 ml tubes | Thermocycler, 96 positions for 0.2 ml tubes (e.g. AB 2720 Thermal Cycler) Static temperature uniformity • ±0.5ºC, 30 seconds after clock-start at 95ºC. Temperature accuracy • ±0.25ºC (range: 35–100ºC) Temperature calibration • Calibrated to standards traceable to the National Institute of Standards and Technology (NIST). Heated cover • Maintains constant temperature of 105ºC for oil-free operation. Ramp time reproducibility • Reaches thermal set points within ±5 seconds. | 1 |  | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | Applied Biosytems | 2720ABI | http://www3.appliedbiosystems.com/cms/groups/mcb_marketing/documents/generaldocuments/cms_040523.pdf
106063 | MOLECULAR BIOLOGY | Instrument | Meters | Flexible thermometer for thermocycler and Twincubator validation: (e.g. GTH 1170 with GTF300) | e.g. : http://www.greisinger.de/index.php?task=2&wg=153&artikel=633#110040 | 1 |  | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | Grisinger | GTH1170 with GTF300 | http://www.greisinger.de/files/upload/en/produkte/kat/8.pdf
106224 | MOLECULAR BIOLOGY | Instrument | Centrifuges | Refrigerated bench top centrifuge | Refrigerated bench top centrifuge, for swing out rotor with a capacity for 1000 ml liquid to be centrifuged at > 3200 g | 1 |  | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | HETTICH | 380R | http://www.hettichcentrifuge.co.uk/pdf/rotina380.pdf
106225 | MOLECULAR BIOLOGY | Instrument | Centrifuges | 4-place swing out rotor | 4-place swing out rotor for bench top centrifuge | 1 | 106224 | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | HETTICH | 1754
106227 | MOLECULAR BIOLOGY | Instrument | Centrifuges | Aerosol tight lids for rotor buckets | Aerosol tight lids for rotor buckets | set of 4 | 106048 | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | HETTICH | 1751
106229 | MOLECULAR BIOLOGY | Instrument | Centrifuges | Reducing adapter for 15 ml/12 ml tubes | Reducing adapter for 15 ml/12 ml Falcon tubes or equivalent tubes | set of 4 | 106224 | Regular needed | YES | 1 | [15 empty] | HETTICH | 1771
106228 | MOLECULAR BIOLOGY | Instrument | Centrifuges | Reducing adapter for 50 ml tubes | Reducing adapter for 50 ml Falcon tubes or equivalent tubes | set of 4 | 106224 | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | HETTICH | 1772
106226 | MOLECULAR BIOLOGY | Instrument | Centrifuges | Rotor buckets | Rotor buckets | set of 4 | 106224 | Always needed | YES | 1 | To be available via India local procurment | [14 empty] | HETTICH | 1752
106061 | MOLECULAR BIOLOGY | Tools | Pipetting | Hand-dispenser pipette, 5 dosage steps - LPA | Hand-dispenser pipette, 5 dosage steps, suited for 5ml / 12.5ml tips | 1 |  | Optional | YES | 1 | [15 empty] | BRAND | 612-5530 | http://www.brand.de/fileadmin/user/pdf/Leaflets/HandyStep_S_EN.pdf
106425 | MOLECULAR BIOLOGY | Instrument | Incubators | Thermoblock / Thermo heater | Thermo heater, exchangeable heating blocks for different standard reaction tubes, range + 5 ºC above ambient to >100 ºC, heating time 25 ºC to 100 ºC ≤ 20 min., uniformity within the block +/- 0,2°C (at 37°C) | 1 |  | Optional | YES | 1 | [15 empty] | PEQLAB | 91-D1100, 91-D115B, 91-D120B | http://www.peqlab.co.uk/wcms/uk/pdf/91-D1100.pdf
106115 | LIQUID CULTURE | Tools | Pipetting | Hand-dispenser pipette, 5 dosage steps, suited for 10ml Tips / Eppendorf | Hand-dispenser pipette, 5 dosage steps, suited for combitips - e.g. Eppendorf Hand-dispenser pipette, Multipipette ® Plus or similar to fit with 10 ml tips | 1 |  | Regular needed | YES | 2 | To be available via India local procurment | [14 empty] | EPPENDORF | Multipipette ® Plus | http://www.eppendorf.com/int/index.php?sitemap=2.3&pb=a5d0f5d8bf3bd6ac&action=products&contentid=1&catalognode=9620&productpage=1
106180 | LIQUID CULTURE | Tools | Pipetting | Multipette® stream, includes charging adapter | Multipette® stream, includes charging adapter - for MGIT Supplementation 0.8ml each - (e.g Eppendorf Multipette® stream, includes charging adapter) | 1 |  | Regular needed | YES | 1 | [15 empty] | EPPENDORF | Multi pipette ®stream | http://www.eppendorf.com/int/index.php?sitemap=2.3&pb=a5d0f5d8bf3bd6ac&action=products&contentid=1&catalognode=22350&productpage=1
106079 | BIOSAFETY | Instrument | Meters | Thermic anemometer with telescope probe | Thermic anemometer with telescope probe | 1 | 106076 | Optional | YES | 1 | [15 empty] | Airflow Instruments | Airflow TA 5
106099 | CLEANING KITCHEN | Instrument | Sterilization | Universal Hot-air oven - Memmert UFE 600 | Universal Hot-air oven, temperature range 10 °C above ambient to 250 °C, approx.110 L or 250 L (depending on the amount of reused items and washed glassware) with 2 Perforated stainless steel shelf for 254 L Hot-air oven | 1 |  | Optional | YES | 1 | [15 empty] | MEMMERT | UFE 600 | http://www.memmert.com/fileadmin/templates/memmert/pdfdata/UFE%20600_EN.pdf
106100 | CLEANING KITCHEN | Instrument | Sterilization | Perforated stainless steel shelf for Universal Hot-air oven |  | 1 | 106099 | Optional | YES | 1 | [15 empty] | MEMMERT | UFE 600 E3
106131 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Water bath 22 Liters | Water bath, capacity approximately 20 L with microprocessor- controlled temperature regulator, including S/S cover for water bath . Accuracy of water temperature : +/- 0.5°C | 1 |  | Regular needed | YES | 1 | To be available via India local procurment | [14 empty] | MEMMERT | WNB22 | http://www.memmert.com/fileadmin/templates/memmert/pdfdata/WNB%2022_EN.pdf
106510 | MEDIA KITCHEN | Instrument | Laboratory, Miscellaneous | Water bath 7 Liters | Water bath, capacity approximately 7 L with microprocessor- controlled temperature regulator, including S/S cover for water bath - Accuracy of water temperature : +/- 0.5°C | 1 |  | Optional | YES | 1 | To be available via India local procurment | [14 empty] | MEMMERT | WNB7
106369 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Masks 3M 9320 | FFP2 (N95) or FFP3 (N99) Masks (according to national policy), individually packed, foldable mask | 10 |  | Always needed | YES | 5000 packs per year | To be available via India local procurment | [14 empty] | 3M | 9320

## 6 - Autoclave lot
Item Number | Method | Family | Sub Family | Item Name | GDF SPECIFICATION | Unit / package GDF | Link to | Importance | Detailled Specification (Y/N) | Supplier Specifications | Gross weight in kg | Depth (in cm) or sizr | Width (in cm) | Height (in cm) | Volume in M3 | Average Delivery time in week(s) | Manufacturer Name | Model or Item Number (Manufacturer) | Country of Origin | Custom Tariff code | Voltage / Amp / Frequency | Please mention quality standard under which the item is produced (if any)
106116 | MEDIA KITCHEN | Instrument | Sterilization | Autoclave basic unit for Media Kitchen | Autoclave basic unit with Program control unit - Liquid temperature control - Liquid Sterilization and recooling system, vertical, capacity approx. 75 litres | 1 |  | Regular Needed | Yes
106117 | WASTE MANAGEMENT | Instrument | Sterilization | Autoclave basic unit for Waste Management |  | 1 |  | Regular Needed | Yes
106118 | WASTE MANAGEMENT / MEDIA KITCHEN | Instrument | Sterilization | Exhaust air filtration system for Autoclave | Exhaust air filtration system FA for autoclave BSL3-Lab (waste handling) and vacuum pump | 1 |  | Regular Needed | Yes
106119 | WASTE MANAGEMENT / MEDIA KITCHEN | Instrument | Sterilization | Autoclave stainless steel buckets | - Three stainless steel buckets with lid, diameter and height adjusted that two fit into the autoclave at the same time (usually 380 x 290 mm) - Two stainless steel buckets with lid, diameter and height adjusted so that one fits into the autoclave | 1 kit |  | Regular Needed | Yes
106120 | WASTE MANAGEMENT / MEDIA KITCHEN | Instrument | Sterilization | Autoclave stainless steel baskets | -Three Stainless steel wire baskets, diameter and height adjusted that two fit into the autoclave at the same time. - Two stainless steel wire baskets, diameter and height adjusted that one fits into the autoclave | 1 kit |  | Regular Needed | Yes
106121 | WASTE MANAGEMENT / MEDIA KITCHEN | Instrument | Sterilization | Printer for Autoclave |  | 1 |  | Regular Needed | Yes
106490 | WASTE MANAGEMENT | Instrument | Sterilization | Installation and Training - Autoclave | Installation of autoclave in the laboratory with if needed addtional pre installation work as plumbing, electricity…. Training on use and troubleshooting of autoclave | 1 |  | Regular needed |  |  | n.a | n.a | n.a | n.a | n.a |  | n.a | n.a | n.a | n.a | n.a | n.a

## 7 - BSC Lot
Item Number | Method | Family | Sub Family | Item Name | GDF SPECIFICATION | Unit / package GDF | Importance | Detailled Specification (Y/N) | Supplier Specifications | Gross weight in kg | Depth (in cm) | Width (in cm) | Height (in cm) | Volume in M3 | Warranty | Average Delivery time in week(s) | Manufacturer Name | Model or Item Number (Manufacturer) | Country of Origin | Custom Tariff code | Voltage / Amp / Frequency | Please mention quality standard under which the item is produced (if any)
106076 | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Bio Safety Cabinet - 3 filters, 5 feet | class II A2 | 1 | Regular needed | Yes
New item | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Bio Safety Cabinet - 3 filters, 4 feet | class II A2 | 1 | Regular needed | Yes
New item | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Bio Safety Cabinet - 3 filters, 6 feet | class II A2 | 1 | Regular needed | Yes
New item | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Bio Safety Cabinet - 2 filters, 5 feet | class II A2 - To be used with an additional Exhaust filter to re - circulate air in the room if needed | 1 | Regular needed | Yes
New item | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Bio Safety Cabinet - 2 filters, 4 feet | class II A2 - To be used with an additional Exhaust filter to re - circulate air in the room if needed | 1 | Regular needed | Yes
New item | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Bio Safety Cabinet - 2 filters, 6 feet | class II A2 - To be used with an additional Exhaust filter to re - circulate air in the room if needed | 1 | Regular needed | Yes
106077 | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Base Support for BSC - 5 feet | Base Support for Biosafety Cabinet (for sitting working position) | 1 | Regular needed | Yes
New item | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Base Support for BSC - 4 feet | Base Support for Biosafety Cabinet (for sitting working position) | 1 | Regular needed | Yes
New item | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Base Support for BSC - 6 feet | Base Support for Biosafety Cabinet (for sitting working position) | 1 | Regular needed | Yes
106078 | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Thimble for Class IIA, hood for BSC | Thimble for class IIA, hood for exhaust air connection to the ventilation system, one suited for each BSC. Air outlet should either vertical or Horizontal. With Anti blow back valve (if needed) | 1 | Regular needed | Yes
106471 | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Fromaldehyde/neutralization fumigation top plate |  | 1 | Optional
106472 | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | HVP fumigation plate |  | 1 | Optional
106469 | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Double Exhaust filter | Frame and HEPA Filter | 1 | Optional
106285 | BIOSAFETY | Consumables | Hoods, Cabinets, Workstations | Exhaust HEPA Air Filter | HEPA filter, class H14 according to the European Norm EN 1822:2009, metal frame (preferred), filter media, glass fiber preferred, appropriate size for [Company, model, size additional information for precise choice] - Filter has to be packed in Sturdy Box | 1 | Regular needed
106047 | BIOSAFETY | Consumables | Hoods, Cabinets, Workstations | Main HEPA Filter for air down flow | HEPA filter, class H14 according to the European Norm EN 1822:2009, metal frame (preferred), filter media, glass fiber preferred, appropriate size for [Company, model, size additional information for precise choice] - Filter has to be packed in Sturdy Box | 1 | Regular needed
106123 | BIOSAFETY | Consumables | Hoods, Cabinets, Workstations | Pre HEPA filter for down flow and in take air | HEPA filter, class H14 according to the European Norm EN 1822:2009, metal frame (preferred), filter media, glass fiber preferred, appropriate size for [Company, model, size additional information for precise choice] - Filter has to be packed in Sturdy Box | 1 | Regular needed
106489 | BIOSAFETY | Instrument | Hoods, Cabinets, Workstations | Installation and Training - BSC | Installation of BSC in the laboratory. If necessary duct the BSC outside with a Thimble construction Certification of BSC after installation Training on use and troubleshooting of BSC and if necessary on certification | 1 | Regular needed |  |  | n.a | n.a | n.a | n.a | n.a | n.a |  | n.a | n.a

## Evaluation Method for sample
Item Number | METHOD | Family | Sub Family | ITEM NAME | GDF specification | Unit / package GDF | Specific compliance test | Sample to provide
106375 | BIOSAFETY | Consumables | Disinfectants and Cleaning | Single-use paper towels | Single-use paper towels, approx. 230 x 300 mm folded to take as single pieces off a wall mounted box. 150 towels/pack, 30 packs/ carton | 4500 | Ability to suck water (eg. How much to dry hand : 2 pieces maximum )- Box must fit with size of towels | 1 pack
106345 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size S | powder-free disposable latex gloves, size S, >0,16 mm thickness, length >240 mm, 1000 per pack | 1000 | Thickness - lenght - feeling with the hand to work - how it fit with different person - sensitivity to work - Grip of the gloves when it is wet | 1 pack
106346 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size M | powder-free disposable latex gloves, size M, >0,16 mm thickness,length >240 mm, 1000 per pack | 1000 |  | 1 pack
106347 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Latex gloves size L | powder-free disposable latex gloves, size L, >0,16 mm thickness, length >240 mm, 1000 per pack | 1000 |  | 1 pack
106403 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 6-7 | nitril gloves, size 6-7, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) | 12 |  | 1 pack
106404 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 7-8 | nitril gloves, size 7-8, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) | 12 |  | 1 pack
106405 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 8-9 | nitril gloves, size 8-9, CE category III, against chemical risks (sulphuric acid, ethidiumbromide, auramine staining etc. thickness >0.36 mm) | 12 |  | 1 pack
106497 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 6-7 - Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. | 100 |  | 1 pack
106498 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 7-8 -Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. | 100 |  | 1 pack
106499 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | nitril gloves, size 8-9 - Disposable | Powder-free, latex-free disposable gloves; rated as CE category III or equivalent (against chemical risks and resistant to ethidium bromide or auramine staining); . 0.3 mm thick and . 249 mm long; stocked in sizes (small, medium and large) appropriate for the laboratory staff. | 100 |  | 1 pack
106348 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Transparent polypropylene bag | Transparent polypropylene bag for waste bins, at least 40 µm thickness, up to 134 °C steam sterilisable, approx 24 L, approx. 540 x 550 mm, in accordance with countiy's norms for environmental and hygiene assignments, 1000 pieces per pack | 1000 | how stable it is with liquid, thight - bleach resistant - sturdy e.g with soft pressure it should resist. | 1 pack
106444 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Plastic bags - 30L | Red color, autoclavable with Bio Hazardous Logo | 200 |  | 1 pack
106399 | BIOSAFETY | Consumables | Plastic & Glassware | Plastic bags made from PP | small disposable plastic bags made from PP, at least 40 µm thick for tripod (approx. 200 mm width and 300 mm length), liquid tight, for not infections material (PCR, material from test kits), 100 pieces per pack (or roles with 100 pieces, perforated for teare off) | 100 |  | 1 pack
106349 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Tri-pod stand for waste bags of ca. 2 litres | Stands, 3 foot, for small plastic bags (size ca. 2 litre) | 1 | stability with liquid / waste… - resistant to desinfectant in use. | 1 pack
106445 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Stand for 30L plastic bags | Stable, should stay in a fix position especially when bag is filled | 1 |  | 1 pack
106355 | PROCESSING | Consumables | Laboratory, Miscellaneous | Marker pen | Marker pen, resistant against water and alcohol, also against autoclaving, writing on nearly all surfaces, even when cold and wet, fast drying, black, line 0.8 - 1 mm broad, black | 1 | water proof - writing should stay under all condition on the tube | 1 pack
106356 | PROCESSING | Consumables | Laboratory, Miscellaneous | Adhesive labels | Adhesive labels on rolls or sheets (e.g. 50 x 19 mm, 24 per sheet and 200 sheets per box = 4800 label) | 1 | should be sticky to all temperature - test in a deep freezer or oven | 1 pack
106357 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Cryo-tags | Cryo-tags sized to fit for use on cryo-tubes and standard micro-tubes (approx.32 mm x 12 mm), made from smear resistant material, but should accept all marking devices, rolles of 1000 | 1 |  | 1 pack
106360 | CLEANING KITCHEN | Consumables | Laboratory, Miscellaneous | Aluminium foil | Aluminium foil, 30 µm thickness, 600 mm width, 100 m long | 1 | Compliance | 1 pack
106363 | CLEANING KITCHEN | Consumables | Laboratory, Miscellaneous | Tube brush | Tube brush, total length min.280 mm, diameter 15-16 mm, brush length 80 mm - too fit witt culture tubes. Soft at the end to not brake the tube | 5 | easy to use with culture tube. | 1 pack
106366 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size L - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent | 100 | Flame resistant/retardant : should be made of non-melting materials to maximize thermal performance and minimize the potential for burn injury - Resistant to bleach - antistatic - easy to work with - water-repellent | at least 20 units
106367 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size M - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent | 100 |  | at least 20 units
106368 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory coat size S - Disposable - Sterile | Laboratory coat size L - Disposable - Sterile - with elastic wrist - water-repellent | 100 |  | at least 20 units
106435 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size S | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent | 20 |  | 1 pack
106436 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size M | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent | 20 |  | 1 pack
106437 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gowns non sterile - Size L | Disposable - with wrist - antistatic - resistant to fire and disenfectant (bleach, alcool, phenol…) -Overlap at the back for any body - water-repellent | 20 |  | 1 pack
106500 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - S | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body | 1 |  | 1 pack
106501 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - M | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body | 1 |  | 1 pack
106502 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Surgical gown - Cotton - L | Surgical gown - back opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning - Overlap at the back for any body | 1 |  | 1 pack
106503 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - S | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning | 1 |  | 1 pack
106504 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - M | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning | 1 |  | 1 pack
106505 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Laboratory Coat cotton - L | Laboratory coat - front opening - Cotton - with elastic wrist or equivalent Need to be autoclavable - Not easily burning | 1 |  | 1 pack
106440 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Shoe cover | Plastic, water resistant | 1000 |  | 1 pack
106441 | BIOSAFETY | Consumables | Laboratory, Miscellaneous | Hair Cover | 0 | 2000 |  | 1 pack
106181 | LIQUID CULTURE | Consumables | Pipetting | Dispenser-tips Universal 10ml sterile | Dispenser-tips Universal 10ml sterile - DNA- /RNase- /ATP- and endotoxin free - suitable for pipettes offered | 100 | Accuracy and reproductibility according to manufactures norm and ISO 8655-5 | 1 pack
106395 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 5 ml - bag of 10 | Tips for Multipette/Combi 100 pieces/pack - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered | 100 |  | 1 pack
106396 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 5 ml, individually packed | Tips for Multipette/Combi 100 pieces/pack - sterile, DNA- /RNase- /ATP- and endotoxin- free -suitable for pipettes offered | 100 |  | 1 pack
106397 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 12.5ml- bag of 10 | Tips for Multipette/Combi 12.5 ml - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered | 100 |  | 1 pack
106398 | MOLECULAR BIOLOGY | Consumables | Pipetting | Tips for Multipette/Combi 12.5 ml - individually packed | Tips 12.5 ml - sterile, DNA- /RNase- /ATP- and endotoxin- free - suitable for pipettes offered | 100 |  | 1 pack
106386 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 0.1 - 10 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 0.1 - 10 µl, | 480 | Compliance with spec - Accuracy - Reproducibility ‘The maximum permissible errors always apply to the total system of piston pipette and tip. Test will be done for both pipettes and tips. The chosen tip should fit tightly onto the pipette, ensuring accurate sealing with minimal force and without needing to bang the pipette into the tip | at least 100 units
106387 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 1.0 - 20 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 1.0 - 20 µl, | 480 |  | at least 100 units
106388 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 10 - 100 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered, 10 - 100 µl, | 960 |  | at least 100 units
106389 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS, 20 - 200 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 20 - 200 µl, | 960 |  | at least 100 units
106390 | MOLECULAR BIOLOGY | Consumables | Pipetting | Sterile, DNA-/RNAse-free TIPS 100 - 1000 µl | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 100 - 1000 µl, | 960 |  | at least 100 units
106419 | LIQUID CULTURE | Consumables | Pipetting | TIPS, PP, 20-200 μl, sterile, autoclavable with Filter | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 20 - 200 µl, | 960 |  | at least 100 units
106420 | LIQUID CULTURE | Consumables | Pipetting | TIPS, PP, 100-1000 μl, sterile, autoclavable with Filter | Pipette tips, high-quality, transparent (PP), with aerosol (PE) filter, sterile, autoclavable, certified to be free of DNase, RNase, DNA, pyrogens, metals and PCR-inhibitors. Inside area prepared for minimal specimen adhesion. compatible for 1-channel pipette offered 100 - 1000 µl, | 960 |  | at least 100 units
106053 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 0.5 - 10 µl - LPA | 1-channel pipette, variable from 1 - 10 µl, - accuracy < 0.025 to 0.100, precision < 0.012 to 0.040 | 1 |  | 1 set of all pipettes
106054 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 2,0 - 20 µl - LPA | 1-channel pipette, variable from 2,0 - 20 µ, - accuracy < 0.10 to 0.20, precision < 0.03 to 0.06 | 1
106055 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 10 - 100 µl - LPA | 1-channel pipette, variable from 10 - 100 µl | 1
106056 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-chanel pipette, variable from 20 -200 ul - LPA | 1-chanel pipette, variable from 20 -200 ul - accuracy < 0.50 to 1.60, precision < 0.20 to 0.30 | 1
106057 | MOLECULAR BIOLOGY | Tools | Pipetting | 1-channel pipette, variable from 200 - 1000 µl - LPA | 1-channel pipette, variable from 100 - 1000 µl - accuracy < 3.00 to 8.00, precision < 0.60 to 1.50 | 1
106072 | LIQUID CULTURE | Tools | Pipetting | 1-chanel pipette, variable from 20 -200 ul - LC | 1-chanel pipette, variable from 20 -200 ul - accuracy < 0.50 to 1.60, precision < 0.20 to 0.30 | 1 |  | No sample if same brand than for Moleculart Biology
106075 | LIQUID CULTURE | Tools | Pipetting | 1-channel pipette, variable from 200 - 1000 µl - LC | 1-channel pipette, variable from 100 - 1000 µl - accuracy < 3.00 to 8.00, precision < 0.60 to 1.50 | 1
106394 | LIQUID CULTURE | Consumables | Pipetting | Single use plastic Pasteur-pipettes sterile individually packed | sterile, single use plastic Pasteur-pipettes volume in a range of 1 to 3 ml, individually packed, graduated 0.5ml, lenght : minimum 155mm ideal 170mm. Drop size maximum 30ul | 500 | Accuracy of marking - Reproductibily - unbreakable | at least 100 units
106343 | LIQUID CULTURE | Consumables | Pipetting | Single use plastic Pasteur-pipettes sterile pack by 10 | sterile, single use plastic Pasteur-pipettes volume in a range of 1 to 3 ml, packed by 10, graduated 0.5ml, lenght : minimum 155mm ideal 170mm. Drop size maximum 30ul | 500 |  | at least 100 units
106434 | MOLECULAR BIOLOGY | Consumables | Pipetting | Disposable pasteur pipettes, graduated, non sterile, 155 mm, 3 ml | Unsterile, plastic Pasteur-pipettes, Length 155 mm, precision-formed, unbreakable, non toxic and with controlled, reproducible drop size (e.g : Carl roth EA65.1 or similar) | 500 |  | at least 100 units
106338 | SOLID CULTURE | Consumables | Tubes | Culture tubes | Culture tubes, diameter 16 mm, length 125 mm, wall strength 1mm, thread outside; inner diameter min.13 mm; borosilicate glass with ISO winding, autoclavable screw caps with gasket out of PP or Teflon. Some prefer 18 mm diameter | 100 | Compliance with specificaton | at least 100 units
106340 | MOLECULAR BIOLOGY | Consumables | Tubes | PP-tubes for centrifuge non sterile 50 ml | PP-tubes for centrifuge, 50 ml with screw cap, tight, size 115x30 mm (Falcon or equivalent), 500 pieces/pack - Resistance minimum 3,200 g in a swing out bucket | 500 |  | at least 100 units
106341 | LIQUID CULTURE | Consumables | Tubes | PP-tubes for centrifuge, sterile, 50 ml | PP-tubes for centrifuge, 50 ml with screw cap, tight, size 115x30 mm (Falcon or equivalent), sterile, packed 25 in racks, 20 racks per pack. | 500 |  | at least 100 units
106342 | LIQUID CULTURE | Consumables | Tubes | PP-tubes for centrifuge, sterile, 15 ml | PP-tubes for centrifuge, 15 ml with screw cap, tight, size 120x17 mm, 500 pieces/pack - Resistance minimum 3,200 g in a swing out bucket | 500 |  | at least 100 units
106384 | MOLECULAR BIOLOGY | Consumables | Tubes | PP-tubes for centrifuge, non sterile, 15 ml | PP-tubes for centrifuge, 15 ml with screw cap, tight, size 120x17 mm, 500 pieces/pack | 500 |  | at least 100 units
106350 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, 1.5 ml | Cryo-vial, sterile with standard screw cap, 1.5 ml, PP with a medical-grade silicone O-ring seal, style conical, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). It can be centrifuged up to 17 000 x g - e.g. SIGMA V4131 | 1000 |  | at least 100 units
106344 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, 2 ml | Cryo-vial, sterile with standard screw cap, 2 ml round with outer winding, PP, graduated, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). pack of 1000 | 1000 |  | at least 100 units
106383 | MOLECULAR BIOLOGY | Consumables | Tubes | Cryo-vial, sterile with cap, for one hand operation | Cryo-vial, sterile with Tethered Screw Caps, 2 ml round with outer winding PP, graduated, frosted marking area and coloured cap inserts for easy sample identification. Temperature stability to -196 °C. γ-ray sterile (SAL 10-6) and autoclavable at 121 °C (20 min.). (sterile, RNase-, DNase-, DNA and endotoxin-free). pack of 1000 | 1000 |  | at least 100 units
106385 | MOLECULAR BIOLOGY | Consumables | Tubes | PCR tubes | PCR tubes, the tubes should be extremely thin to ensures speedy and even heat transfer of every sample, pyrogen free, DNase- /RNase-free and autoclavable, 0,2 ml with attached caps, fits all standard thermocyclers with heated lids especially the one provided. | 1000 |  | at least 100 units
106421 | MOLECULAR BIOLOGY | Consumables | Tubes | Standard reaction tubes 1.5 ml | Standard reaction tubes 1.5 ml, conical with screw top (twist top vials) with O-ring (seal ring) with Uniform lid thickness, PP, autoclavable at 121 °C (20 min.), High clarity, Frosted marking area, subdivisions for approx. volume determination Can withstand centrifugation up to 20 000 g | 500 |  | at least 100 units
106422 | MOLECULAR BIOLOGY | Consumables | Tubes | Standard reaction tubes 2 ml | Standard reaction tubes 2 ml, conical with screw top (twist top vials) with O-ring (seal ring) with Uniform lid thickness, PP, autoclavable at 121 °C (20 min.), High clarity, Frosted marking area, subdivisions for approx. volume determination Can withstand centrifugation up to 20 000 g | 500 |  | at least 100 units
106230 | LIQUID CULTURE | Tools | Centrifuges | Test tube rack 18/8 stainless steel 16 mn | diameter 16 mm, 3x12 positions, sturdy | 1 | Compliance with specification Sturdy / Stable Compatible and easy to use with tubes and cap provided | 1 pack
106231 | PROCESSING | Tools | Centrifuges | Test tube rack 18/8 stainless steel 35 mn | diameter 35 mm, 2x4 positions | 1 |  | 1 pack
106068 | MOLECULAR BIOLOGY | Tools | Tubes | Stand (96) for 0.2 ml pcr tubes / PCR Rack | Stand (96) for 0.2 ml, plastic resistant to bleach and alcoohol at least | set of 10 |  | 1 pack
106496 |  |  |  | Slanted racks fr BD LJ Tubes | Made of PC, autoclavable, for simple and speedy preparation of angular agar with two different inclination angles. ( 5° and 20°) With 40 alphanumerically marked holes (4x10). - Tube diameter between 20 and 22mm Compatible with 221413 - L-J Medium Slants (10/sp) |  |  | 1 pack
106217 | MOLECULAR BIOLOGY | Tools | Tubes | Racks with lid (plastic preferred), suitable for cryovial | Racks with lid (plastic preferred), suitable for cryovial (2ml, round outer winding), Operating range -196 °C to +121 °C, autoclavable at 121 °C (20 min). | 1 |  | 1 pack
106237 | PROCESSING | Tools | Tubes | Universal test tube rack 16mm | Universal test tube rack, PP, temperature resistant form -15 °C to +130 °C, autoclavable, for 16 mm diameter tubes, dimensions (W x D x H) 250 mm x 105 mm x 70 mm, 60 positions. (not swimming) | 1 |  | 1 pack
106238 | PROCESSING | Tools | Tubes | Universal test tube rack 18mm | Test tube rack PP, diameter 18mm, 4x12, sturdy | 1 |  | 1 pack
106246 | MOLECULAR BIOLOGY | Tools | Tubes | Racks, desinfectable / autoclavable for 1.5/2 ml tubes | Racks, desinfectable / autoclavable for tubes 1.5 - 2 ml, 10 mm diameter, at least 20 positions | 1 |  | 1 pack
106410 | MOLECULAR BIOLOGY | Tools | Tubes | Racks for cryovials for One hand operation | 0 | 1 |  | 1 pack
106413 | MOLECULAR BIOLOGY | Tools | Tubes | Rack for 50 ml PP-tubes | Rack for PP-tubes for centrifuge 50 ml, stable, | 1 |  | 1 pack
106432 | SOLID CULTURE | Tools | Tubes | Slanted racks | Made of PC, autoclavable, for simple and speedy preparation of angular agar with two different inclination angles. ( 5° and 20°) With 40 alphanumerically marked holes (4x10). Dimensions: L 285 x W 115 mm. (e.g. carl roth item N°E100.1) | 1 |  | 1 pack
106433 | MOLECULAR BIOLOGY | Tools | Tubes | Rack for 15 ml PP-tubes -universal test tube rack 17mm | Universal test tube rack, PP, temperature resistant form -15 °C to +130 °C, autoclavable, for 17 mm diameter tubes, dimensions (W x D x H) 250 mm x 105 mm x 70 mm, 60 positions. (not swimming) or comparable | 1 |  | 1 pack
106454 | MOLECULAR BIOLOGY | Tools | Tubes | Rack 15 ml / 50 ml - 4-Way Microtube Rack | Rack four different usable sides, high chemical resistant, high temperature resistant and autoclavable | 1 |  | 1 pack
106455 | MOLECULAR BIOLOGY | Tools | Tubes | Rack Heat resistant for Cryo vial 1.5ml - Modular Heating Blocks | Rack Heat resistant for Cryo vial 1.5ml - Modular Heating Blocks * close contact of tubes to block walls allows for maximum heat retention * Each block should have a thermometer well for measuring block temperature | 1 |  | 1 pack
106242 | PROCESSING | Tools | Sample Collection and Media Transport | Transport boxes for specimen | Transport boxes for specimen from collection points, styrofoam covered with PP, including cooling elements , number according to the needs | 1 | Disinfectable - about 30 liters volume -compatible for racks provided for 50 ml tubes | 1 pack
106243 | PROCESSING | Tools | Sample Collection and Media Transport | Cool boxes for sputum transport | Cool boxes for sputum transport by car, inner volume approx. 27 litre. Inner dimensions (W x D x H) ≥ 335 mm x ≥ 230 mm x ≥ 380 mm Inside and outside durable, disinfectable surface (e.g. LDPE), isolated with PU foam, Peltier elements, 12 V plug for cigarette lighter in a car cable length ≥ 2.5 m, possibility to change the polarity for cooling or heating alternatively. Cooling capacity 20 °C below ambient and heating 40°C above ambient. This transportation box is used in areas with extreme temperature deviations over the year and long duration of transport. | 1 | Compliance with specification | 1 pack
106254 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Molecular grade water | DNAase, RNAase free water - not chemically treated molecular grade water - 1.7mlx10 or 10ml | 10x1.7ml | Compliance with specification | 1 pack
106400 | MOLECULAR BIOLOGY | Consumables | Laboratory, Miscellaneous | Forceps, individually wrap, sterile | Forceps, preferably plastic for handling strips, individually wrap, sterile | 100 | Compliance with specification | 1 pack
106453 | MOLECULAR BIOLOGY | Tools | Laboratory, Miscellaneous | Thermometer for Oven | Minimum temp : 40°C | 1 | Accuracy at least +/- 1°C | 1 pack
106088 | BIOSAFETY | Tools | Meters | Electronic maxima-minima-thermometer | Electronic maxima-minima-thermometer, two channels, one for room temperature and the second for temperature inside the equipment (incubator/refrigerator), switchable Celsius/Fahrenheit, thin cable with sensor, battery, measuring range for room temperature -10°C/+50°C (-14°F/+!122°F) and for equipment from -50°C/+60°C (-58°F/+140°F), reset function, number depending on the sum of incubators refrigerators, freezers and deep freezers | 1 |  | 1 pack
106442 | BIOSAFETY | Tools | Meters | Deep-freeze thermometer | With bimetallic indicator. Range -50 to +50 °C. Graduation 1 °C. Round plastic housing, e.g. Carl roth H717.1 or similar | 1 |  | 1 pack
106358 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Sterile indicator tape, hot air oven | Sterile indicator tape (roles), approx.19mm broad and 55m long, DIN EN 867-5, for hot air sterilize (green/brown) | 1 | Compliance with specification | 1 pack
106359 | WASTE MANAGEMENT | Consumables | Laboratory, Miscellaneous | Sterile indicator tape, autoclave | Sterile indicator tape (roles), approx.19mm broad and 55m long, DIN EN 867-5, for autoclave (beige/dark brown) | 1 |  | 1 pack
106207 | MEDIA KITCHEN | Tools | Laboratory, Miscellaneous | Set of spatula | Set of spatula: stainless steel 18/8 electrolytically polished, length approx. 150 mm, blade approx. 40 mm in length and width ranging from 3 mm to 9 mm, 6 pieces in most common sizes of double and micro spatulas, 3 of them with spoon on one side | 1 |  | 1 pack
106506 |  |  |  | Marker pen PCR Tube | Permanent - to use with PCR Tubes |  |  | 1 pack
106509 | PROCESSING | Consumables | Laboratory, Miscellaneous | Special pH indicator strips | Range from pH 6.5 to pH 10; Reading in small increments, especially around pH 8 the incremental reading should be 0.2 pH units, according to a provided colour scale; Colour not “bleeding” from the strip | 100 | Test strips with PH meter | 1 pack