## Slide 1: GenoType® MTBDRplus
Specimen preparation

## Slide 2: Content outline
PCR laboratory layout
Checklist
Personal protective equipment (PPE)
Working in a Biological Safety Cabinet (BSC)
Specimen preparation for PCR
Contamination control

## Slide 3: Optimal PCR laboratory layout
DNA-“dirty” area
DNA-free
area
Amplicon-
“dirty” area

## Slide 4: Working in BSC (I)
Prepare a written checklist of materials and tools needed
Prepare a worksheet with list of specimens (identifiers) to be tested
Notes: <number>
Use of written checklists avoids the unnecessary interruption of work in the BSC that can compromise the quality of the specimen testing or the biosafety of laboratory staff.

## Slide 5: Use appropriate BSL3 PPE
Gowns
  Must have solid front and can be tied in the back
  Long-sleeved with elastic cuffs
  Must be either autoclavable/washable (soak in bleach) or disposable
Respirators
  N-95
Gloves
  Single-use

## Slide 6: BSL 3 PPE is suitable for:
Specimen processing
Specimen inoculation
Manipulating cultures
Centrifuging
Inoculation of DSTs
Making smears from cultures and processed specimens
Biochemical testing to identify TB
Preparing samples for PCR (before heat-kill)

## Slide 7: Working in BSC (2)
Decontaminate the BSC with freshly diluted 0.5% bleach (20 min), followed by 70% alcohol
Cover the working area with paper towels and spray it with suitable disinfectant
Use a discard container with a plastic bag and freshly diluted 1% bleach
DO NOT block the grill
Perform all operations at least 12 cm away from the front grill on the work surface
Place all materials and aerosol-generating equipment away from the front grill
Notes: <number>
Place absorbent paper towelling on the work surface to minimize the potential of exposure to infectious material or cross contamination in case of a spill.
Blocking the grill disrupts the integrity of the cabinet air barrier and can compromise both staff and product protection.
Keeping clean materials at least 12 cm away from aerosol-generating activities will minimise the potential for cross-contamination

## Slide 8: Preparation of clinical specimens
Label 1.5 ml screw-cap conical tubes (not the caps) for each specimen
Vortex processed clinical specimen for 1 min
Process only one specimen at a time; do not leave open containers or centrifuge tubes in the BSC
Transfer 500 l processed clinical specimen into 1.5 ml conical screw-capped vial; change tips!
Infectious material: all steps have to be performed in a BSC at BSL3 Laboratory
Clinical specimens have to be processed by N-acetyl-L-cysteine/NaOH method before specimen preparation

## Slide 9: Sample preparation of isolates on solid culture
Label 1.5 ml screw cap conical tubes (not the caps) for each sample
Add 300 µl molecular biology grade water to each tube
Re-suspend 1–2 colonies of M. tuberculosis culture on solid medium; change loops!
Do not scrape the surface of the medium since components of egg-based medium can inhibit PCR reaction
Vortex suspension for 1 minute to break up clumps
Infectious material: all steps have to be performed in a BSC at BSL3 Laboratory

## Slide 10: Sample preparation of liquid cultures
Label 1.5 ml screw-cap conical tubes (not the caps) for each sample
Vortex liquid culture tube to homogenize sample for 1 minute. Let tube stand for 10 minutes before opening to allow aerosols to settle
Transfer 1 ml of liquid M. tuberculosis culture into 1.5 ml conical tubes; change transfer pipettes between samples
Infectious material: all steps have to be performed in a BSC at BSL3 Laboratory

## Slide 11: DNA extraction options
Option One
  Extract DNA in BSL3 lab
Option Two
  Heat kill samples in a dry oven at 80 °C for 1 hour
  Transfer samples to BSL2 lab for DNA extraction and remainder of LPA procedure

## Slide 12: DNA extraction: Option One
Pellet bacteria by centrifuging specimen for 15 min at 10,000 x g
Use table top centrifuge with an aerosol-tight rotor in the BSC
Mark the top side of the tubes as they are placed in the rotor. Pellet is located on the bottom of the tube, opposite to mark on tube
Let the vials stand in the rack for 1–2 min to avoid aerosol generation when opening tubes
Carefully aspirate the supernatant (without removing the pellet) with a 1 ml pipette with filtered tips; change tips!
Add 100 µl molecular grade water (pre-aliquoted in 15ml sterile tubes) to pellet using a 200 l pipette with filtered tips; change tips!
Vortex tubes until pellet is completely resuspended.
Infectious material: all steps have to be performed in a BSC at BSL3 Laboratory
pellet
supernatant
water

## Slide 13
Wipe tubes with disinfectant before removing them from the BSC
Incubate tubes in a heat block or water bath at 95 °C for 20 min to kill MTBC and disrupt cell wall
Always check temperature before incubation!
Water bath temperature is more accurate than temperature in heat block
Heat block diminishes the chance of contamination from the water
Discard water after each use
Clean water bath with 0.5% freshly prepared bleach (20 min) followed by 70% alcohol
MTB in drops inside vial caps may remain viable!
Use of a small oven may be the best and safest solution (80 °C for 1h)
If heat-kill was properly carried out, specimen is considered non-infectious after this point
Water bath
Infectious material: all steps have to be performed in a BSC at BSL3 Laboratory
Specimen preparation (3)

## Slide 14
Fill sonicator with filtered distilled water
Turn on sonicator for 15 min and de-gas before use
Check function with a piece of aluminum foil
Place tubes into a floating rack, turn on sonicator and incubate tubes for 15 min at room temperature
Water temperature in sonicator should not exceed 30 °C. Higher temperature (37 °C) may activate DNAse and RNAse in specimen
Discard water after each use
Clean the sonicator tank with 5% freshly prepared bleach (20 min) followed by 70% alcohol
sonicator
Below 30°C
Specimen preparation (4)

## Slide 15
Centrifuge specimens for 15 min at 13,000 x g
Use table top centrifuge with an aerosol-tight rotor in the BSC
Transfer supernatant containing DNA to new 1.5 ml screw-cap vials using a 1 ml pipette with filtered tips; change tips!
Do not stir up pellet – cell debris to discard
Freeze leftover specimens at -20 °C
Specimens may be kept at 4 °C for 24 - 48h
DNA ready for addition to master mix
Specimen preparation (5)
! Always use screw-cap, not flip-cap, vials in order to avoid contamination

## Slide 16
Decontaminate any materials (vials, tubes, etc.) before removal from the BSC and PCR hood
Clean pipettes, racks, instruments and the BSC with freshly diluted 0.5% bleach (20 minutes), followed by 70% alcohol wash
Use area-dedicated spray bottles or beakers (separate beakers for surface cleaning and instruments)
Do not take anything from this area to Reagent Preparation Area
Consider performing specimen preparation on a separate day to enable unidirectional work flow between areas
Infectious material: all steps have to be performed in a BSC at BSL3 Laboratory
Specimen preparation (7)

## Slide 17: DNA extraction in BSL2
Heat-kill in samples in dry oven in BSL3
Centrifuge
Sonicate
Centrifuge
Remove DNA supernatant to new 1.5 ml conical screw cap tubes
Freeze DNA until ready to add to master mix

## Slide 18: Contamination control Specimen preparation area
Tools:
  Tools and instruments used here have to be labelled accordingly and should not be used elsewhere
  Never take anything from this site to the Reagent Preparation Area
  You may only take the PCR tubes in an aluminum foil-covered beaker to Amplification/Detection Area.
  Do not bring this beaker back to Specimen or Reagent Preparation Area. Wash and autoclave before reusing
  Always use filtered tips
  Do not try to put cotton plugs into unfiltered tips
  Do not use flip-cap tubes, always use screw-cap tubes
Good Laboratory Practice:
  Do specimen preparation on a separate day (earlier) and freeze extracted DNA
  Proceed with reagent preparation and amplification/detection on the following day
  This way work can be unidirectional
  Add DNA extract to PCR tubes in a PCR hood, use BSC only for DNA extraction
  If applicable, dedicate a BSC only for molecular work
  Change gloves if they get contaminated by specimens

## Slide 19: Contamination control Specimen preparation area
Reagents:
  Follow the instructions for Na-hypochlorite as described earlier
  Use 70% alcohol instead of distilled water for cleaning to avoid contamination of distilled water stocks
  Aliquot sterile molecular grade water into 15 ml tubes
  Use an Eppendorf repeat pipette with individually wrapped, sterile combitips for aliquoting, if possible
Laboratory cleaning:
  Always properly clean the work area after completing the work
  Establish a regular (e.g. weekly) and thorough laboratory cleaning protocol (floors, doors, walls)

## Slide 20: Summary
Organize your work (use checklist)
Work unidirectionally (laboratory layout)
Develop and follow SOP for procedure and instrument maintenance
Keep tools and workplace clean!

## Slide 21: Acknowledgement
<number>