## Slide 1: GenoType® MTBDRplus
Reagent preparation

## Slide 2: Content outline
PCR laboratory layout
Checklist
Calculation of volume of master mix components
Personal protective equipment (PPE)
Reagent preparation for PCR
Contamination control

## Slide 3: Optimal PCR laboratory layout
DNA-“dirty” area
DNA-free
area
Amplicon-
“dirty” area

## Slide 4: Reagent preparation (1) Preparations DNA-free area
Prepare a written checklist of materials and tools needed
Prepare a worksheet with list of specimens (identifiers) to be tested
Worksheets used in Specimen Preparation area should not be brought over to this area!
Prepare a worksheet to calculate the volume master mix ingredients
Notes: <number>
Use of written checklists avoids the unnecessary interruption of work in the BSC that can compromise the quality of the specimen testing or the biosafety of laboratory staff.

## Slide 5: Reagent preparation (2) Calculation of components DNA-free area
Determine the number of specimens and controls
Use this number plus 1 to calculate the volume of the different components of the master mix
The volume of master mix is 45 l per specimen

## Slide 6: Reagent preparation (3) Setting up workplace DNA-free area
Put on gloves and laboratory coat
Decontaminate work area with freshly diluted 0.5% bleach (20 min), followed by 70% alcohol
Cover the work area with clean paper towels

## Slide 7: Reagent preparation (4) Master mix preparation DNA-free area
Store molecular grade water, and other reagents, in the DNA-free area
Completely thaw all reagents
Pipette calculated amount of different components of master mix into 1.5 ml conical screw-cap vial using filtered tips
Change tips between reagents
Always use a new aliquot of molecular grade water
Prepare master mix, beginning with the PNM and adding DNA polymerase last
Always open only one reagent vial at a time, and place vials back in mini cooler after use
Gently (avoid breaking up primers!) vortex master mix vial and spin in Minispin for 2 seconds

## Slide 8: Reagent preparation (5) Dispensing master mix DNA-free area
Label PCR tubes in line with numbers of specimens and controls on your worksheet
Place PCR tubes onto the PCR tube rack
Dispense 45 µl master mix into each PCR tube using filtered tip
Add 5 µl of same sterile, molecular grade water used to prepare master mix into negative control PCR tube
Close PCR tubes with PCR tube caps and place them in a sterile glass beaker covered with aluminum foil (make holes in the foil for vials)

## Slide 9: Reagent preparation (6) Clean-up of workplace DNA-free area
After each use, clean pipettes, racks, instruments and the BSC with freshly diluted 0.5% bleach (20 minutes), followed by 70% alcohol and UV light
Use area-dedicated spray flasks or beakers (separate beakers for surface cleaning and instruments)

## Slide 10: Contamination control Reagent preparation area
Tools:
  Tools and instruments that are used here have to be labelled accordingly and cannot be used elsewhere
  PCR tubes can only be taken to specimen preparation area in an aluminum foil-covered beaker
  Do not bring back this beaker to reagent preparation area
  Use 0.5 ml or 0.2 ml individual PCR tubes instead of 0.2 ml strips
  Use a PCR hood with UV
Good laboratory practice:
  Change shoes or use shoe cover before entering clean room Autoclave and calibrate pipettes regularly
  After you have finished working in the specimen preparation area, do not come back and clean the reagent preparation area workplace

## Slide 11: Contamination control Reagent preparation area
Reagents:
  Always prepare fresh Na-hypochlorite solution
  Check undiluted bleach concentration
  Final concentration of hypochlorite should not fall below 0.5%
  Working solutions of 5 g/l are necessary for dealing with biohazard spills
  Containers containing bleach CANNOT be autoclaved!
  To avoid contamination of distilled water stocks, use 70% alcohol instead of distilled water for cleaning
  Keep master mix reagents in desktop cooler, especially Taq polymerase
  Aliquot sterile molecular grade water into 15 ml tubes
  Use an Eppendorf repeat pipettor with individually wrapped sterile combitips for aliquoting, if possible
Laboratory cleaning:
  Always properly clean the work area after completing the work
  Establish a regular (e.g. weekly) and thorough laboratory cleaning protocol (floors, doors, walls)

## Slide 12: Summary
Organize your work (use checklist)
Work unidirectionally (laboratory layout)
Develop and follow SOP for procedure and instrument maintenance
Keep tools and workplace clean!

## Slide 13: Acknowledgement
<number>