## Slide 1: GenoType® MTBDRplus
Amplification and detection

## Slide 2: Content outline
PCR laboratory layout
Check list
Personal protective equipment (PPE)
Amplification
Detection
Contamination control

## Slide 3: Optimal PCR laboratory layout
DNA-“dirty” area
DNA-free
area
Amplicon-
“dirty” area

## Slide 4: Amplification Thermocycler setup
Place PCR tubes into thermocycler
Programme and run the following protocol for direct specimens:
  denaturation at 95 °C for 15 min
  10 cycles of denaturation at 95 °C for 30 sec and elongation at 58 °C for 120 sec
  an additional 30 cycles of denaturation at 95 °C for 25 sec, annealing at 53 °C for 40 sec, elongation at 70 °C for 40 sec
  final extension at 70 °C for 8 min

## Slide 5: Amplification Contamination control and amplicon storage
Take the beaker that was used to carry the PCR tubes to the autoclave room for disposal
Do not take anything from this area back to the specimen or reagent preparation area
After the cycles are complete, the tubes should be placed in a 4 °C refrigerator in this area; proceed directly to hybridization if refrigerator is not available

## Slide 6
Overview of reverse hybridization
Binding of an enzyme-conjugated protein
Enzymatic conversion of dye
Separation of amplicons in single-stranded DNA
Removal of unspecifically bound DNA probes
Binding of labelled amplicons to probes
Substrate reaction
Denaturation
Hybridization
Stringent wash
Conjugate reaction

## Slide 7: Detection (1) Preparations
Prepare a written checklist of materials and tools needed
Put on gloves and laboratory coat
Decontaminate the work area with freshly diluted 0.5% bleach (20 minutes), followed by 70% alcohol
Cover the work area with clean paper towels
Use a discard container with a plastic bag and freshly diluted 0.5% bleach

## Slide 8: Detection (2) Preparation of reagents
Using clean/new tweezers, place strips onto a clean sheet of paper and label
Pre-warm rinse solution and sterile distilled water to room temperature
Using a suitable tube (15 or 50 ml sterile Falcon), dilute conjugate concentrate 1:100 with conjugate diluent
Using a suitable tube (15 or 50 ml sterile Falcon) dilute substrate concentrate 1:100 with substrate diluent - protect from light
For each strip, add 10 µl concentrate to 1 ml of the respective diluting buffer
Keep conjugate concentrate and substrate concentrate refrigerated

## Slide 9: Detection (3) Instrument and reagent preparation
Turn on TwinCubator and start the programme to enable it to warm to 45 °C
Pre-warm hybridization (HYB) buffer and stringent wash (STR) solution to 37–45 °C
Reagents must be free of precipitates!

## Slide 10: Detection (4) Denaturation
Use clean/new tray
Dispense 20 l denaturation (DEN) solution in a corner of each trough
Add 20 l amplicon to DEN and carefully mix by pipetting up and down 5 times
Do not splash mixture and contaminate neighbouring trough

## Slide 11: Detection (5) Concept of denaturation

## Slide 12: Detection (6) Hybridization
Carefully add 1 ml of pre-warmed HYB buffer to each trough; change tips between samples!
Gently rock trough to homogenously mix HYB with amplicon
Do not make any spills into neighbouring troughs – CROSS- CONTAMINATION!

## Slide 13: Detection (7) Placing the strips
Place numbered strips face side up using clean/new plastic tweezers
Strips must be completely covered by HYB. If strips turn over, re-position them with a fresh pipette tip
Place tray onto TwinCubator and press START
Incubate tray at 45 °C for 30 min
Completely aspirate HYB using transfer pipettes or vacuum aspirator
Do not decant the contents of the tray into a sink or jar; this could contaminate the work area!

## Slide 14: Detection (8) Concept of hybridization

## Slide 15: Detection (9) Stringent wash of non-hybridized amplicons
Wipe off condensation that forms on lid before every incubation step
Add 1 ml of STR solution to each trough
Press right arrow and incubate tray for 15 min at 45 °C on TwinCubator
Completely aspirate STR solution by transfer pipettes or vacuum aspirator
Do not decant the content of the tray into a sink or jar; this could contaminate the work area!

## Slide 16: Detection (10) Concept of stringent washing
A.
B.

## Slide 17: Detection (11) Removal of stringent wash solution
Wash each strip once with 1 ml of rinse (RIN) solution for 1 minute on TwinCubator
Completely aspirate the RIN solution by transfer pipettes or vacuum aspirator
Do not decant the content of the tray into a sink or jar; this could contaminate the work area!
25 °C temperature for rest of procedure

## Slide 18: Detection (12) Conjugation
Add 1 ml of diluted conjugate (CON-D) solution to each trough
Press right arrow and incubate for 30 min on TwinCubator
Completely aspirate CON-D solution by transfer pipettes or vacuum aspirator
Do not decant the content of the tray into a sink or jar; this could contaminate the work area!
1 ml diluted conjugate per trough

## Slide 19: Detection (13) Concept of conjugation reaction

## Slide 20: Detection (14) Removal of unbound conjugant
Wash each strip TWICE with 1 ml of RIN solution for 1 minute on TwinCubator
Completely aspirate RIN solution by transfer pipettes or vacuum aspirator
Do not decant the content of the tray into a sink or jar; this could contaminate the work area!
Rinse strips TWICE with RIN
to remove unbound CON-D!

## Slide 21: Detection (15) Removal of rinse solution
Wash each strip once with 1 ml of sterile distilled water (H2O) for 1 minute on TwinCubator
Completely aspirate H2O by transfer pipettes or vacuum aspirator
Do not decant the content of the tray into a sink or jar; this could contaminate the work area!
Wash strips ONCE with H2O
to remove RIN and unbound CON-D!

## Slide 22: Detection (16) Colour development
Add 1 ml of Diluted Substrate (SUB-D) to each trough
Press right arrow and incubate for 30 min on TwinCubator
Cover tray with aluminum foil
Completely aspirate SUB-D solution by transfer pipettes or vacuum aspirator
Do not decant the content of the tray into a sink or jar; this could contaminate the work area!
1 ml Diluted Substrate per trough.
Wrap SUB-D tube with aluminum foil

## Slide 23: Detection (17) Concept of substrate reaction/colour development
z
substrate
colourless
substrate
brown
DNA•STRIP®

## Slide 24: Detection (18) Removal of unbound substrate
Wash each strip TWICE with 1 ml of sterile distilled water (H2O) for 1 min on Twincubator to stop colour development reaction
Completely aspirate H2O by transfer pipettes or vacuum aspirator
Do not decant the content of the tray into a sink or jar; this could contaminate the work area!
Using clean tweezers, remove strips from the tray and dry them between two layers of absorbent paper
Wash strips TWICE with H2O
to remove unbound SUB-D!

## Slide 25: Detection (19) Clean-up of workplace
Clean pipettes, racks, instruments and the work area with freshly diluted 0.5% bleach (20 minutes), followed by 70% alcohol after every use
Clean strip trays, tweezers with freshly diluted 0.5% bleach (20 min), followed by sterile distilled water after every use
Use area-dedicated spray flasks or beakers (separate beakers for surface cleaning and instruments)
Do not take anything from this area either to reagent or specimen preparation area!

## Slide 26: Contamination control Amplification/detection area
Tools
  Tools and instruments used here have to be labelled accordingly and cannot be used elsewhere
  Never take anything from this site to reagent or specimen preparation area
  Use an Eppendorf repeat pipettor with individually wrapped sterile combitips for aliquoting reagents into trough BUT ONLY after the STR step
  Denaturation solution can be aliquoted by Eppendorf repeat pipette
  Do not re-use Falcon tubes for conjugate and substrate preparation
Good laboratory practice
  Do specimen preparation on a separate day (earlier) and freeze extracted DNA
  Proceed with reagent preparation and amplification/detection on the following day
  This way work can be unidirectional
  Add DNA extract to PCR tubes in a PCR hood (use BSC only for DNA extraction)
  If applicable, dedicate a BSC only for molecular work
  Do not decant content of the strip tray into sink or any other container. This could contaminate the work area!
  Change gloves if they get contaminated amplicons
  Decontaminate trays and tweezers (Stratalinker UV)
  Keep clean(ed) strip trays and tweezers in a clean and sealed plastic bag

## Slide 27: Contamination control Amplification/detection area
Reagents:
  Follow the instructions for Na-hyphochlorite as described earlier
  To avoid contamination of distilled water stocks, use 70% alcohol to clean workplace and instruments after bleach instead of distilled water
  However, use sterile distilled water to rinse strip trays and tweezers after bleach decontamination
  Aliquot sterile molecular grade water into 15 ml tubes
Results:
  Mixed susceptible/resistant pattern might be a result of cross-over or carry-over contamination
Laboratory cleaning:
  Always properly clean work area after completion of tasks
  Establish a regular (e.g. weekly) and thorough laboratory cleaning protocols (floors, doors, walls)

## Slide 28: Summary
Organize your work (use checklist)
Work unidirectionally (laboratory layout)
Develop and follow SOP for procedure and instrument maintenance
Keep tools and workplace clean!

## Slide 29: Acknowledgement
<number>