## Slide 1: Biosafety in the molecular laboratory

## Slide 2: Which BSL for TB laboratory activities?
<number>
BSL1 – AFB smear microscopy
  Can use countertop if airflow is directed away from laboratory staff
  Need for exhaust fan to draw air away from area
  BSC preferred
BSL2 – Specimen processing for TB culture
  Work performed in a BSC
  Safety centrifuge cups opened only in BSC
BSL2 – Molecular laboratory
  Three areas with separate airflow
  NO LIVE TB; dead cells only
BSL2 and/or BSL3 – TB Culture: Identification and DST
  All work performed in a BSC
  Isolation and air exhaust needed for manipulating large quantities of live TB

## Slide 3: TB culture laboratory
Follow all biosafety protocols for work in the TB culture lab
  Collect samples for molecular testing
  Disinfect outside of tubes before removing from BSC
  Heat kill samples if DNA extraction occurs in the moderate-risk TB laboratory
<number>

## Slide 4: Molecular laboratories
Three separate rooms
  No mixing of air between the three rooms
  No sharing of supplies, laboratory coats and gloves
Good laboratory practices
  Minimize aerosol generation – cross contamination
  Disinfect areas with freshly prepared 0.5% concentration of bleach
    Know concentration of hypochlorite in bleach
<number>

## Slide 5: PPE
<number>
Three laboratory areas with PPE stored in each area
Laboratory coat
  Front-closing
  Long-sleeved
  Washable cloth or disposable
Gloves
  Single use
  Latex-free
Notes: <number>

## Slide 6: Disinfectants
Freshly prepared 0.5% bleach
Follow by a 70% ethanol wash inside the BSC and where metal surfaces are present
Nucleic acid contamination
  DNA Away removes nucleic acid contamination from pipettes and metal and plastic surfaces
<number>

## Slide 7: Equipment

## Slide 8: Genotype MTBDRplus equipment
Centrifuge
Vortex
Water bath
Heating block/dry oven
Sonicator
PCR hood
Thermocycler
Twincubator
GT-Blot
<number>
Notes: <number>

## Slide 9: Centrifuge
Sedimentation by centripetal force
RPM (revolutions per minute) or x g
RPM = the number of full rotations completed in one minute
x g = centrifugal force
Sedimentation of cells is to one side of the tube – mark tube for pellet orientation before centrifugation
Different sizes: microcentrifuge/picofuge (0.2 ml – 2.0 ml)
<number>
Notes: <number>

## Slide 10: Centrifuge
Important:
  Wipe with a clean cloth after use
  Record use on the log sheet provided by the individual laboratories
  Record maintenance activities
<number>
Notes: <number>

## Slide 11: Vortex
Distribute contents uniformly
(mix reagents)
Temperature stability for liquids
It is crucial to heat samples in
order to totally lyse cells and
inactivate vegetative bacteria
<number>
Notes: <number>

## Slide 12: Vortex and water bath
Important:
  The vortex must be disinfected on a regular basis
  Water in the waterbath must be replaced weekly
  Clean with mild detergent
  Keep the heating element covered
<number>
Notes: <number>

## Slide 13: Heating block
Temperature stable
Heat several tubes at the same time
Important:
Clean the heating block
with 70% ethanol after
use
<number>
Notes: <number>

## Slide 14: Dry oven
80 °C for 1 hour to kill tubercle bacilli
May use instead of heating block and waterbath
External thermometer to record temperature
Place specimen racks inside oven and do not remove once temperature is stable
Transfer specimen tubes to heated racks
Clean as required with 0.5% bleach and 70% alcohol (turn off before cleaning)
<number>

## Slide 15: Sonicator
Ultrasound waves agitate particles and break bonds
Cavitation bubbles implode and generate powerful shockwaves that disrupt bacterial cell walls
<number>
Notes: <number>

## Slide 16: Sonicator
Important:
  Fill chamber with deionized water and de-gas to remove bubbles
  Do not place items directly on the bottom of the unit
  Make sure tubes are submerged in the water and the lids are floating on top
Maintenance:
  Weekly:
    Drain water out of chamber
    Clean inside and outside chamber with a soft cloth
    Do not touch coils on bottom
    Fill with deionized water
<number>
Notes: <number>

## Slide 17: PCR hood
Used in the ‘clean room’/DNA-free
  room
For preparation of the “master mix”
NO DNA in this room
Prevent contamination
<number>
Notes: <number>

## Slide 18: PCR hood
Important:
  Clean before and after with bleach (prepared daily)
  Switch on the UV light after use
  Make sure the UV light is switched off while working in the PCR hood
  Allocate equipment for the “clean room” and DO NOT USE IN OTHER ROOMS
  Do not enter the clean room after DNA extraction/entering of amplification room
<number>
Notes: <number>

## Slide 19: Thermocycler
Designed for the amplification of the specific gene target by polymerase chain reaction
High accuracy – amplification profiles can be modified to optimize results
Important:
  Switch off and unplug before
  cleaning
  Clean outside housing with
  ethanol and wipe clean after use
<number>
Notes: <number>

## Slide 20: TwinCubator
Designed for use with DNA strips in reverse hybridization assays
Manual procedure
Developed for the Hain molecular diagnostic assays
Dry incubation (no water bath
needed)
ONLY 12 SAMPLES
<number>
Notes: <number>

## Slide 21: Twincubator
Important:
  Clean before and after use
  Switch off and unplug before cleaning
  Clean outside housing with 70% ethanol, wipe clean
<number>
Notes: <number>

## Slide 22: GT-Blot
GT-Blot 20 and GT-Blot 48
The GT-Blot is a fully automated system for hybridization
Optimization of assay (programmable)
<number>
Notes: <number>

## Slide 23
Scanner: The strips can be scanned and interpreted in less than 1 minute
<number>
Notes: <number>

## Slide 24: GT-Blot
Important:
  Wash machine before and after use
  Remove reagents after the last run of the day – wash, take “connectors” out of water at the back
  Do not leave the connectors in water
  Clean outside housing with ethanol
  Wash trays with water after each run
  Wash trays (not the plastic trays) with bleach once a week and do not leave the tray in the bleach
<number>
Notes: <number>

## Slide 25: Summary
GOOD LABORATORY PRACTICE AT ALL TIMES
TAKE CARE OF YOUR MACHINES
<number>
Notes: <number>

## Slide 26: Acknowledgement
<number>