## Sheet1
LPA Equipment Use, Maintenance and Servicing Summary
Equipment | Biological Safety Cabinet (BSC) | Centrifuge | Microcentrifuge | Hot air oven | Sonicator | Thermocycler | Twincubator | Water bath | Vortex | Pipette | PCR Hood | Thermometer
Use | BSC is enclosed and ventilated work station, used when working with hazardous microorganisms. Protects operator and environment from microbiological hazards MUST be properly installed, used and maintained Recommend leave BSC fan running at all times - or switch on at least 15 mins before use | Must have aerosol-free safety cups for TB work Safety cups provide protection from aerosols produced during centrifugation and from tube breakages Safety cups must have O-rings to provide air-tight seal Only open safety cups inside BSC | Used to spin down heat-killed, non-infectious samples | Sterlisation of samples by dry heat Use a calibrated thermometer to record temperature - do not rely on digital readout only | Uses ultrasound energy to agitate particles in a sample and break bonds Cavitation bubbles implode and generate powerful shockwaves that disrupt bacterial cell walls Fill with deionized water and degas to remove air bubbles Do not place items directly on the bottom of the unit | Designed for nucleic acid amplification by polymerase chain reaction (PCR) Highly accurate, programmable heat block - raises and lowers the temperature in pre-programmed steps | Programmable, shaking incubator Designed for use with DNA strip assays (Hain Lifescience) | Used to heat and maintain reagents at a constant temperature, e.g. for reverse hybridisation reaction | Used to mix small volumes of samples and solutions. Most vortex mixers have variable speeds and can run continuously, or to run only when a tube is inserted in the rubber cup | Set of pipettes for each work area - keep in one place, store in carousel Number pipettes for use and calibration purposes Use appropriate pipette according to volume being measured | Reduces cross-contamination - protects against airbourne or surface contamination UV light inactivates DNA (denatures nucleic acids) between experiments | Use thermometers that measure maximum and minimum temperature Different temperature ranges available depending on equipment being monitored
Maintenance and Servicing
With every use | Disinfect with mycobactericidal disinfectant before and after use. Wipe afterwards with water if corrosive agent (bleach) used. Appropriate disposal of infectious liquid and solid waste | Check for spills after cycle . Disinfect interior of rotor after use with 70% ethanol. Record use in logbook. | Check for spills after cycle . Disinfect interior of rotor after use with 70% ethanol. Record use in logbook. | [6 empty] | Disinfect surface with 70% ethanol | Clean surfaces with fresh 1% bleach solution before and after use and immediately after a spill
Daily | Check airflow on magnahelic gauge, check alarm and check airflow manually with taped piece of tissue at BSC opening Call service technician in case of any problems. Switch on UV light for at least 30 mins at end of work day (optional) |  |  | Monitor temperature with max-min thermometer (calibrated) |  | Disinfect surface with 70% ethanol | Disinfect surface with 70% ethanol | Disinfect surface with 70% ethanol | Disinfect surface with 70% ethanol
Weekly | Thorough disinfection including removal and cleaning of tray followin manufacturer's instructions | Disinfect surfaces with 70% ethanol. Disinfect then wash safety cups and rotor, lubricate O-rings with silicone before replacing | Disinfect surfaces with 70% ethanol. | Disinfect surfaces with 70% ethanol. | Empty unit and clean inside and out with a soft cloth Re-fill with deionised water | Disinfect surfaces with 70% ethanol. | Disinfect surfaces with 70% ethanol. | Disinfect surfaces with 70% ethanol. | Disinfect surfaces with 70% ethanol. | Take apart barrel and wipe inner surfaces with 70% ethanol, also if solution aspirated into pipette
Monthly | [5 empty] | Switch off and unplug before cleaning Clean outside with 70% ethanol | Switch off and unplug before cleaning Clean outside with 70% ethanol
Every 3 months | [5 empty] | Run calibration cycle according to manufacturer's recommendation |  |  |  | Calibrate pipettes. Need calibrated balance for pipette calibration
Bi-annually | Service by qualified technician - HEPA filter, down-flow and face velocity, negative pressure/ventilation rate, air-flow smoke pattern, electrical system, UV light MUST decontaminate with formaldehyde vapour before servicing (preferably every 6 months, minimum annually) |  |  |  |  | Temperature check with wire sensor - check different wells randomly Return to manufacturer for temperature calibration if temperature checks fail | Temperature check with wire sensor - check different wells randomly Return to manufacturer for temperature calibration if temperature checks fail
Annually | Service by qualified technician - see above. | Service by qualified technician - to include electrical check, centrifuge brushes, verify speed and timer | Service by qualified technician - to include electrical check, centrifuge brushes, verify speed and timer | Service by qualified technician - check electrical system, temperature | Service by qualified technician - check electrical system, temperature | Service by qualified technician - check electrical system, temperature and timing | Service by qualified technician - check electrical system, temperature and timing | Calibration by qualified technician - electrical check. Store calibration certificates in equipment file | Calibration by qualified technician - electrical check, change rubber cup as needed. Store calibration certificates in equipment file | Calibration by qualified technician Store calibration certificates in equipment file | Calibration by qualified technician. Replace UV bulb and filters according to manufacturer's recommendation. Store calibration certificates in equipment file | Calibration by qualified technician. Store calibration certificates in equipment file
As required | Needs validation at installation and re-valildation if BSC is moved MUST decontaminate with formaldehyde vapour before servicing | Refer to Biosafety module for cleaning after broken tube in centrifuge |  | Clean and disinfect immediately after spill | Clean and disinfect immediately after spill | Clean and disinfect immediately after spill | Clean and disinfect immediately after spill | Clean and disinfect immediately after spill | Clean and disinfect immediately after spill | Clean and disinfect immediately after spill |  | Clean and disinfect immediately after spill
SOP
Forms | BSC logbook | Centrifuge logbook | Microcentrifuge logbook | Temperature logbook
Other references