## Slide 1: GenoType® MTBDRplus V2.0
Amplification and detection

## Slide 2: Content outline
PCR laboratory layout
Check list
Personal protective equipment (PPE)
Amplification
Detection
Contamination control

## Slide 3: Optimal PCR laboratory lay out
DNA-“dirty” area
DNA-free
area
Amplicon-
“dirty” area

## Slide 4: Amplification Program for smear-positive or smear-negative pulmonary specimens
Place PCR tubes into thermocycler
Program and run the following protocol for clinical specimens:
  denaturation at 95°C for 15 min;
  20 cycles of denaturation at 95°C for 30 s and elongation at 65°C for 2 min;
  an additional 30 cycles of denaturation at 95°C for 25 sec, annealing at 50°C for 40 sec, elongation at 70°C for 40 sec;
  final extension at 70°C for 8 min.
MDR DIR program for direct clinical specimens

## Slide 5: Amplification Program for cultured specimens
Place PCR tubes into thermocycler
Program and run the following protocol for cultured specimens:
  denaturation at 95°C for 15 min;
  10 cycles of denaturation at 95°C for 30 s and elongation at 65°C for 2 min;
  an additional 20 cycles of denaturation at 95°C for 25 sec, annealing at 50°C for 40 sec, elongation at 70°C for 40 sec;
  final extension at 70°C for 8 min.

## Slide 6: Amplification Contamination control and amplicon storage
Take the beaker that was used to carry the PCR tubes to the autoclave room for disposal
Do not take anything from this area back to the Specimen or Reagent Preparation areas
After the cycles are complete, the tubes should be placed in a 4°C (max. 1 week) household refrigerator in this area if one is not proceeding directly to hybridization
Avoid overnight PCR
MDR CUL program for cultured specimens

## Slide 7
Detection 1
Overview of Reverse Hybridization
Binding of an enzyme-conjugated protein
Enzymatic conversion of dye
Separation of amplicons in single-stranded DNA
Removal of unspecifically bound DNA probes
Binding of labeled amplicons to probes
Substrate Reaction
Denaturation
Hybridization
Stringent Wash
Conjugate Reaction

## Slide 8: Detection 2 Preparation of reagents
Using clean/new tweezers place strips onto a clean sheet of paper and label them
Pre-warm rinse solution and sterile distilled water to room temperature
Using a suitable tube (15ml or 50ml Falcon tube), dilute the conjugate concentrate (CON-C) 1:100 with conjugate diluent (CON-D)
Using a suitable tube (15ml or 50ml Falcon tube), dilute the substrate concentrate (SUB-C) 1:100 with substrate diluent (SUB-D)
Use fresh tube for dilutions for each run
For each strip, add 10 µl concentrate to 1 ml of the respective diluting buffer,
Keep Conjugate concentrate and Substrate concentrate refrigerated

## Slide 9: Detection 3 Instrument and reagent preparation
Turn on Twincubator and start the program to enable it to warm to 45°C
Check temperature
Pre-warm Hybridization (HYB) buffer and Stringent wash (STR) solution to 37-45°C
Reagents must be free of precipitates!

## Slide 10: Detection 4 Denaturation
Use clean/new tray
Dispense 20 ml Denaturation (DEN) solution in a corner of each trough, no need to change tips
Add 20 µl amplicon to DEN and carefully mix by pipetting up and down 5 times, change tips
Do not splash mixture and contaminate neighbouring trough

## Slide 11: Detection 5 Concept of denaturation

## Slide 12: Detection 6 Hybridization
Carefully add 1 mL of pre-warmed HYB buffer to each trough using filtered tips. Change tips!
Put the necessary amount of reagents (HYB, STR, RIN) into a 50 mL tube and use this as stock. Use long tips!
Gently shake trough to homogenously mix HYB with amplicon
Do not make any spills into neighboring troughs

## Slide 13: Detection 7 Placing the strips
Place numbered strips face side up using clean/new plastic tweezers
Strips must be completely covered by HYB. If strips turn over, re-position them with a fresh pipette tip
Place tray onto Twincubator and press START
Incubate tray at 45°C for 30 min
When alarm goes off, press right arrow to stop
Completely aspirate HYB using transfer pipettes, change pipettes
Do not decant the contents of the tray into a sink or jar; this can contaminate the work area!

## Slide 14: Detection 8 Concept of hybridization

## Slide 15: Detection 9 Stringent wash of non-hybridized amplicons
Wipe off condensation that forms on lid before every incubation step
Add 1 mL of STR solution to each trough
Press right arrow and incubate tray for 15 min at 45°C on Twincubator
When alarm goes off, press right arrow to stop
Completely aspirate STR solution by transfer pipette, change pipette
Do not decant the content of the tray into a sink or jar, that can contaminate the work area!

## Slide 16: Detection 10 Concept of stringent washing
A.
B.

## Slide 17: Detection 11 Removal of stringent wash solution
Wash each strip once with 1 ml of Rinse (RIN) solution for 1 min on Twincubator, no need to change tip between troughs
When alarm goes off, press right arrow to stop
Completely aspirate the RIN solution by transfer pipette, no need to change pipette between troughs
Content of the tray may be decanted into a container with 1% bleach from this point
Work at room temperature from this step forward.

## Slide 18: Detection 12 Conjugation
Add 1 mL of Diluted Conjugate (CON-D) solution to each trough using, no need to change tip
Press right arrow and incubate for 30 min on TwinCubator
When alarm goes off, press right arrow to stop
Completely aspirate CON-D solution by transfer pipette, no need to change pipette between troughs
Content of the tray may be decanted into a container with 1% bleach
1 ml Diluted Conjugate per trough

## Slide 19: Detection 13 Concept of conjugation reaction

## Slide 20: Detection 14 Removal of conjugate solution
Wash each strip TWICE with 1 ml of Rinse (RIN) solution for 1 min on Twincubator, no need to change tip between troughs
When alarm goes off, press right arrow to stop
Completely aspirate the RIN solution by transfer pipette, no need to change pipette between troughs
Content of the tray may be decanted into a container with 1% bleach from this point
Work at room temperature from this step forward.

## Slide 21: Detection 15 Removal of rinse solution
Wash each strip once with 1ml of sterile distilled water (H2O) for 1 min on Twincubator, no need to change tip
When alarm goes off, press right arrow to stop
Completely aspirate H2O by transfer pipettes, no need to change pipette
Content of the tray may be decanted into a container with 1% bleach
Wash strips ONCE with H2O
to remove RIN and unbound CON-D!

## Slide 22: Detection 16 Color development
Add 1 mL of Diluted Substrate (SUB-D) to each trough, no need to change tip
Press right arrow and incubate for 3 - 20 min on TwinCubator
When alarm goes off, press right arrow to stop
Cover tray with aluminum foil
Completely aspirate SUB-D solution by transfer pipettes, no need to change pipette
Content of the tray may be decanted into a container with 1% bleach
1 mL Diluted Substrate per trough.
Wrap SUB-D tube with aluminum foil

## Slide 23: Detection 17 Concept of substrate reaction/ color development
z
substrate
colorless
substrate
brown
DNA•STRIP®

## Slide 24: Detection 18 Removal of unbound substrate
Wash each strip TWICE with 1 ml of sterile distilled water (H2O) for 1 min on Twincubator to stop color development reaction
Completely aspirate H2O by transfer pipettes
Content of the tray may be decanted into a container with 1% bleach
Using clean tweezers, remove strips from the tray and dry them between two layers of absorbent paper
Wash strips TWICE with H2O
to remove unbound SUB-D!

## Slide 25: Detection 19 Clean-up of work place
Clean pipettes, racks, instruments and the work area with freshly diluted 1% bleach (20 min), followed by 70% alcohol after every use
Clean strip trays, tweezers with freshly diluted 1% bleach (20 min), followed by sterile distilled water after every use
Use area-dedicated spray flasks or beakers (separate beakers for surface cleaning and instruments)
Do not take anything from this area either to Reagent or Specimen Preparation Area!

## Slide 26: Contamination control Amplification/Detection Area
Tools:
  Tools and instruments used here have to be labeled accordingly and cannot be used elsewhere
  Never take anything from this site to Reagent or Specimen Preparation Area
  Do not re-use Falcon tubes for Conjugate and Substrate preparation
Good Laboratory Practice:
  Do specimen preparation on a separate day (earlier) and freeze extracted DNA
  Proceed with reagent preparation and amplification/detection on the following day
  This way work can be unidirectional
  Add DNA extract to PCR tubes in a PCR hood (use BSC only for DNA extraction)
  If applicable, dedicate a BSC only for molecular work
  Do not decant content of the strip tray into sink or any other container during th <hyb and STR steps. This will contaminate the work area!
  Change gloves if they get contaminated amplicons
  Decontaminate trays and tweezers (Stratalinker UV)
  Keep clean(ed) strip trays and tweezers in a clean and sealed plastic bag

## Slide 27: Contamination control Amplification/Detection Area
Reagents:
  Follow the instructions for Na-hyphochlorite as described earlier
  Use 70% alcohol to clean work place and instruments after bleach instead of distilled water for cleaning to avoid contamination of distilled water stocks
  However, use sterile distilled water to rinse strip trays and tweezers after bleach decontamination
  Aliquot sterile molecular grade water into 1.5 ml tubes
Results:
  Mixed susceptible/resistant pattern might be a result of cross over or carry over contamination
Laboratory cleaning:
  Always properly clean work area after completion of tasks
  Establish a regular (e.g. weekly) and thorough laboratory cleaning protocols (floors, doors, walls)

## Slide 28: Summary
Organize your work (use checklist)
Work unidirectionally (laboratory layout)
Develop and follow SOP for procedure and instrument maintenance
Keep tools and work place clean!

## Slide 29: Acknowledgement
<number>