## Slide 1: GenoType® MTBDRplus
Troubleshooting
Contamination control
Positive and negative controls
Quality control

## Slide 2: Troubleshooting (1)
Overall weak or no signals, including Conjugate Control zone:
  Check instruments, reagents and temperatures
  Room temperature was too low or reagents were not properly pre-warmed to room temperature
  Incubation temperature too low
  No, or insufficient, CON-C and/or SUB-C was added
  Reagent expired

## Slide 3: Troubleshooting (2)
Weak or no signals, except for Conjugate Control zone:
  Check instruments, reagents and temperatures
  PMN and Taq were frozen and thawed too many times (optimal 5–6 times)
  PMN, Taq and other Master Mix components were exposed to heat during reagent preparation; use table-top coolers
  Incubation temperature too high or too low
  Check purity of processed specimen (contaminated parallel culture, Gram or AFB smear)
  Quality (presence of inhibitors) and/or quantity of extracted DNA do not allow an efficient amplification
  Check Amplified Product for presence of amplicons on 2% agarose gel
    In case of no visible amplicons, repeat DNA extraction and amplification
    In case of weak band, increase the amount of DNA added from 5 to 8 l
    In case of intense and smeared bands, dilute extracted DNA

## Slide 4: Troubleshooting (3)
Non-homogeneous staining of bands:
  Strips were not completely immersed by reagents during incubation steps
  Tray was not shaken properly
High background colour:
  Check instruments, reagents and temperatures
  CON-C and/or SUB-C were too concentrated
  Washing steps were not thorough enough
  Wash solutions were too cold
  Depending on the amount of amplified DNA used and on the specific reaction conditions, a strong and fast colour development may occur
  In such cases, discontinue incubation with SUB-D as soon as signals are clearly visible, so as to prevent development of smeared bands

## Slide 5: Contamination control General
Record specimens that were processed (NALC/ NaOH) and tested (GenoType MTBDRplus) in the same batch
Record which technician performed the test

## Slide 6: Contamination control Reagent preparation area (1)
Tools:
  Tools and instruments that are used here have to be labelled accordingly and cannot be used elsewhere
  PCR tubes can only be taken to specimen preparation area in an aluminum foil-covered beaker
  Do not bring back this beaker to reagent preparation area
  Use 0.5 ml individual PCR tubes instead of 0.2 ml strips (Stratalinker)
  Use a PCR hood with UV
Good laboratory practice:
  Change shoes or use shoe covers before entering specimen preparation area. Autoclave and calibrate pipettes regularly
  After you are finished working in the specimen preparation area; do not come back and clean the reagent preparation area workplace

## Slide 7: Contamination control Reagent preparation area (2)
Reagents:
  Always prepare fresh sodium hypochlorite solution
  Check undiluted bleach concentration
  Final concentration of bleach should not fall below 1%
  Working solutions of 5 g/l are necessary for dealing with biohazard spills
  Containers containing bleach CANNOT be autoclaved!
  To avoid contamination of distilled water stocks, use 70% alcohol instead of distilled water for cleaning
  Keep master mix reagents in desktop cooler, especially Taq polymerase
  Aliquot sterile molecular-grade water into 1.5 ml tubes, if possible
Laboratory cleaning:
  Always properly clean the work area after completing the work
  Establish a regular (e.g. weekly) and thorough laboratory cleaning protocol (floors, doors, walls)

## Slide 8: Contamination control Specimen preparation area (1)
Tools:
  Tools and instruments that are used here have to be labelled accordingly and cannot be used elsewhere
  Never take anything from this site to reagent preparation area
  PCR tubes can only be taken to amplification/detection area in an aluminum foil- covered beaker
  Do not bring back this beaker to specimen or reagent preparation area
  Always use filtered tips
  Do not try to put cotton plugs into unfiltered tips
  Do not use flip-cap tubes, always use screw-cap tubes
Good laboratory practice:
  Do specimen preparation on a separate day (earlier) and freeze extracted DNA
  Proceed with reagent preparation and amplification/detection the following day
  This way, work can be unidirectional
  Add DNA extract to PCR tubes in a PCR hood, use BSC only for DNA extraction
  If applicable, dedicate a BSC for molecular work only
  Change gloves if they get contaminated by specimens

## Slide 9: Contamination control Specimen preparation area (2)
Reagents:
  Follow the instructions for sodium hypochlorite as described earlier
  To avoid contamination of distilled water stocks, use 70% alcohol instead of distilled water for cleaning
  Aliquot sterile molecular-grade water into 1.5 ml tubes
Laboratory cleaning:
  Always properly clean the work area after completing the work
  Establish a regular (e.g. weekly) and thorough laboratory cleaning protocol (floors, doors, walls)

## Slide 10: Contamination control Amplification/detection area (1)
Tools:
  Tools and instruments that are used here must be labelled accordingly and cannot be used elsewhere
  Never take anything from this site to reagent or specimen preparation area
  Do not re-use Falcon tubes for conjugate and substrate preparation
Good laboratory practice:
  Do specimen preparation on a separate day (earlier) and freeze extracted DNA
  Proceed with reagent preparation and amplification/detection the following day
  This way work can be unidirectional
  Add DNA extract to PCR tubes in a PCR hood; use BSC only for DNA extraction
  If applicable dedicate a BSC only for molecular work
  Do not decant the content of the strip tray into sink or any other container during the HYB and STR steps; this will contaminate the work area!
  Change gloves if they get contaminated with amplicons
  Decontaminate trays and tweezers (Stratalinker UV)
  Keep clean(ed) strip trays and tweezers in a clean, sealed plastic bag

## Slide 11: Contamination control Amplification/detection area (2)
Reagents:
  Follow the instructions for sodium hypochlorite as described earlier
  To avoid contamination of distilled water stocks, use 70% alcohol instead of distilled water to clean workplace and instruments after bleach
  However, use sterile distilled water to rinse strip trays and tweezers after bleach decontamination
  Aliquot sterile molecular-grade water into 1.5 ml tubes
Results:
  Mixed susceptible/resistant pattern might be a result of cross-over or carry-over contamination
Laboratory cleaning:
  Always clean the work area properly after completing the work
  Establish a regular (e.g. weekly) and thorough laboratory cleaning protocols (floors, doors, walls)

## Slide 12: Positive and negative controls
Positive control(s):
  Always add DNA of positive control at specimen preparation area
  Use a susceptible and/or MDR heat-killed control
  Negative strips with this control indicate unsuccessful amplification
  Mixed (susceptible/resistant) pattern with positive control may indicate carry-over or cross-over contamination
Negative control(s):
  Prepare FIRST negative control at reagent preparation area using same molecular-grade water as was used for the Master Mix
  Do not open this tube at specimen preparation area
  Prepare a SECOND negative control at specimen preparation area using the same molecular-grade water as was used for specimen preparation at the very end of the procedure
  Negative control strip must have CC and AC bands, but no other bands should be visible
  Positive results with other bands of these controls may indicate carry-over or cross-over contamination

## Slide 13: Quality control (1)
Monitor the number (%) of unsuccessful amplifications for AC/TUB/rpoB/katG/inhA
Monitor the number (%) of weak amplifications for AC/ TUB/rpoB/katG/inhA
Monitor the number (%) of different mutation “types” (common and uncommon)
Monitor the number (%) of tests with mixed pattern (rpoB/katG/inhA)
Record specimens tested in one batch and technician who performed the run
Monitor turn-around time

## Slide 14: Quality control (2)
Culture and drug susceptibility testing of all specimens with resistant results
Culture 10% of specimens with susceptible result
Confirm discrepant results (molecular vs conventional) by another molecular method, if possible
Perform lot-to-lot runs with well-characterized DNA extracts

## Slide 15: Summary (1)
Organize your work (use checklist)
Work unidirectionally (laboratory layout)
Develop and follow SOPs for procedure and instrument maintenance
Keep tools and workplace clean!

## Slide 16: Summary (2)
Compare all molecular results with other laboratory results, i.e. culture, identification, conventional drug susceptibility testing and, if possible, with clinical findings

## Slide 17: Acknowledgement
<number>