# TABLE OF CONTENTS

TABLE OF CONTENTS	1

1.	PURPOSE	2

2.	SCOPE	2

3.	RESPONSIBILITIES	2

4.	CROSS-REFERENCES	2

5.	PROCEDURES	2

5.1. Equipment	2

5.2. Materials	2

5.3. Safety	3

5.4	Use of GT Blot instrument	3

5.5	Interpretation of Results	5

5.6	Maintenance of the GT Blot	6

5.6.1	Cleaning of GT Blot trays	6

5.6.2	Cleaning the GT Blot 48	6

6.	REFERENCES	7

7.	 CHANGE HISTORY	7

# PURPOSE

This SOP provides an outline of the use of the GT Blot 48 instrument.  The GT Blot 48 instrument is an automated hybridization washer that incubates line probe assay (LPA) strips at the appropriate temperature and performs the various incubation and wash steps automatically.

# SCOPE

This SOP relates to molecular LPA using Genotype kits (Hain Lifescience) performed in the ___________ TB Laboratory.

# RESPONSIBILITIES

All laboratory staff members working in the ____________ TB Laboratory are responsible for the implementation of this operating procedure.

All users of this procedure who do not understand it or are unable to carry it out as described are responsible for seeking advice from their supervisor.

# CROSS-REFERENCES

| See: | Document Matrix_TB 01-01_VX.Y.doc |
| --- | --- |
| Location: |  |

| See: | Genotype MTBDRplus TB 06-01_VX.Y.doc |
| --- | --- |
| Location: |  |

Refer to SOPs listed under TB 02 (General Procedures), TB 03 (Specimen Handling), TB 07 (Equipment Use and Maintenance) and TB 06 (Molecular Methods).

# PROCEDURES

## 5.1. Equipment

GT Blot 20 or 48 machine (Hain Lifescience)

P200 pipette

Gloves

Tweezers

## 5.2. Materials

GenoType MTBDRplus kit (Hain Lifescience): DEN solution, HYB buffer, STR buffer, TwinCubator, DNA STRIP test membranes, Conjugate Concentrate (Con-C), Conjugate Buffer (CON-D), Substrate Concentrate (Sub-C), Substrate Buffer (Sub-D)

15 ml sterile graduated conical tubes for dilutions

Absorbent paper

Forceps

Sterile filtered tips

Disposable Pasteur pipettes

Discard containers

## 5.3. Safety

Do not touch any moving parts of the instrument while it is in operation.

Do not attempt to open the electrical supply or interfere with the electrical components of the instrument at any time.

It is important not to attempt to run the instrument without liquids.  Serious damage to the instrument may occur.  The operator must position the sensor in the tray prior to starting the assay.

## Use of GT Blot instrument

Before using the instrument for analyzing samples the following series of instrument checks should be carried out:

- Ensure the power is switched OFF

- Ensure the waste tubes are routed into a suitable empty container

- One of the waste tubes is gravity fed.  Ensure the waste container is below the level of the machine and the waste tubes are not trapped under feet.

- Switch power ON (back left of the machine) and ensure that the instrument initializes properly.

- The initialization sequence of events is as follows:

  - The display will indicate “BeeBlot Ready Press Start”

  - The aspirating needles move upwards to their home position

  - The arm will move left to its home position

  - The tray mechanism will move into its home position

  - The green LED at the front of the instrument will come ON

  - The rear fan should be rotating

- Place all tubings in a container filled with distilled/deionised water

- Follow the keypad commands, and ensure the water flows though the appropriate dispensing needles under the arm when washing is in progress.

  - “BeeBlot press start”

- Using either the left or right arrow select “Washing – Assay 11” then press START three times

  - “Cleaning cycle A”  Press START

After cleaning cycle A is complete press START

  - “Cleaning cycle B” Press START

- While the cleaning cycle is in progress, prepare dilutions of the conjugate and substrate solutions.

- After cleaning cycle B is complete switch the power to the instrument OFF

- Dry tubings with a dry cloth

- Place all solutions in their respective places (colour matched( and insert the conduction tubes (e.g. the HYB solution is green, and the slot is marked with a green dot).  Check that there is sufficient solution in each bottle for the number of strips being loaded into the tray.

Switch the power to the instrument ON.

“BeeBlot start”  Press START.

- By pressing the right arrow key select “Genotype 45” Press START three times.

- It is important to locate the temperature sensor in the tray before the assay starts.  Lift the temperature sensor up carefully and put into the tray.  Make sure that the sensor is properly positioned down before starting the assay.

- Reagent preheat of HYB and STR solutions for 15 minutes will start.

- While the pre-heat is in progress, dispense 20µl of denaturation “DEN” solution in a corner of each of the wells in the tray.  One tip can be used, but should be changed in case of possible contamination.

- Add 20µl of amplified sample to the DEN solution, pipette up and down to mix and incubate at room temperature for 5 minutes.  Take great care not to contaminate the neighbouring wells during addition of amplified samples.  Use a fresh pipette tip for each addition.

- Take the LPA strips out of the tube using tweezers and label with a pencil (or Hain red marker pen) under the coloured mark on the strip.  Always use gloves when handling strips.

- After the pre-heat is complete, press START and select the number of wells to be used using the left or right arrow key.  The choice must be an even number between 2 and 48.  Press START.

- Position sensor and press START.

- Begin assay – press START

- Close lid – press START

After the priming sequence is completed the instrument will dispense pre-warmed hybridization reagent (HYB) into the tray and it will dispense water into the wells that are not used, in order to gain optimal temperature.  Check to ensure that all the wells were filled correctly.

“Add Amplicon” – open the door and place a strip in the respective well using tweezers.

The strips must be completely covered by the solution and the coated side must be facing upwards (identifiable by the mark near the lower end).

Using tweezers (or sterile pipette tip), turn over any strips that may have turned over during immersion.  Clean tweezers using bleach to avoid cross contamination.

Close the door and press START.

The tray should start to rock and the green LED will come ON.  Check that the temperature starts to rise.

Once the temperature of 45ºC is attained, hybridization will start and continues for 30  minutes incubation.

According to the assay program, the instrument will aspirate and add the necessary reagents at the correct time interval and incubate reagents at the appropriate temperature throughout the test method.

HYB (hybridization)		green			30 minutes at 45ºC

STR (stringent wash)		red			15 minutes at 45 ºC

RIN (rinse solution)					1 minute

CON (diluted conjugate)	orange			30 minutes at 25 ºC

RIN							1 minute at 25 ºC

Distilled water wash					1 minute at 25 ºC

SUB (diluted substrate)		yellow			3 minutes at 25 ºC

Distilled water wash (twice)				1 minute each

- After the reaction has stopped “?ASPIRATE tray? Will be displayed.  Press START.

- After aspiration is complete open the door and remove the tray and put it on the bench.  Take out the tubings from the reagent containers and place them in the container filled with 1% bleach solution and run the washing assay twice and repeat the washing cycle twice with de-ionised water.

- Dry tubings and place them in an empty container.  Close the reagent containers and store at 2-8ºC.

- To remove the tray, remove the three tray clamps at the bottom of the tray, lift sensor and slide tray outwards.

## Interpretation of Results

- Use forceps to transfer strips to the GenoType MTBDRplus Results Sheet provided with the kit or downloadable from Hain Lifescience website.

- Be careful to follow instructions on the pdf file which can be downloaded from the website to ensure that the correct size sheet is printed (otherwise the size of the strip on the form is smaller than the actual size and the bands will not line up).

| Use: | Genotype MTBDRplus Results Sheet.pdf |
| --- | --- |
| Location: |  |

- Read results by lining strips up to code provided with kit.

- In order for results to be valid, CC (conjugate control) and AC (amplification control) bands must appear for every sample.  The presence of TUB band indicates that M.tuberculosis complex is present in the sample. (For positive samples the AC band may be very weak or absent, but as long as TUB band is present the result can be evaluated.  For negative samples, however, the AC band MUST be present or the result can not be evaluated).

- A mutation in the relevant gene (and resistance to the relevant drug) is signified by either an absent wild type band and/or the presence of a mutant band for each gene cluster.

- The rpoB, katG and inhA each have a control band which must be present in order to interpret the results.

- rpoB predicts RIF resistance, katG predicts high level INH resistance, inhA predicts low level INH resistance.

- For results to be valid the bands must be of intensity approximately equal to or greater than the intensity of the AC band.

- In order for a batch of results to be valid, the negative control strip must have a CC and AC band present, but no other bands must be visible.

- If a positive result is obtained with the negative control, the results of the whole batch must be repeated and measures taken to remove amplicon contamination from all rooms and equipment.

- Refer to the product insert for more details of interpretation of banding patterns and troubleshooting.

- Record the results on the report form and store forms in a file.

## Maintenance of the GT Blot

## Cleaning of GT Blot trays

Soak the trays thoroughly using diluted bleach solution in the sink . Remove excess liquid and then spray with 70% ethanol and wipe wells thoroughly to remove residues (using a cotton tipped applicator stick).

After cleaning, wash well using distilled water to ensure all bleach is removed.  Bleach residues not removed may affect the colour development of the strips.

Discard the water and wash a second time with distilled water.

Pat remaining liquid on a fresh paper towel and leave to air dry.

The tray can be reused once dry.

Keep several trays available, so that they can be alternated.  Trays can be reused for quite some time, but check during cleaning that there are no holes in the bottom of the black trays.

Discard trays once a hole appears.

## Cleaning the GT Blot 48

Clean the outside of the instrument on a monthly basis using a moist lint free cloth.  Switch off the instrument and close the lid prior to cleaning.

On a weekly basis clean the insert for the tray inside the instrument carefully using 70% ethanol and a cotton tipped applicator stick to remove any residues from between the wells.  Do not use bleach inside the instrument.

# REFERENCES

GenoType MTBDRplus product insert.  Hain Lifescience. Version 1.0; 2/2007.

Barnard M, Albert H, Coetzee G, O'Brien R, Bosman ME. Rapid molecular screening for drug-resistant tuberculosis in a high-volume public health laboratory in South Africa. Am J Respir Crit Care Med. 2008 Apr 1;177(7):787-92. Epub 2008 Jan 17.

- 7.		CHANGE HISTORY

| New version # / date | Old version # / date | No. of changes | Description of changes | Source of change request |
| --- | --- | --- | --- | --- |