## Slide 1: GenoType® MTBDRplus
Specimen preparation

## Slide 2: Content outline
PCR laboratory layout
Checklist
Personal protective equipment (PPE)
Working in a Biological Safety Cabinet (BSC)
Specimen preparation for PCR
Contamination control

## Slide 3: Optimal PCR laboratory layout
DNA-“dirty” area
DNA-free
area
Amplicon-
“dirty” area

## Slide 4: Working in BSC (1)
Prepare a written checklist of materials and tools needed
Prepare a worksheet with list of specimens (identifiers) to be tested
Notes: <number>
Use of written checklists avoids the unnecessary interruption of work in the BSC that can compromise the quality of the specimen testing or the biosafety of the laboratorian.

## Slide 5: Use appropriate TB containment laboratory PPE
Gowns
  Must have solid front and can be tied at the back
  Long-sleeved with elastic cuffs
  Must be either autoclavable/washable (soak in bleach) or disposable
Respirators
  N-95
Gloves
  Single-use

## Slide 6: TB containment laboratory PPE is suitable for:
Specimen processing
Specimen inoculation
Manipulating cultures
Centrifuging
Inoculation of DSTs
Making smears from cultures and processed specimens
Biochemical testing to identify TB
Preparing samples for PCR (before heat-kill)

## Slide 7: Working in BSC (2)
Decontaminate the BSC with freshly diluted 1% bleach (20 min), followed by 70% alcohol
Cover the working area with paper towels and spray it with suitable disinfectant
Use a discard container with a plastic bag and freshly diluted 1% bleach
DO NOT block the grill
Perform all operations at least 12 cm away from the front grill on the work surface
Place all materials and aerosol-generating equipment away from the front grill
Notes: <number>
Place absorbent paper towelling on the work surface to minimize the potential of exposure to infectious material or cross contamination in case of a spill.
Blocking the grill disrupts the integrity of the cabinet air barrier and can compromise both staff and product protection.
Keeping clean materials at least 12 cm away from aerosol-generating activities will minimise the potential for cross-contamination

## Slide 8: Preparation of clinical specimens
Work in sets equivalent to one centrifuge load (8 - 16 specimens at a time)
Label 1.5 ml screw-cap conical tubes (both tubes and caps) for each specimen
Completely thaw frozen and processed clinical sediments
Vortex processed clinical specimen for 1 min
Open only one specimen at a time; do not leave open containers or centrifuge tubes in the BSC
Transfer 500 l processed clinical specimen into 1.5 ml conical screw-capped vial; change tips!
Infectious material: all steps have to be performed in a BSC
at least in an appropriate moderate TB risk precaution laboratory
Clinical specimens have to be processed by N-acetyl-L-cysteine/NaOH method before specimen preparation

## Slide 9: Sample preparation of isolates on solid culture
Work in sets equivalent to one centrifuge load (8–16 specimens at a time)
Label 1.5 ml screw cap conical tubes (both tubes and caps) for each sample
Re-suspend a half loopful (but representative of a no. of colonies) of M. tuberculosis culture on solid medium; change loops!
Do not scrape the surface of the medium since components of egg-based medium can be PCR inhibitors
Vortex suspension for 1 min to completely break up colonies and clumps
Infectious material: all steps have to be performed in a BSC at a TB containment
laboratory

## Slide 10: Sample preparation of liquid cultures
Work in sets equivalent to one centrifuge load (8–16 specimens at a time)
Label 1.5 ml screw-cap conical tubes (both tubes and caps) for each sample
Vortex liquid culture for 1 min to break up clumps
Transfer 1 ml of liquid M. tuberculosis culture into 1.5 ml conical tubes; change transfer pipettes or filtered tips!
Vortex suspension for 1 min to completely break up flakes and clumps
Infectious material: all steps have to be performed in a BSC at a TB containment
laboratory

## Slide 11: Specimen preparation (1)
Pellet bacteria by centrifuging specimen for 15 min at 10,000 x g
Use table-top centrifuge with an aerosol-tight rotor (open rotor in the BSC)
Mark the cap of the tubes to make visualization of pellet easier when removing the supernatant
Let the vials stand in the rack for 1–2 min to avoid aerosol generation when opening the tubes
Using a 1 ml pipette with filtered tips, carefully aspirate the supernatant without removing the pellet; change tips!
Add 100 l molecular grade water (pre-aliquoted in 1.5 ml sterile tubes) to pellet using a 200 l pipette with filtered tips; change tips!
Vortex tubes until pellet is completely re-suspended. This may take 1–2 min
Infectious material: all steps have to be performed in a BSC at a TB containment
laboratory

## Slide 12
Wipe tubes with disinfectant before removing them from the BSC
Incubate tubes in an dry heat block for 30 min to kill MTB and disrupt cell wall
Always check temperature before incubation!
MTB in drops on vial caps may remain viable in heat blocks or water baths!
Use of a small oven may be the best and safest solution
If heat-kill was properly carried out, specimen may be considered non-infectious after this point
Water bath
Infectious material: all steps have to be performed in a BSC at a TB containment
laboratory
Specimen preparation (2)

## Slide 13
Fill sonicator with (filtered) distilled water
Turn on sonicator for 15 min before use and check efficacy with a piece of aluminum foil
Place tubes into a floating rack, turn on sonicator and incubate tubes for 15 min at room temperature
Discard water after each use
Clean the sonicator tank with 1% freshly prepared bleach (20 min) followed by 70% alcohol
sonicator
Specimen preparation (3)

## Slide 14
Centrifuge specimens for 5 min at 13,000 x g
Use table top centrifuge with an aerosol-tight rotor (open rotor in the BSC)
Transfer supernatant containing DNA to new 1.5 ml screw-cap vials using a 1 ml pipette with filtered tips, change tips!
Do not stir up pellet
Freeze leftover specimens at -20 °C
Specimens may be kept at 4 °C for not more than 7 days
Specimen preparation (4)
! Always use screw-cap, not flip-cap, vials in order to avoid contamination

## Slide 15
Work in a PCR hood (optional)
Completely thaw supernatant
Pipette up and down
Add 5 l of each sample to the corresponding tube containing the master mix
Change filtered pipette tips between each addition and dispose in 1% bleach solution
Close PCR tubes with new PCR tube caps and place them in a sterile glass beaker covered with aluminum foil (make holes in the foil for vials)
Specimen preparation (5)

## Slide 16: Contamination control Specimen preparation area
Tools:
  Tools and instruments used here have to be labelled accordingly and should not be used elsewhere
  Never take anything from this site to the Reagent Preparation Area
  You may only take the PCR tubes in an aluminum foil-covered beaker to Amplification/Detection Area
  Do not bring this beaker back to Specimen or Reagent Preparation Area
  Always use filtered tips
  Do not try to put cotton plugs into unfiltered tips
  Do not use flip-cap tubes, always use screw-cap tubes
Good laboratory practice:
  Do specimen preparation on a separate day (earlier) and freeze extracted DNA
  Proceed with reagent preparation and amplification/ detection on the following day
  This way work can be unidirectional
  Add DNA extract to PCR tubes in a PCR hood, use BSC only for DNA extraction
  If applicable, dedicate a BSC for molecular work only
  Change gloves if they get contaminated by specimens

## Slide 17: Contamination control Specimen preparation area
Reagents:
  Follow the instructions for sodium hypochlorite, as described earlier
  Use 70% alcohol instead of distilled water for cleaning, in order to avoid contamination of distilled water stocks
  Aliquot sterile molecular grade water into 1.5 ml tubes
Laboratory cleaning:
  Always clean the work area properly after completing the work.
  Establish a regular (e.g. weekly) and thorough laboratory cleaning protocol (floors, doors, walls)

## Slide 18: Summary
Organize your work (use checklist)
Work unidirectionally (laboratory layout)
Develop and follow SOP for procedure and instrument maintenance
Keep tools and workplace clean!

## Slide 19: Acknowledgement
<number>