# TABLE OF CONTENTS

TABLE OF CONTENTS	1

1.	PURPOSE	2

2.	SCOPE	2

3.	RESPONSIBILITIES	2

4.	CROSS-REFERENCES	2

5.	PROCEDURES	2

5.1. Equipment and Materials	2

5.2	Master Mix Preparation	5

5.3. DNA Extraction using GenoLyse®	5

5.3.1	DNA extraction directly from clinical specimens	6

5.3.2  DNA extraction from liquid culture	6

5.3.3  DNA extraction from solid culture	6

5.4	Amplification	6

5.5	Hybridization	7

5.6.	Interpretation of Results	8

6.	REFERENCES	9

7.	 CHANGE HISTORY	9

# PURPOSE

This SOP describes procedure for determination of Mycobacterium tuberculosis positivity and rifampicin/isoniazid resistance by utilizing the GenoType MTBDRplus test, Version 2.0 (Hain Lifesciences), including DNA extraction using GenoLyse procedure.  Traditional microscopy takes weeks to establish a diagnosis of tuberculosis, with additional time required for conventional drug susceptibility testing.  Testing may be performed on DNA isolated from cultures as well as directly on clinical specimens.  The Genotype MTBDRplus assay is based on line probe assay (LPA) technology involving polymerase chain reaction (PCR) amplification and binding of amplicons to specific oligonucleotide probes immobilized on a membrane strip.

# SCOPE

This SOP relates to molecular assays for mycobacterial speciation performed in the _____________ TB Laboratory.

# RESPONSIBILITIES

All laboratory staff members working in the ___________ TB Laboratory are responsible for the implementation of this operating procedure.

All users of this procedure who do not understand it or are unable to carry it out as described are responsible for seeking advice from their supervisor.

# CROSS-REFERENCES

| See: | Document Matrix_TB 01-01_VX.Y.doc |
| --- | --- |
| Location: |  |

Refer to SOPs listed under TB 02 (General Procedures), 03 (Specimen Handling), 07 (Equipment Use and Maintenance) and 06 (Molecular Methods).

# PROCEDURES

## 5.1. Equipment and Materials

|  | SPECIMEN PREPARATION ROOM, TB LABORATORY |
| --- | --- |
| 1 | Latex free disposal gloves |
| 2 | Surgical gown |
| 3 | N95 mask |
| 4 | Hair cover |
| 5 | Hot air oven (must heat up to at least 105C) |
| 6 | Microcentrifuge (1.5ml tubes) |
| 7 | Screw cap tubes, 1.5ml |
| 8 | Racks for 1.5ml tubes |
| 9 | Adjustable pipette, 20µl |
| 10 | Adjustable pipette, 200µl |
| 11 | Adjustable pipette, 1000µl |
| 12 | Pipette stand |
| 13 | Filter tips for above mentioned pipettes |
| 14 | Sterile molecular grade water (nuclease free), aliquoted into 1.5ml vials |
| 15 | Discard container with plastic bag and freshly prepared 1% bleach solution |
| 16 | Paper towels and dispenser |
| 17 | Biohazard waste disposal box |
| 18 | Spray bottle for 1% bleach solution (freshly prepared) |
| 19 | Spray bottle for 70% alcohol |
| 20 | Thermometer |
| 21 | Timer |
| 22 | Marker |
| 23 | Vortex mixer |
| 24 | PCR Hood (optional), for addition of DNA |
| 25 | Positive control |
| 26 | Specimens |
| 27 | GenoLyse kit reagents |

|  | REAGENT PREPARATION (MASTER MIX) ROOM |
| --- | --- |
| 1 | Dedicated laboratory coats + coat hooks |
| 2 | Latex free disposable gloves |
| 3 | PCR hood (optional) |
| 4 | Adjustable pipette, 10µl |
| 5 | Adjustable pipette, 20µl |
| 6 | Adjustable pipette, 200µl |
| 7 | Adjustable pipette, 1000µl |
| 8 | Pipette stand |
| 9 | Filter tips for above mentioned pipettes |
| 10 | PCR tubes, 200µl plus caps |
| 11 | Screw cap tubes, 1.5ml |
| 12 | Racks for 1.5ml tubes |
| 13 | Paper towels and dispenser |
| 14 | Vortex mixer |
| 15 | Marker |
| 16 | Discard container with plastic bag and freshly prepared 1% bleach solution |
| 17 | Biohazard waste disposal box |
| 18 | Spray bottle for 1% bleach solution (freshly prepared) |
| 19 | Spray bottle for 70% alcohol |
| 20 | Timer |
| 21 | Glass beaker |
| 22 | Aluminium foil |
| 23 | Mini-spin centrifuge |
| 24 | GenoType MTBDR plus Master mix reagents (AM-A and AM-B) |
| 25 | Benchtop cooler |

|  | AMPLIFICATION AND DETECTION ROOM |
| --- | --- |
| 1 | Dedicated laboratory coats + coat hooks |
| 2 | Latex free disposable gloves |
| 3 | GT Blot (Automated) or Twincubator (Manual) |
| 4 | Thermal cycler with heated lid |
| 5 | Adjustable pipette, 200µl |
| 6 | Adjustable pipette, 1000µl |
| 7 | Pipette stand |
| 8 | Filter tips for above mentioned pipettes |
| 9 | Plastic transfer pipettes, 1ml, non-sterile |
| 10 | Discard container with plastic bag and freshly prepared 1% bleach solution |
| 11 | Biohazard waste disposal box |
| 12 | Spray bottle for 1% bleach solution (freshly prepared) |
| 13 | Spray bottle for 70% alcohol |
| 14 | Paper towels and dispenser |
| 15 | Marker |
| 16 | Tweezers / forceps (clean) |
| 17 | Absorbent paper |
| 18 | Clear adhesive tape |
| 19 | Pair of scissors |
| 20 | Waste bucket, 20l (GT Blot only) |
| 21 | Sterile distilled water |
| 22 | Aluminium foil |
| 23 | Timer |
| 24 | Plastic tray with 1% bleach solutiion to discard hybridisation tray content |
| 25 | Sterile Falcon tubes, 15ml, plus rack |
| 26 | Clean tray for strips |
| 27 | Vortex mixer |
| 28 | GenoType MTBDRplus detection reagents |

## 5.2	Master Mix Preparation

Always perform at the beginning of the day in Pre-Amplification Room wearing gloves. Wear dedicated lab coats which should remain in the pre-amplification room.

- Thoroughly apply 1% bleach solution to all surfaces of hood and worktop.

- For preparation of bleach solution:

| See: | Laboratory Cleaning and Maintenance_TB 02-04_VX.Y.doc |
| --- | --- |
| Location: |  |

- Thaw the two amplification mixes, A and B (AM-A and AM-B).

- Carefully remove PCR tubes and close lids. Place in rack in hood and label.

- Determine the number of samples to be tested plus control samples.  A table of volumes for common numbers of tubes is given below:

|  | Number of tubes |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 | 2 | 3 | 5 | 7 | 8 | 10 | 11 | 12 |
| AM-A | 10 µl | 20 | 30 | 50 | 70 | 80 | 100 | 110 | 120 |
| AM-B | 35 µl | 70 | 105 | 175 | 245 | 280 | 350 | 385 | 420 |

- Prepare a master mix containing AM-A and AM-B according to the table above.

- Add 45 µl of the master mix to each PCR tube.

- Add 5 µl H2O to negative control tube.

- Disinfect all surfaces of PCR hood, and worktop, with 1% sodium hypochlorite solution.

- Remove lab coats and store in room.

- Bring PCR tubes out of room to separate area for addition of DNA to PCR tubes.

## 5.3. DNA Extraction using GenoLyse®

All steps to be performed in BSL 3 Laboratory.

- Thoroughly clean all surfaces of hood as well as external surfaces of all equipment in hood with appropriate concentration of mycobactericidal disinfectant followed by 1% bleach solution.

- Check temperature of the hot air oven or water bath.  Temperature must be monitored using a calibrated thermometer.

- For preparation of cleaning solutions:

| See: | Laboratory Cleaning and Maintenance_TB 02-04_VX.Y.doc |
| --- | --- |
| Location: |  |

### 5.3.1	DNA extraction directly from clinical specimens

- Sputum specimens should be decontaminated according to the Specimen Decontamination procedure using N-acetyl-L-cysteine-sodium hydroxide (NALC-NaOH).

- Smear positive and smear negative sputum specimens may be used with the version 2.0 assay.

| See: | Specimen Decontamination_TB 03-02_VX.Y.doc |
| --- | --- |
| Location: |  |

- Following suspension of the pellet in phosphate buffer, use a disposable Pasteur pipette to pipette 500μl of decontaminated sample to a 1.5ml microcentrifuge tube with screw cap.

- Centrifuge for 15 minutes at 10,000 x g in a centrifuge with aerosol tight rotor.

- Discard supernatant and suspend pellet in 100ul Lysis Buffer (A-LYS) using a vortex mixer.

- Incubate for 5 minutes at 95°C in a hot air oven or water bath.

- Briefly spin down.

- Add 100 µl Neutralisation Buffer (A-NB) and vortex the sample briefly (5 seconds).

- Centrifuge for 5 minutes at 10,000 x g.

- Tubes should be loaded in the centrifuge and unloaded in the same tube alignment to make visualization of pellet easier for removing the supernatant.

- Carefully carry the tubes to the BSC.  Uncap tubes one at a time, and carefully transfer 40-80ul of supernatant to a clean microcentrifuge tube using a P100 pipette (for storage).

- Careful attention must be paid to avoid touching the pipette tip to the pellet.  If this is done, and pellet is drawn into the pipette tip, a small amount of liquid can be expelled back into the tube until only clear supernatant remains.

- Add 5μl of DNA supernatant to corresponding PCR tubes in the PCR hood (if available).

### 5.3.2  DNA extraction from liquid culture

- Pipette 1 ml of liquid culture directly to microcentrifuge tube.

- Centrifuge for 15 minutes at 10,000 x g in a centrifuge with aerosol tight rotor.

- Discard supernatant and proceed as above, from addition of Lysis Buffer.

### 5.3.3  DNA extraction from solid culture

- Pipette 300μl molecular grade water into sufficient 200μl screw top tubes for 1 tube per culture.

- Use 1μl sterile disposable inoculation loop to collect bacteria from media with sufficient growth.

- Centrifuge for 15 minutes at 10,000 x g in a centrifuge with aerosol tight rotor.

- Discard supernatant and proceed as above, from addition of Lysis Buffer.

## 5.4	Amplification

All steps take place in Post-Amplification Room.

- If PCR tubes have bubbles at base, remove by swinging arm with tubes in hand in arc.

- Transfer PCR tubes to middle section of thermocycler.

- For first use, set up the thermocycler to the correct amplification profiles according to the instruction manual.  For thermocyclers provided by Hain Lifescience, the PCR protocols are pre-installed (MDR CUL for cultures and MDR DIR for clinical specimens)

- Check program parameters and follow menu options to start appropriate program.

Amplification Profile:

Culture samples	Clinical specimens

15 min 95°C	1 cycle			1 cycle

30 sec 95°C	10 cycles		20 cycles

2 min 65°C

25 sec 95°C

40 sec 50°C	20 cycles		30 cycles

40 sec 70°C

8 min 70°C	1 cycle			1 cycle

Note: the PCR protocols are different from the ones used in GenoType MTBDRplus version 1.0.

## 5.5	Hybridization

All steps to be performed in Post-Amplification Room wearing gloves.

- Pre-warm HYB and STR solutions (green and red) to 45°C in water bath (15 minutes total).

- Pre-warm TwinCubator to 45°C (tray must be in H20 to at least 1/3 its height).

- Pipette 20 ul DEN (denaturing solution) to each well of tray to be used.

- For each repetitive addition step, may use same pipette tip if wells/samples not touched.  If the wells are touched a fresh pipette tip must be used

- Add 20 ul of corresponding amplified DNA sample to each well, and mix well by pipetting up and down several times.

- Incubate for 5 minutes at room temperature.

- Remove DNA strips from tube (shake strips down to end of tube then remove carefully holding the end of the strip with forceps) and mark them with provided red pen or pencil.

- Add 1 ml HYB (hybridization solution) to each well and gently shake to homogenize solution.

- Add 1 strip to each well with colored marker facing up.  If strips turn over, re-position them with a fresh pipette tip.

- Place tray on TwinCubator and press “START” to incubate for 30 minutes at 45°C. From this point, press right arrow on TwinCubator once to advance steps in protocol.

- When alarm goes off, press right arrow to stop.

- Pour off HYB carefully or pipette with individual sterile pipettes into a small plastic discard container containing undiluted bleach solution. Wipe off condensation that forms on lid before every incubation step.

- Add 1 ml STR (stringent buffer) per well and incubate for 15 minutes in TwinCubator at 45°C.  Press right arrow to start.

- Prepare diluted Conjugate and Substrate in 15 ml conical vials by diluting 1:100 with corresponding Con-D and Sub-D. Colours of small tubes (concentrates) correspond to colors of dilution buffer tubes.  Wrap Substrate dilution in aluminum foil.  Prepare fresh Conjugate and Substrate dilutions every day. When alarm goes off, press right arrow. Completely remove STR as previously described for HYB removal.

- Add 1 ml RIN (rinse solution) per well.   Press right arrow to incubate for 1 minute on TwinCubator..

- Hit right arrow after alarm. Remove RIN and add 1 ml of diluted Conjugate per well.

- Press right arrow to incubate for 30 minutes on TwinCubator.

- Rinse sink with 1% sodium hypochlorite solution.

- Press right arrow upon alarm. Remove solution and wash for 1 minute with 1 ml RIN per well on TwinCubator.  Pour out solution and repeat rinse with 1 ml RIN per well for 1 min.

- Remove RIN and wash with 1 ml distilled water per well on TwinCubator.

- Remove water and add 1 ml of diluted substrate per well.

- Place on TwinCubator under aluminum foil for a maximum of 10 minutes.  Look for colour reaction to indicate reaction completion after 4-5 minutes.  If colour reaction is too weak, replace the foil and re-incubate for several more minutes, up to a maximum of 10 minutes.

- Wash twice for 1 minute each with distilled water.  Do not remove the volume of water as this will assist transfer of strips to result sheet.

- Trays can be reused a few times. Wash in water and repeat rinse in distilled water.  Occasionally wash in 1% sodium dodecyl sulphate solution.

- Record cleaning of laboratories and equipment in Laboratory Cleaning and Maintenance Logbook.

| Use: | Laboratory Cleaning and Maintenance Logbook_form.doc |
| --- | --- |
| Location: |  |

## 5.6.	Interpretation of Results

- Use forceps to transfer strips to the GenoType MTBDRplus Results Sheet provided with the kit or downloadable from Hain Lifescience website.

| Use: | Genotype MTBDRplus Results Sheet.pdf |
| --- | --- |
| Location: |  |

- Read results by lining strips up to code provided with kit.

- Attach the strips to the results sheet using clear adhesive tape.  Store the strips in a folder away from light.

- In order for results to be valid, CC (conjugate control) and AC (amplification control) bands must appear for every sample.  The presence of TUB band indicates that M.tuberculosis complex is present in the sample.

- A mutation in the relevant gene (and resistance to the relevant drug) is signified by either an absent wild type band and/or the presence of a mutant band for each gene cluster.

- The rpoB, katG and inhA each have a control band which must be present in order to interpret the results.

- rpoB predicts RIF resistance, katG predicts high level INH resistance, inhA predicts low level INH resistance.

- For results to be valid the bands must be of intensity approximately equal to or greater than the intensity of the AC band.

- In order for a batch of results to be valid, the negative control strip must have a CC and AC band present, but no other bands must be visible.

- If a positive result is obtained with the negative control, the results of the whole batch must be repeated and measures taken to remove amplicon contamination from all rooms and equipment.

- Refer to the product insert for interpretation of banding patterns and troubleshooting.

# REFERENCES

GenoLyse VER1.0.  IFU-51610-08.  10/2011. Hain Lifescience GmbH.

GenoType MTBDRplus  VER 2.0. IFU-304A-01.  Hain Lifescience GmbH.

Barnard M, Albert H, Coetzee G, O'Brien R, Bosman ME. Rapid molecular screening for drug-resistant tuberculosis in a high-volume public health laboratory in South Africa. Am J Respir Crit Care Med. 2008 Apr 1;177(7):787-92. Epub 2008 Jan 17.

Ling DI, Zwerling AA, Pai M.  GenoType MTBDR assays for the diagnosis of multidrug-resistant tuberculosis: a meta-analysis.  Eur Respir J  2008; 32: 1165-1174.

- 7.		CHANGE HISTORY

| New version # / date | Old version # / date | No. of changes | Description of changes | Source of change request |
| --- | --- | --- | --- | --- |