## Slide 1: GenoType® MTBDRplus
Interpretation of results

## Slide 2: Interpretation of results
Align the bands Conjugate Control (CC) and Amplification Control (AC) on each strip with the respective lines on the worksheet
Tape each strip onto the corresponding line on the worksheet
Determine the band positivity and positions on each strip using the reference reading chart of the kit and mark the results on the worksheet
Bands should only be considered positive if they are approximately as strong as the AC control
Store GenoType worksheets together with strips; keep protected from light

## Slide 3
The Conjugate Control (CC) line must be present
If CC is negative, the conjugation or substrate reaction was unsuccessful, either due to error in the procedure or due to problems with the reagent

## Slide 4
The Amplification Control (AC) line detects an internal control in PNM that is specific for all currently known mycobacteria as well as Gram+ bacteria of high G+C content. Therefore, this control is not mycobacteria specific!
If AC is positive, errors during extraction and amplification set-up and presence of amplification inhibitors in the specimen can be excluded
Signal of AC can be weak or even absent while results for other bands (TUB, rpoB, katG and inhA locus controls) may be positive. This might be due to competitive reactions between AC and TUB, ropB, katG, inhA during amplification. In this case, the strip can be evaluated.

## Slide 5
A weak or missing AC band with negative test result for TUB, rpoB, katG and inhA locus controls may indicate potential mistakes during extraction and amplification set-up, or the presence of amplification inhibitors; in this case, the test has to be invalidated
All bands (except of CC) should be compared to the AC control band for density
Bands that are less dense than the AC band should not be reported
The AC band always has to be positive even in negative controls, indicating that the amplification reaction was proper and not inhibited

## Slide 6
A positive M. tuberculosis control (TUB) band indicates the presence of members of the M. tuberculosis complex
If the TUB zone is negative, the tested bacterium does not belong to the M. tuberculosis complex; therefore, the presence or absence of any other bands (except CC and AC) cannot be considered for evaluation

## Slide 7
Locus control zones (rpoB, katG, inhA) detect a gene region specific for their respective genes
If the locus control zones are negative, then their respective mutation- specific positive bands cannot be considered for evaluation

## Slide 8
Prediction of susceptibility:
Only those bands whose intensities are about as strong as (or stronger than) that of the AC are to be considered
Wild type probes encompass common resistance-associated mutation sites of the respective genes
When all wild type probes of a gene stain are positive, there is no detectable mutation within the examined regions; the tested strain may be considered sensitive for the respective antibiotic

## Slide 9
Prediction of dug resistance:
In case of mutation, the respective amplicon cannot bind to the corresponding wild type capture probe on the strip due to the mismatch
The absence of a signal for at least one of the wild type probes may predict resistance to the respective antibiotic indirectly
Positive hybridization signal with a mutation-specific capture probe (for common mutations only!) may predict resistance to the respective antibiotic directly
Presence of rare mutations that do not have mutation-specific capture probes may only be indicated by the lack of hybridization with one or more wild type probes

## Slide 10: Rifampicin resistance: rpoB gene
rpoB wild type probes: WT 1 to WT 8
rpoB mutation-specific probes: MUT D516V, H526Y, H526D, S531L
Detection of mutations:
  absence of wild type signals
  presence of common mutation-specific signals

## Slide 11: High-level Isoniazid resistance: katG gene
Mutations in katG and the corresponding wild type and mutation probes
missing wild type probe analysed codon mutation probe mutation
katG WT 315 katG MUT1 S315T1
katG MUT2 S315T2

## Slide 12: Low-level Isoniazid resistance: inhA gene
Mutations in the inhA promotor region and the corresponding wild type and mutation probes

## Slide 13: Mixed pattern
Hetero-resistance = equal representation of susceptible and resistant mutants of the same strain
Mixed pattern = mutual presence of a resistant strain and a second, susceptible strain
Not pure culture
Carry-over contamination

## Slide 14: Summary (1)
All bands (except CC) should be compared to the AC control band for density
If the TUB zone is negative, the tested bacteria does not belong to the M. tuberculosis complex and presence or absence of any other bands (except CC and AC) cannot be considered for evaluation
If the locus control zones are negative, the respective mutation-specific positive bands cannot be considered for evaluation
Absence of signal with wild type probes may predict resistance, while positive hybridization signal with a mutation-specific capture probe (for common mutations only!) may predict resistance
Presence of rare mutations may only be indicated by the lack of hybridization with one or more wild type probes

## Slide 15: Summary (2)
Compare all molecular results with other laboratory results, i.e. culture, identification, conventional drug susceptibility testing and, if possible, with clinical findings

## Slide 16: Acknowledgement
<number>