## Slide 1: Module PRINCIPLES OF RECHECKING AND SAMPLE SIZE DETERMINATION

## Slide 2: Content Overview
Background of rechecking
LOT Quality Assurance Sampling (LQAS) and sample size determination

## Slide 3: Aim Of Rechecking
Improve AFB microscopy quality
  Screening for substandard performance and corrective action
  Continuous motivation
Not correction of diagnosis
High quality of AFB microscopy essential for patient management
  Missed or delayed diagnosis: death, advanced, disabling disease
  Incorrect AFB pos: side-effects of drugs; social, economic consequences
Notes: STRESS that high quality AFB microscopy services are highly important for the management of patients. The principal aim of rechecking is to improve the quality of AFB microscopy services by identification of microscopy centers with unacceptable performance.

## Slide 4: Principles Of Rechecking (1)
Random, representative sample of slides from each laboratory
Rechecked blindly at higher level- first controller
Discordant slides rechecked by second controller (gold standard)
Errors assigned to laboratory and first controller
Notes: Customisation required: specify here the actual designations and level of first- and second-level controllers.
EXPLAIN the principles of rechecking in a simple way: a random and representative number of slides from every laboratory are rechecked blindly (without knowledge of the peripheral result) at a higher level. Discordant smears are again checked by a second controller. Errors can thus be assigned to the laboratory or to the first controller, and the work of the first controller can also be evaluated.

## Slide 5
Analysis to identify centres possibly performing below standard
Confirmation by other checks in time and on-site visits
Identification of possible sources of errors
Principles Of Rechecking (1)
Notes: Stress that not all labs that have errors may perform below standard, so an in-depth analysis has to be made according to pre-set criteria to identify those which are more likely to do so (because there may be a problem with the microscope, stains or staining, reading etc..) Then a selection has to be made of the labs that need a problem-identification and –solving visit by an expert. These will be the labs with the poorest performance. The others with only slightly too many errors may not have a real problem, or only a minor one such as slightly superficial reading. These labs will be followed-up by further rechecking, and visits will be made only if the excessive number of errors persists. The aim of the visits must be to probe into the causes of the errors, and it should result in solving them so as to prevent further errors.

## Slide 6: Rechecking Merits
Final aim: improve quality
  Screening for unsatisfactory performance
  Problem solving
Control of routine smears: real situation
Continuous motivation and learning

## Slide 7: Rechecking Disadvantages
Workload, especially national and intermediate level
Only possible with good TB programme
  Regular supervision
  Highly functional laboratory network
  Corrective measures
Needs considerable inputs
  Personnel
  Materials
  Funds

## Slide 8: Inputs for Rechecking (1)
Personnel
  Supervisors for sample collection and feedback
  Rechecking coordinators at intermediate levels
  National level coordinator
  First level controllers
    Sufficient time
    Motivated
    Seniority not required
  Second level controllers
    Experienced, dedicated
    Sufficient time
    High level qualifications/position not required
Notes: Customisation required: specify here the actual designations and level of supervisors, first- and second-level controllers and coordinators.

## Slide 9: Inputs For Rechecking (2)
Materials and funds
  Microscopy centre
    Slide boxes
  Supervisors collecting the slides
    Slide boxes, sampling forms, transport
  Controllers
    TIME, staining facilities/equipment, transport
  TB coordinators
    Forms, tools for analysis (computer), transport
Notes: Customisation required: specify here the actual designations and level of supervisors and coordinators.

## Slide 10: Sampling Principles
As few slides as possible
  overload is counterproductive
  balance between statistical and technical accuracy
Notes: Stress that often striving at statistical accuracy makes correct technical execution of rechecking impossible due to overload. Statistics on incorrect numbers are of course useless, so it is better to limit the sample size to what can be done by recheckers with good quality.

## Slide 11: Sample Size Concepts (1)
Representative and random selection of any slide, irrespective of result and type (diagnostic, follow-up smear)
Statistically accurate evaluation is limited
  reproducibility limitations of scanty results
Notes: EXPLAIN that smears with scanty results (1-9 AFB / 100 HPF) have a special status. It is both easy to miss them and difficult to confirm and thus statistically accurate evaluation is limited.

## Slide 12: Sample Size Concepts (2)
Sampling method for rechecking: LQAS = LOT Quality Assurance Sampling
One-sided test: no exact results, wide confidence limits, except for bigger area, i.e. country;
    Example: 3% false negatives in a sample of 90 slides per centre
    if for 1 centre: 0.8 to 9.1% confidence limits
    if average for 12 centres: 2.1 to 4.2 %
    if average for 850 centres: 2.9 to 3.1%
Notes: STATE the Lot Quality Assurance Sampling (LQAS) is a statistically based sampling method used for rechecking.
EXPLAIN that LQAS is a one-sided test with wide confidence limits except for large areas. If e.g. 3% false negatives are found in a sample of 90 slides per microscopy centre, the confidence limits are 0.8%-9.1% if used for one centre, 2.1%-4.2% if used as the average of 12 centers and 2.9% to 3.1% if used as the average of 850 centers.

## Slide 13: Sample Size Concepts (3)
LQAS
  Smears reported negative by microscopy centre
  Allows smallest possible sample
  Statistics to calculate sample size not applied for positive smears
  Assurance of quality of majority of centres
  No excess errors in sample- no more FN than allowed
  Excess errors- further checks needed

## Slide 14: Sample Size Concepts (4)
False negatives (FN)
  Few unavoidable
  Check underlying problem if too many or if H(igh)FN
  HFN and LFN counted together= total FN
False positives (FP)
  Should not occur
  No statistics needed to decide on significance
  Systematic errors or real problem?
  Ignore occasional LFP
Notes: MAKE SURE participants understand abbreviations used (HFN, LFN, etc.)

## Slide 15
How to determine significance of FN?
  If not more than a set number of false negatives are found in the sample, the laboratory performs satisfactorily
  If more than a set number of false negaties are found in the sample,
    No assurance of satisfactory performance
    Does not mean bad performance
    Further investigation needed
Sample Size Concepts (4)
Notes: Any technician, however good, will from time to time make a false negative, and these low rates must be distinguished from an excessive rate due to underlying problems. This needs a statistical system.
This is not so for false positives (except low scanty, reproducibility problem): clear false positives should not happen at all, and are always significant. No statistics are needed. And we are only interested in more frequent false positives due to a problem that must be solved: these can be found easily from small samples.
For convenience and direct comparability, positives and scanty smears are sampled at their frequency in each lab register, by selecting one random sample regardless of the result recorded. Sample sizes needed for negatives have been inflated to accomodate these positive/scanty smears.

## Slide 16: Sample Size System
LQAS
  sample sizes taken from tables, according to certain parameters
  positives/scanty added to sample of negatives proportionally to their frequency in laboratory registers
Notes: Explain that the LQAS sample sizes have been determined by statisticians beforehand, and that we can take these from published tables. Although these sample sizes have been calculated targeting only false negative errors, the sample sizes shown in the microscopy EQA handbook tables contain also a number of positive / scanty proportional to their occurrence in the lab registers. Stress that only one sample of the chosen size needs to be taken.

## Slide 17: LQAS Parameters (1)
Parameters to determine sample size
  The LOT (N) = total number of negative slides registered during 1 year
  The acceptance number (d) = maximum number of FN errors allowed in the sample before action taking
  The critical value of FN = maximum accepted % of FN
Notes: EXPLAIN that a number of parameters are used to determine a sample size based on LQAS method.

## Slide 18
  Critical value of FN:
    Not straightforward parameter
    Varies with prevalence of smear positives
    Centers with higher prevalence of positive are allowed more FN for same sensitivity
LQAS Parameters (2)
Notes: Explain that FN is not an easy parameter since its normal value will vary with positivity rate (or prevalence of positive smears) in the labs. It will thus need to be calculated based on this positives prevalence and a critical value for what is considered to be acceptable sensitivity. Stress that sensitivity (and specificity) are objective parameters, not influenced by positives prevalence.

## Slide 19: Influence Of Smear Positive Prevalence On False Negatives (1)
Notes: This is an example of how positives prevalence influences greatly the FN percentage even when sensitivity remains the same. Here it is done via calculations in cross-tables, on the left for a lab with 10% positives, on the right another lab with 5% positives. Sensitivity and specificity are the same, 80% and 100%, but the calculations show that FN is more than double with 10% positives than with 5%.
How the calculations are done: start from total slides (1000) in the bottom right corners, and apply the prevalence of positives (10% or 5% of 1000), which yields the numbers 100 respectively 50 at the right outer border of the cross-tables. Calculate the difference from 1000 to find the negatives in the lab register. Since there are no false positives (specificity = 100%), a zero can be filled in the upper right cell within the borders of the cross-tables (= false positives). The difference with total positives in the register can then be calculated, i.e. 100 on the left and 50 on the right. These values (100, resp. 50) constitute the positives detected at sensitivity 80%, so we can calculate total positives dividing by 80, then multiplying by 100. This way we find the numbers 125 (left table) and 62.5 (right table) you see at the left on the outer bottom border of the tables. The numbers 875 and 937.5 have then been obtained by subtracting these total positives from total slides 1000.
Finally, false negatives (controlled negative / controller positive results within the cross-tables) is calculated subtracting the true positives (100 or 50) from the total positives (125 or 62.5), yielding 25 or 12.5. Percent false negatives is then found dividing these numbers by total negatives in the register, so right outer border of the cross-tables (900 or 950), yielding these 2.8% and 1.3%

## Slide 20: Influence of a 10% and 20% Prevalence on FN
Notes: This diagram illustrates the same influence of positives prevalence as the cross-tables in previous slide. Both the left and right parts contain each 100 squares: the blue ones are true negatives, the red ones true positives, and the green ones are false negatives. There are twice the number of positives (red) on the right, but also twice the number of false negatives (green). So the sensitivity remains the same (9/10 or 18/20 were detected). But not the percent FN, expressed as false negative divided by all negatives reported (so 91 on the left but 82 on the right).

## Slide 21: LQAS Parameters (3)
  Critical value of FN replaced by two other parameters on which it depends
    Side positivity rate (SPR)= percentage of positive slides out of total slides (diagnostic and follow-up) examined during one year
    Relative sensitivity= the ability of the technician to detect AFB relative to controllers

## Slide 22: Sample Size Determination (1)
LOT and SPR calculated from laboratory registers
d and relative sensitivity decided by National TB Reference Laboratory/NTP management; strongly recommended to keep:
  d=0 and
  Relative sensitivity at 75% or 80%
Other parameters:
  Specificity: set at 100%: no false positives are allowed
  Confidence level: sample sizes have been set to determine if the centre has met expected sensitivity within 95% confidence level
Notes: Customisation: specify here the values of the parameters (d, relative sensitivity) set by your country.
Stress that to make technically correct rechecking possible, sample sizes must be kept low. For this reason d is always set at zero, and the critical value of relative sensitivity is kept low. Explain that this does not mean that in routine work no false negative errors are allowed at all. And also that this low critical relative sensitivity is not to be seen as a target for performance, but just as a warning level. Even with this small sample size, every laboratory performing below this level will be considered as having too many errors.

## Slide 23: Sample Size Determination Example (1)
Step 1 | Laboratory | Slides/yr | Pos/Yr | Neg/Yr
Make a list of the laboratories in your country (or region in large countries( with the following information: number of slides done per year number of positive slides per year number of negative slides per year | A B C D E F G H I J Total | 1 500 2 550 1 990 2 085 900 1 158 1 250 885 2 569 500 15 387 | 200 351 156 151 85 100 125 101 335 55 1 659 | 1 300 2 199 1 834 1 934 815 1 058 1 125 784 2 234 445 13 728

## Slide 24
Procedure | Example |  |  | 
Step 2 | Laboratory | Slides/yr | Pos/yr | Neg/yr
Calculate the slide positivity rate (SPR) in each laboratory and round off to the nearest %. This is best done using Laboratory Register data from the previous year or 3 quarters of one year and one quarter previous year. Both diagnostic and follow-up slides should be included. | A B C D E F G H I J Total | 1 500 2 550 1 990 2 085 900 1 158 1 250 885 2 569 500 15 387 | 200 351 156 151 85 100 125 101 335 55 1 659 | 1 300 2 199 1 834 1 934 815 1 058 1 125 784 2 234 445 13 728
Sample Size Determination Example (2)

## Slide 25
Procedure | Example
Step 3 | Average SPR= (1659/15387) x 100 = 10.8% or 10% (rounded off)
Calculate the average SPR for your country (or region) and round off to the nearest % Average SPR= total positive slides/total number of slides) x 100 | 
Sample Size Determination Example (3)

## Slide 26
Procedure | Example
Step 4 | Average negative slides = 13728/10 = 1373 or 1000 (rounded off) Note: because the sample size does not vary considerably for annual workloads above 1000, rounding off will not affect calculation
Calculate the average annual number of negative slides per laboratory, round off to nearest 1000 for averages above 1000 and nearest 100 for averages below 1000 Average negative slides = total number of the negative slides /number of laboratories | 
Sample Size Determination Example (4)

## Slide 27
Procedure | Example
Step 5 | If: - Chosen sensitivity is 80% and d=0 - Average Number of Negative Slides = 1000 - Average SPR = 10% The Sample Size = 96
Select appropriate sample size from tables in Appendix D2 of “External Quality Assessment for AFB Smear Microscopy”. Heading of the tables: the sensitivity relative to the controllers First column: average numbers of negatives examined annually per centre Second column: d Remaining columns: SPR Locate sample size by selecting the sensitivity first; then d, followed by the number of negatives. Finally locate sample size under corresponding SPR | 
Sample Size Determination Example (5)

## Slide 28
Procedures | Example
Step 6 | 96/4 = 24 slides to be collected every quarter
Decide on a convenient interval to select the slides Recommended quarterly Divide the required sample size by the interval to calculate the number of slides to be collected at every interval | 
Sample Size Determination Example (6)

## Slide 29
LQAS sample size table (simplified) Sensitivity (relative to controller)= 80%
    Specificity=100%
Acceptance number d=0
Number of Slide positivity rate
negative
slides/yr 5% 10% 15% 20% 25% 30%
107 72 54 43 36 30
154 89 62 48 39 31
180 96 66 49 40 33
208 103 69 50 40 33
50000 216 104 69 51 40 33
Sample Size Determination Example (7)
Notes: Stress that the sample size only slightly increases with increased workload above 1000 negative slides per year. The more important variable is the SPR.

## Slide 30
Alternative
  Decide on workable number of slides per centre for rechecking
  Calculate average number of negative slides per centre and SPR
  Find the sensitivity in the tables which gives the workable number of slides
  Keep d=0
  No fixed rules: sensitivity may be adjusted, keep
  number of slides to workable minimum
  Justify what has been decided and implemented
Sample Size Determination (2)

## Slide 31
Aim is fixed sample size for all centres
  sample size varies very little with the number of negative slides; this variation is thus disregarded
  If major fluctuation in SPR a few sample sizes may be selected
  If extreme variability, a sample size per laboratory may be considered
  The more sample sizes, the more complex field operations and the higher the risk of mistakes
Sample Size Determination (3)

## Slide 32
Compared with old rechecking method- all positives and 10% of negatives- the sample size of LQAS method is much smaller
Comparison of Sample Sizes: Old and LQAS Method
Notes: STATE this slide shows a comparison of sample sizes using the in the past frequently applied sampling of all positives +10% of negatives, versus LQAS samples sizes taken from the LQAS sample size tables.
STRESS that LQAS samples are smaller except if positivity rates are very low or laboratories have very low workload. All three laboratories in this example have either high positivity rates (first one) or high workload (other two).

## Slide 33: Stratified Sampling (1)
Solution for medium prevalence situations
  SPR around 5%, when non-stratified sampling leads to overload
  Cannot be used when almost no positives are detected
Notes: “Stratified sampling” is the term used for sampling a certain number of slides first from among suspect smears, then from follow-up smears (or reverse).

## Slide 34
Rationale stratified sampling:
  Follow-up smears from initial smear positives contain, on average, 10% of positive/scanty results.
    Most at end intensive phase of treatment
    % depends on prevalence MDR-TB, frequency follow-up smears, prevalence HIV
  Follow-up smears are relatively difficult,- low positive and scanty, with lower reproducibility than diagnostic smears
Stratified Sampling (2)

## Slide 35
  Minimal acceptable sensitivity can be set lower than for all smears; 65% recommended
  If d is set at 0, the annual sample size of follow-up smears is 40
  50-60 slides will be rechecked per laboratory per year; all results together
    40 randomly selected follow-up smears
    10-20 randomly selected smears of suspects
Stratified Sampling (3)

## Slide 36: Stratified Sampling (4)
If <40 follow-up smears per year and <3% SPR, disregard rechecking
Instead regular panel testing and intensified supervision

## Slide 37: Stratified Sampling (4)
Refined analysis is still possible
Greater difficulty of follow-up smears is accounted for in the analysis
  sensitivity of first controller calculated on rechecking sample
  performance of each laboratory compared to controller’s performance
Notes: Explain that objective and comprehensive analysis is possible along the principles explained in the module: “Reporting and analysis of AFB-smear laboratories performance, interpretation and feedback”; in the part on the use of the Workbook.

## Slide 38: Conduct Of Rechecking (1)
Requirements for correct rechecking results
  Blinding of first controller to results of the laboratories
  Second control (countercheck) on discordants
  Restaining of slides (some or all)
  Validation of results by comparing error-rates of the controllers with average for the laboratories
Notes: Explain that here just the requirements are listed. Details are discussed in the module: Rechecking procedures”.

## Slide 39: Conduct Of Rechecking (2)
Collaboration between different levels
  Supervisors
  Intermediate level rechecking coordinator
  High level coordinator(s)
    level and number depends on size of the country, population, number of administrative areas, means of communication, experience with rechecking, availability of staff at intermediate level
Notes: CUSTOMISE this slide by stating the exact designation of the staff and level of the service responsible for various tasks in your program.
STATE implementation of a rechecking program requires collaboration between different levels. Some functions of different levels can be combined.

## Slide 40: Conduct Of Rechecking (3)
Supervisors
  Select samples for rechecking and deliver these to intermediate coordinator
  Give feedback of rechecking to microscopy centres
  Identify possible problems and their solutions, including requirement for a visit by a laboratory professional, for laboratories with unsatisfactory results
Notes: Explain that the supervisor can be a NTP supervisor, a TB supervisor or a general supervisor, depending on the situation in the country. It will usually not be a laboratory professional. The last bullet implies that the supervisors can identify and solve minor practical problems, such as poor staining technique or poor microscope and can identify problems that require a visit by a laboratory professional.

## Slide 41: Conduct Of Rechecking (4)
Intermediate level rechecking coordinator (1)
    Forwards slides and forms with slide identification numbers to first controllers
    Compares results of first controller with those of laboratories
    Forwards discordant slides to second controller
    Gives feedback to laboratories and first controllers
Notes: CUSTOMISE this slide by stating the exact designation of the staff and level of the service responsible for various tasks in your program.
LIST responsibilities of the intermediate level rechecking coordinator as stated on the slide.

## Slide 42: Conduct Of Rechecking (5)
Intermediate level rechecking coordinator (2)
    Requests supervisors to identify possible problems and give possible solutions for laboratories with unsatisfactory results, takes steps for remedial action, if appropriate
    Requests immediate problem solving by laboratory professional to centres with extremely bad results
    Prepares quarterly / annual reports
    Sends quarterly / annual reports to higher level(s)
Notes: CUSTOMISE this slide by stating the exact designation of the staff and level of the service responsible for various tasks in your program.
CONTINUE listing responsibilities of the intermediate level rechecking coordinator as stated on the slide.

## Slide 43: Conduct Of Rechecking (6)
High level coordinator
  Compiles quarterly / annual reports
  Facilitates remedial actions: replacement of poor microscopes, visits by laboratory professional, additional training, etc
Notes: CUSTOMISE this slide by stating the exact designation of the staff and level of the service responsible for various tasks in your program.
LIST responsibilities of the high level coordinator as stated on the slide.

## Slide 44: Key Messages
The purpose of rechecking is to screen for possible problems in AFB microscopy based on a representative and random sample of smears examined in routine work
Rechecking requires a substantial investment of financial, material and human resources, besides smooth collaboration at all levels
LQAS allows the smallest possible sample size to assess the quality of routine work
Stratified sampling could be a possible solution for situations where the smear positivity rates are low, around 5%.