RECHECKING - QUALITY ASSURANCE COORDINATOR AND RECHECKER LEVEL

________________________________________________________________________

BACKGROUND AND PRINCIPLES OF RECHECKING

___________________________________________________________________

| Purpose: | To provide participants with an understanding of the background and principles of rechecking |
| --- | --- |
| Learning Objectives: | At the end of the module, the participant will be able to: - Explain the purpose of rechecking - Explain the principles of rechecking - List prerequisites for rechecking - Define sample size |
| Content Outline: | - Background of rechecking - Principles of rechecking |
| Appendix: | 1. Task analysis for the different levels of the microscopy network |

BACKGROUND OF RECHECKING

The purpose of rechecking is to screen for possible problems in AFB microscopy based on a sample of smears examined in routine work. The sample should be representative and random. Ideally it should allow for evaluating the work of the individual technician over a defined period, and the sample should be sufficiently large to attain statistical significance of the results. However, in practice this may not be possible, because of the danger of overburdening the recheckers (controllers). Therefore sampling is usually done per center over a longer period, i.e. one year, and using a statistical method that will allow for identification of centers that may be performing below the minimum level targeted by the NTP and thus in need of further action. Rechecking is not meant to confirm the diagnosis of patients and is also not a substitute for internal quality control and regular supervision. If properly executed, rechecking is capable of gradually improving the quality of the AFB microscopy network.

We always have to keep in mind that a high quality AFB microscopy service is very important for the management of patients. If the diagnosis is missed because the smears are incorrectly reported as negative, the so called false-negative (FN), the patient will not receive anti-TB treatment. This patient may die or be diagnosed later with advanced disease which, even if treated, may seriously affect his/her quality of life. If a patient is incorrectly diagnosed as smear-positive, the so called false-positive (FP), they will unnecessarily be treated with drugs that may cause side-effects. In addition, there may be social and economic consequences for the patient.

- A rechecking system is considered to be the best method for evaluating performance because it reveals the reality of the daily performance of the laboratory. Experience shows that blinded rechecking can be highly motivating for laboratory staff, and can be seen as a method of continuous education. Properly executed, rechecking can gradually bring the quality of the network to a high level.

However, this EQA method is the most resource intense and the most expensive to implement, and constitutes a great deal of work for intermediate and central level. The workload is proportional to the number of centers to be controlled, and inversely proportional to the prevalence of positive smears in the laboratories. Thus the more centers a controller has to recheck and the lower the TB prevalence, the higher the workload. A highly decentralized AFB microscopy service or a very low TB prevalence will make rechecking impossible.

PRINCIPLES OF RECHECKING

The principle of blinded rechecking is rather simple: a random and representative number of slides from every laboratory are rechecked at a higher level by a rechecker/controller who does not know the original results of the laboratory. Discordant smears, (positive at the laboratory and negative by the controller, or the reverse) are rechecked by a second controller who serves as the gold standard. This will result in errors assigned to the laboratory as well as to the first controller, so that controllers can also be evaluated and controls are validated. Analysis is meant to identify centers that may be performing below the standard. The sample of smears rechecked is relatively small, so when poor performance is detected it must be validated by other checks, preferably an on-site visit. The on-site supervision visit should identify possible sources of errors which require remedial action.

Prerequisites for rechecking are:

- a well functioning TB control program with regular supervision,

- a highly functional TB laboratory network,

- the potential to implement corrective measures when unsatisfactory laboratory performance is identified, and

- available resources to support the work, including personnel, materials and funds.

Resources needed:

Personnel:

- Supervisors for sample collection and routine feedback. These do not need to be laboratory technicians; they are usually the regular non-laboratory NTP (district) supervisors.

- Rechecking coordinators at the intermediate levels. This can be the non-laboratory district or regional TB coordinator.

- First level controllers, who can be junior but motivated laboratory staff. The required number depends on the number of centers, the sample of slides for rechecking per center/per year, and whether or not the controllers are full time or part time assigned to rechecking.

- Second level controllers, can be staff at the central or at the (same) intermediate laboratories. Controllers should be experienced and motivated staff with adequate time to perform the work, but they do not need to be high level supervisors. They need plenty of time per slide to ensure they provide the most correct result, so they may only be able to process 5-10 discordant slides (if part-time) or 15 discordant slides (if full-time) per day. Based on experiences in several countries, the number of discordant slides is usually higher during the initial phase of EQA implementation and initially requires more second level controller time per unit of population.

Materials and funds:

- For the laboratories:

  - Sufficient slide boxes to keep all slides for at least two quarters

  - Tools for permanent identification of slides: either pencils for marking frosted slides or diamond markers

- For the supervisors collecting the slides:

  - Small slide boxes (20 slots) for slide transport; 1-2 per laboratory supervised

  - Forms for sampling

  - Transport

- For the controllers

  - Xylene to clean slides

  - Staining materials, equipment and facilities

  - Transport

- For the TB coordinators

  - Forms for listing discordant results

  - Forms for sending routine feedback back to the microscopy center (if different from the sampling form)

  - Forms for reporting upward to TB Program

  - Tools for analysis: either a standard spreadsheet format or computer forms if data management is computerized

  - Transport

The task analysis for the different levels of the microscopy network is presented in Appendix 1.

SAMPLE SIZE DETERMINATION

If the sample is to be truly representative, it must be possible to select any examined slide, whatever the result and whatever the type (suspect first spot specimen, subsequent suspect specimens, follow-up specimen).

The sample size is a fixed number of slides, chosen for the country or region/province, based on a statistical system called the Lot Quality Assurance Sampling (LQAS). It is not a percentage of the slides examined. If conditions allow, the sample size is kept uniform meaning the same number of slides will be selected from each laboratory, irrespective the number of slides examined annually in a center. In case of wide variations between either the workload of the laboratories, or the percentage of positive slides, more than one sample size of slides may be needed.

The LQAS system allows for the smallest possible sample to be used to gain the information most needed to assure the quality of the majority of the laboratories.

Using this statistical system, the annual sample size is usually between 50 and 90 slides rechecked per laboratory. The head of the national TB reference laboratory, together with NTP management, will decide on the sample size based on certain parameters. One of the parameters used to determine the sample size is the maximum accepted number of false negative (FN) slides in the sample. This is usually set at 0, meaning that no FN is allowed in the sample. If no FN is found in the sample, we can conclude that the quality of the laboratory is acceptable. If one or more FN is found in the sample, we cannot be sure whether the quality is acceptable or not; in these cases more checks are needed. Clear cut false positives (FP) 1+ to 3+ are also not allowed; scanty FP are usually ignored, as long as they occur only occasionally.

Using a strict statistical method, which would require rechecking huge numbers of slides for absolute accuracy, is not practical. Therefore, when using the LQAS method, further evaluation is always required to determine if the errors are random or represent a performance problem.

In comparison with the rechecking method that was recommended in the past (rechecking of 100% of the positive slides and 10% of the negative slides), the sample size using LQAS is much smaller. This is shown in the following example:

| Annual volume of slides | Positive slides | Negative slides | Old rechecking method | New (LQAS) rechecking method |
| --- | --- | --- | --- | --- |
| 520 | 74 | 456 | 120 | 62 |
| 2556 | 182 | 2374 | 459 | 135 |
| 6822 | 355 | 6467 | 1001 | 208 |

After the annual sample size per center has been determined, a convenient interval has to be chosen for visiting the centers to select the slides. The usual interval is quarterly. If 96 slides have to be selected annually, this means that 96/4= 24 are selected quarterly. If the sample size is 77 annually, each quarter 77/4= 19.25 slides, rounded off to 20 slides, should be selected or three times 19 slides and one time 20 slides. In cases where visits for selection of slides are made three times annually, the sample size has to be divided by three. If all of the visits for selection of slides cannot be done as frequently as originally determined, then the number of slides to be selected by the supervisor during the last visit of the year or the first visit of the next year should be the annual sample size minus the number of slides already collected. The slides should be selected from those examined since the last visit until the end of the year.

Appendix 1

TASK ANALYSIS FOR DIFFERENT LEVELS OF THE MICROSCOPY NETWORK

National level

- Training of laboratory staff in AFB microscopy and EQA

- Culture and sensitivity testing

- EQA, rechecking second control level (in smaller countries with good communication)

- Final analysis and report of EQA results

- Supervision (on-site visits) for specific quality problems

- Requests for ordering of laboratory supplies to NTP management, distribution of supplies to intermediate levels

Intermediate level

- Culture and sensitivity testing, according to NTP policy

- Selection of slides samples in peripheral centers for rechecking

- EQA, rechecking first control level (and eventually also second level in larger countries)

- Analysis of EQA results, forwarding to national level

- Supervision of peripheral laboratories

- Requests for laboratory supplies to national level, distribution of laboratory supplies to peripheral levels

Peripheral level

- AFB microscopy

- Internal quality control of laboratory

- Keeping all slides

- Ordering of supplies

RECHECKING - QUALITY ASSURANCE COORDINATOR AND RECHECKER LEVEL

________________________________________________________________________

RECHECKING PROCEDURES

_______________________________________________________________________

| Purpose: | To provide participants with knowledge of the procedures for proper storage, collection and reading of slides. |
| --- | --- |
| Learning Objectives: | At the end of the module, the participant will be able to: - Describe how to keep the slides in the laboratory - Explain the process for collection of slides - List factors causing fading of fuchsin stain - Describe restaining policies - Describe the restaining technique - Explain blinded rechecking - Describe the process of blinded rechecking - Assess the smear quality |
| Content Outline: | - Keeping the slides in the laboratory - Collecting the slides - Restaining the slides - Reading the slides - Evaluating the smear quality |
| Exercises: | - Sample collection - Quality of ZN stained smears |

KEEPING THE SLIDES IN THE LABORATORY

In order to perform rechecking, all slides examined at the center have to be kept in slide boxes until the sample of slides to be rechecked has been collected. To keep the slides, the laboratory staff should take the following steps:

- Keep the slides from all patients, irrespective their result, until the sample has been taken.

- Label the slides with the laboratory register number, including the result column number 1, 2 or 3 for the diagnostic slides; the result should not be written on the slide.

- After reading the slides, let the oil soak into absorbent paper until the end of the day or next morning, putting the slides smear downwards on a strip of toilet paper or newspaper. Do not try to remove remaining oil by xylene or any other solvent and never rub the smear.

- Arrange the slides in the boxes by number, leaving empty slots for those expected to arrive over the next few days (2nd and 3rd diagnostic smear). If there is a shortage of boxes, roll slides in toilet paper taking care that they do not touch each other and store them in their original carton packs.

- Start with a new box at the beginning of each sampling period (quarterly, every four months, etc.).

- Each center should have sufficient boxes to store slides examined during 2 sampling periods (i.e., 2 x 3 months if samples are taken quarterly). For example, if the quarterly visit of a supervisor takes place on 28 November, during that visit the supervisor will select the sample from the slides examined during the period 1 July-30 September (3rd quarter of the year). The laboratory has stored the slides for that period in boxes, and, starting with new boxes, also for the period 1 October-28 November. After the sample has been taken, the remaining slides for the quarter 1 July-30 September can be discarded and these boxes will be filled with the slides for the period 1 January-30 March of the following year.

COLLECTING THE SLIDES

Slides for the rechecking sample should be selected by a visiting supervisor. This does not need to be a laboratory specialist; usually it will be the supervisor who also supervises all TB related activities in a facility. During the visit to collect the sample, the supervisor should provide feedback on the results of the previous slides rechecked. This feedback will help to motivate laboratory staff to keep all slides and improve performance if necessary. If sampling is not possible during a supervisory visit, all slides examined since the previous visit, with a copy of the laboratory register, should be forwarded and sampling is then done by the rechecking coordinator. To be certain that a proper, representative sample is collected, the sample should only be selected by a trained supervisor or rechecking coordinator, the laboratory staff of the peripheral center should never be asked to select and send the sample.

When collecting the sample, follow these steps:

- During each visit, collect the fraction of the annual sample covering the period since last visit (i.e., one quarter every 3 months, or, one half year if the last quarterly visit did not take place).

- Mark the period in the laboratory register: 1 January - 31 March, 1 April - 30 September, etc, and draw a line under the last entry for that period.

- Count the total number of smears examined during this period. It may be very helpful if the laboratory staff counts the number of smears routinely and enters these in the laboratory register at the end of each quarter (or month).

- Calculate the sampling step (interval) by dividing the total number of smears by the sample size required to be selected, rounding down to the next integer. For example: 675 smears examined and 24 slides to be selected gives a sampling step of 675/24= 28, meaning every 28th slide should be collected.

- Enter the required information on the “blinded rechecking form of sputum smear examinations for acid-fast bacilli”, including identification of the laboratory, period covered by the sample, and date of sample. If possible use carbon paper to complete a second form at the same time, placing it only under the identification but not under the results column, to obtain one copy with results and one without. If no carbon paper is available, fill out two forms, one with and one without result.

- Use a random number (such as last digit of a bank note) to select the first slide of the sample, counting from the first result in the laboratory register and going left to right over the columns, then downwards over the lines. If the random number is 5, the first slide of the sample will be the 5th slide in the laboratory register, the second slide the 33rd slide (5+28) and so on until the required number of slides has been reached. Note the identification numbers of the slides on both forms and the results on only one form.

- Ask the laboratory staff to retrieve the selected slides from the boxes. Check these against the list. Mark any missing slide on the form and replace the slide with the next slide in the laboratory register. Ensure that each selected slide is labeled with the laboratory code.

- If no slides with a positive or scanty AFB result are selected, add the first two positive or scanty slides found in the laboratory register searching from the start of the period. These are additional slides and should not replace already selected slides.

- Ensure that the remaining slides not selected are discarded.

- Arrange the slides in a transport box in the same order as listed on the form.  The slides should be delivered as soon as possible to the first controller, together with the copy of the form without the results of the microscopy center. The form with the results of the microscopy center should be given to the rechecking coordinator.

RESTAINING THE SLIDES

Ziehl Neelsen (ZN) stained slides

Poor staining at the laboratories may result in AFB not being visible, resulting in FN when read by the technician, and also undetected when rechecked by the controller. Poorly stained AFB may also be missed by the controller when fuchsin stain fades, which may occur slowly in high temperature and high humidity climates. If AFB were weakly stained, the bacilli may loose all color and quickly become invisible. As a result of fading, positives reported by the laboratories cannot be confirmed by the controllers. In this case, there may be false accusations of FP.

It is impossible to predict where and when fading will occur to such an extent that the results of rechecking will become unreliable or even completely wrong.

This graph shows the amount of fuchsin fading at 35 degrees C and different relative humidity (RH).

- The following result from Burundi, a country with a mild climate, shows the profound effect that restaining can have on rechecking results. False positives disappear while false negatives become apparent, and therefore required action is totally different.

Restaining Ziehl Neelsen (ZN) stained slides

Based on the conditions in the country and practical experience with rechecking, the head of the national TB reference laboratory, in consultation with NTP management, will decide whether and at what level the slides selected for rechecking will be restained. Restaining of fuschin stained slides may be performed on:

- all discordant slides at the second control level, or

- all slides selected for rechecking at the first control level.

Restaining Auramine stained slides (Fluorescence Microscopy, FM)

All auramine stained slides should be restained with auramine-phenol stain at the first control level. This recommendation is justified based on country experiences, anecdotal evidence, and theoretical considerations. To make unbiased, blinded rechecking possible, auramine-stained fluorescence positives should be confirmed by changing to higher magnification when needed, but not after restaining with ZN.

Restaining technique

The slides should first be cleaned in xylene. There is no need to decolorize. After cleaning, the slides should be restained with ZN or auramine, using the standard staining procedure.

READING THE SLIDES

Blinded rechecking

First level rechecking should be completely blinded, meaning the first controller should have no access to the results of the centers. The first controller should receive only the slides and a copy of the rechecking form with the identification numbers without the results.

Process of blinded rechecking

First controller:

- Clean all oil-immersion examined slides with xylene.

- Inspect smears with naked eye, record an assessment of size, thickness, evenness, and background color on the form in the column “Remarks”. If slides are not restained, a record of AFB color and artifacts in the background will be added after microscopic examination.  If all slides are routinely restained at the first controller’s level, make a note on the form that restaining was done.

- Restain all auramine stained slides, and all ZN stained slides only if this is NTP policy.

- Read the smears using the appropriate microscopes, bright field for ZN, fluorescence for auramine, and standard magnification.

- Check the routine number of fields, usually 100 fields or one length for a negative smear. Do not check more fields, unless there is a question about the quantification after reading one length.

- Record the first control results on the copy of the rechecking form in the column “First Controls Result”.

- Add an overall appreciation of the quality of smearing (restained smears) or smearing and staining (non-restained smears) at the bottom of the form.

- Do not recheck slides with unclear identification or severely damaged smears; report these as “excluded/ID” or “excluded/damaged”

- Let the oil soak onto absorbent paper (toilet paper, newspaper) as for routine work; replace the slides in the original box in the proper order.

- Keep the box in a refrigerator if the climate is likely to cause fading.

- Send the forms to the rechecking coordinator.

EVALUATING THE SMEAR QUALITY

Smear quality:

The material should have been spread equally over an area of 2 cm by 3 cm or 2 cm by 1 cm. The quality of sputum is good if mucus threads with white blood cells are present; the quality of the smear is poor if the smear is thin with epithelial cells. However, follow–up smears may be thin and not contain white blood cells. Therefore, remarks on poor quality of sputum and smears should only be made if too many thin smears without white blood cells are seen.

Staining quality: A well stained ZN smear shows strong red AFB against weak blue background. A well stained auramine smear shows yellowish fluorescence rods against dark background. Moreover, the background should have very little remaining red in ZN, or fluorescence in auramine stained smears. There should not be any red artifacts in ZN stained smears, and only few fluorescing artifacts in auramine stained smears.

RECHECKING - QUALITY ASSURANCE COORDINATOR AND RECHECKER LEVEL

________________________________________________________________________

IDENTIFICATION AND MANAGEMENT OF DISCORDANT SLIDES

_______________________________________________________________________

| Purpose | To provide participants with an understanding of the procedures for identification and management of discordant slides, including validation of the results and identification and allocation of errors. |
| --- | --- |
| Learning Objectives | At the end of the module, the participant will be able to: - Define discordant slides - Describe the principle of a second control level - Describe the process for managing discordant results - Explain the need for validation of rechecking results - Describe the process for validation of rechecking results - Identify and allocate errors |
| Content outline | - Identification of discordants - Management of discordants - Validation of rechecking results - Identification and allocation of errors |
| Exercises | - Discordants and error identification and validation |

IDENTIFICATION OF DISCORDANTS

To resolve discrepancies, discordant results between the first controller and the laboratory should be counterchecked by a second controller. First controllers can make mistakes, so results have to be validated to ensure the overall reliability of rechecking. Discordant slides are slides called positive (scanty, 1+, 2+ or 3+) by the laboratory and negative by the first controller or the reverse. Quantification errors (QE) are a difference in positive grading of more than one step: scanty versus 2+ or 3+ and 1+ versus 3+.

MANAGEMENT OF DISCORDANTS

Discordant slides (LFN, HFN, LFP, HFP and QE), must be counterchecked by a second controller. This person must be an experienced technician with sufficient time and sense of commitment to yield the most correct result possible. The second controller can be stationed at a higher level laboratory, provided that communication with the first controlling level is easy and fast. Otherwise, it may be more efficient if second controller is stationed at the same level as the first controller, but this obviously will require more discipline to guarantee objective results. Usually this requires a third person, the rechecking coordinator.

The second controller will know both results, but not which result is from the laboratory and which result is from the first controller. Assuming that AFB is present, second level rechecking may require a prolonged search to find the bacilli. Both reported results are used to guide the number of fields to be checked.

- If the discrepancy is negative versus clear positive result (1+, 2+, 3+), two lengths (200 fields) of the slide will be sufficient before concluding that the negative result is correct.

- For scanty versus negative results, 5 lengths of the slide are to be examined before declaring that the negative result is correct.

- For quantification differences examine as many fields as needed to feel confident in making the final judgment.

Positive smears should also be quantified as to the number of AFB seen, e.g. 98/200, 20/500.

Process for managing discordant results

Rechecking coordinator

- Copy the results of the first controller to the original rechecking form with the results of the laboratory, and the results from the laboratory onto the copy from the first controller.

- Identify discordant results between the laboratory and the first controller. Using the discordant rechecking form, list all of the discordants with the name of the laboratory.  Use column “Result 1” for some entries of the result of the laboratory and others for entries of the result of the first controller and reversing for column “Result 2” so that the second controller does not know where each result was produced.

- Request the discordant slides from all laboratories on the discordant form from the first controller to be packed in a separate box.

- Send (give) this box, together with the discordant form, to the second controller.

- At this stage, do not give any feedback to the laboratories with discrepant slides, since errors have not yet been verified. Feedback can be given to a laboratory if no discrepant results were identified in the entire sample.

Second controller

- Restain all discordant ZN and auramine stained slides, if not done previously or if not protected against fading since restaining.

- Recheck the smears, allowing for sufficient time to carefully verify results.

- Record the quantified results on the discordant form in the column “Result” for the second controller; add qualitative remarks about smear quality (using same criteria as specified for the first controller).

- Make one copy for the file; send the original, together with the slides, back to the rechecking coordinator.

Rechecking coordinator

- Copy the results of the second controller to the column “Second Controls Result” on both copies of the laboratory rechecking form.

- Using these results as the gold standard, determine a) whether the error was made by the laboratory or the first controller and b) the type of error, according to the table on classification and definition of errors.

- Complete the cross-tables on the bottom of the rechecking form with numbers of slides rechecked and errors identified, both for the microscopy centers and the first controllers.

VALIDATION OF RECHECKING RESULTS

- It is important to validate the results of rechecking prior to sending reports to higher level and to centers. Final analysis and interpretation of rechecking results depends on proper validation, otherwise conclusions may be completely wrong.  Occasionally, controllers may not perform very well. This can be detected in several ways:

- There may not be any discordants identified. This is unexpected and highly unlikely when controller rechecks hundreds of smears originating from centers where positives are regularly encountered. This suggests that controllers somehow knew the result from the peripheral laboratory (results were not blinded).

- The first controller may have more false-negatives than all his/her centers combined together from one same series, i.e. one quarter.  In extreme cases, this means that the controller's examination has been so superficial that it must be assumed that false-negatives of the microscopy centers were not detected.  This would invalidate the results. However, this validation can only made if all smears were restained prior to reading by the first controller, and when the proportion of positive/scanty smears in the total sample was the same as the average SPR in routine.

- Together with few or no false-negatives, superficial controls will result in false declaration of false-positives committed by laboratories. If second controller counterchecks are done well, this will be recognized as FN of the first controller and corrected. If the second controller is also superficial, this may not be corrected and then centers would show many FP, while FN by laboratories as well as first controllers may be rare.

- If FP are detected in many centers, this should always raise suspicion about the quality of the controls. However, this may also be the result of rechecking without restaining. HFP, in particular, should seldom be a problem with trained people and reasonably good microscopes.

- Under some conditions (rapid fading, staining problems), validation can only be done fairly when all smears were restained prior to reading by the first controller. Otherwise, the first controller will have made many false negative errors, simply because the AFB were not visible.

- Validation of the rechecking results may also include feedback from microscopists and supervisors. When smears with serious errors are returned to the center, microscopists will be able to show supervisors the positives missed by both controllers.

IDENTIFICATION AND ALLOCATION OF ERRORS

- After validation of the rechecking results, the errors will be defined according to type, as specified in the table on classification and definition of errors. The result being controlled are the slides with discordance between the laboratory and first controller, and the second controllers’ result will then serve as the final result or gold standard, to decide who made an error and which type of error was made. Classification and definition of errors is specified in the table below.

- Identification and allocation of errors should be done after rechecking of each sample of slides (i.e. quarterly). This will allow for early investigation and corrective action, especially in the case of serious problems. It will also provide the opportunity to monitor whether centers with problems improve their performance with the next round of rechecking.

RECHECKING - QUALITY ASSURANCE COORDINATOR AND RECHECKER LEVEL

________________________________________________________________________

REPORTING AND ANALYSIS OF AFB-SMEAR LABORATORIES PERFORMANCE, INTERPRETATION OF RESULTS AND FEEDBACK

_______________________________________________________________________

| Purpose: | To provide participants with an understanding of the process for reporting and analyzing AFB-laboratories performance, interpreting results and providing feedback. |
| --- | --- |
| Learning Objectives: | At the end of the module, the participant will be able to: - Report results of AFB-smear laboratories performance - Analyze rechecking results - Explain validation of rechecking results - Interpret rechecking results - Describe the purposes and methods of feedback |
| Content Outline: | - Reporting of AFB-smear laboratories performance - Validation of rechecking results - Analysis of rechecking results - Interpretation of rechecking results - Examples of analysis and interpretation of rechecking reports |
| Exercises: | - Exercise on Reporting of AFB lab performance, Manual system |
| Appendices: | Reporting and analysis documents |

REPORTING OF AFB-SMEAR LABORATORIES PERFORMANCE

Reporting and analysis of rechecking results can be done using a manual system, or together with a refined analysis by means of a workbook which requires computers for recording the data. Regardless of which system is used, data should be compiled and analyzed on a quarterly and annual basis. In addition, reports of laboratory performance should be included that provide essential information on workload and other parameters related to the quality of work. These reports are needed independent of rechecking reports. Analysis of rechecking results using the workbook is only possible with the supplemental data provided in the quarterly/annual report on AFB laboratory performance.

The following documents are used for developing reports on the performance of AFB-smear laboratories:

Quarterly Report “AFB laboratory performance and stock of consumables”: for reporting by the peripheral laboratories (Appendix 1-1).

Quarterly Report “AFB laboratory performance”: for the intermediary level coordinator, to report consolidated performance of the peripheral laboratories upward to higher/national level (Appendix 1-2).

Form 1 “Blinded Rechecking of Sputum Smear Examinations for Acid-Fast Bacilli”: used by the visiting supervisor to register slides sampled for rechecking, and used to provide rechecking feedback to the microscopy center (Appendix 2-1).

Form 2 “Rechecking of Sputum Smears for AFB, List of Discordants”: for use by the rechecking coordinator at the intermediary level to list discordant smears and corresponding results, this form is sent to the second controller to record second controller’s results (Appendix 2-2).

Form 3 “Rechecking of Sputum Smears for AFB, Quarterly / Annual Report Form”: for use by the rechecking coordinator at the intermediary level, to manually compile the rechecking data and forward these upward to higher/national level (Appendix 2-3).

The workbook is initially for use at the national level, eventually it may be used at the intermediate level. The advantage of the workbook is that data is compiled and analyzed in a consistent manner provided the data is entered correctly. It will automatically make the additions quarterly and annually, by laboratory and by first controller. Another advantage is that it will automatically correct for inaccuracies, such as those related to a uniform sample size, and over-sampling of positives.

Use of the laboratory performance documents

The Quarterly Report “AFB laboratory performance and stock of consumables” should be completed quarterly by each peripheral laboratory and sent to the district/region/ province. It can also be completed by a visiting supervisor with help from the laboratory staff. Laboratory staff should, at minimum, count and record the numbers of smears by type of result (positives, scanty, negative) each month, and check and record inventory of the remaining stock quarterly.

The information on smears performed will be used to compile a consolidated performance report for the district/region/province… that will sent to the higher level. For this purpose the individual reports are either summarized on the quarterly report form “AFB laboratory performance”, either directly or in the Workbook. The following information needs to be recorded:

- name and sequence number of the laboratory,

- technique used: Ziehl-Neelsen or Fluorescence Microscopy,

- totals of AFB smears examined during the quarter/year: from the performance reports of the laboratories, totals of positive, scanty 1-9 and negative smears reported, by type of examination (diagnostic or follow-up),

- stocks remaining for the essential consumables.

The stock supply inventory is not reported to higher up, but used to estimate the needs of essential supplies at the intermediary level, for automatic supply to the laboratories. For this purpose a simple spreadsheet should be used. These stock reports are NOT used to calculate the supplies that need to be ordered from higher up, instead the number of smears examined and the number of functional laboratories will be used for this purpose.

Use of the rechecking forms

Form 1 is used by the supervisor collecting the sample to list the selected slides and corresponding results on one copy to be kept by the EQA rechecking coordinator; only slide numbers are listed on another copy which goes to the first controller together with the slides.  Upon receipt of the copy with first controller results, the coordinator will fill in the column of missing results on both copies, and identify discordant results.

Form 2 is used by the rechecking coordinator to list slides with discordant results, together with both results (peripheral and first controller).  Discordants from several laboratories can be entered on the same form, as long as the laboratory is also identified on the form. To ensure that the second controller does not know where each result was produced (laboratory or controller), it is imperative that the columns used for the results from the laboratory and the first controller are switched regularly and randomly. For example, enter laboratory results for a few smears using column 1, and column 2 for the first controller results, and listing laboratory results of other smears in column 2 and first controller results in column 1.

Form 2 with the corresponding discordant slides goes to the second controller, who will add her/his results in the third column, and return the form to the coordinator.

Form 1- the coordinator adds the second controller results to both copies of Form 1, and completes the boxes at the bottom of the form, entering:

  - the numbers of slides checked by result, for both laboratory and first controller: these numbers are obtained by counts made in the respective columns (“Peripheral lab result” and “Result first control”); the number will regularly be slightly different between laboratory and first controller, but their totals should be the same.

  - numbers of errors made by both laboratory and first controller in the boxes below the totals rechecked, classified by error type (HFP, HFN, etc.)

Form 1 is then completed by adding recommendations at the bottom, depending on the findings (errors; qualitative remarks on smear and staining quality). One copy of Form 1 is used to give feedback to the laboratory; the second one is filed by the coordinator and will be used to compile the quarterly report (Form 3).

Form 3 is used by the coordinator to compile rechecking results for all the laboratories controlled in a specified area (district, region, province). Only the summary counts are listed, not individual slides. Alternately Form 3 can be omitted if information from Form 1 is entered directly into the workbook (computer). The workbook is then used to report to the higher level. The following information is entered on Form 3:

  - names of the peripheral laboratories and a sequential number (left-hand columns);

  - peripheral labs: total numbers of positive, scanty 1-9 and negative smears rechecked (from Form 1, totals by result for the peripheral lab) and errors found (from Form 1, summary of errors detected for the peripheral lab);

  - names or codes of the first controllers, by laboratory rechecked;

  - first controller(s) summary counts per laboratory rechecked: total numbers of positive, scanty 1-9 and negative results reported for smears rechecked and errors committed by the first controller (from Form 1, as for the peripheral labs);

  - totals for each column are calculated and entered at the bottom.

Form 3 is then sent to the higher level each quarter.

VALIDATION OF RECHECKING RESULTS

To check the reliability of the first controller, compare the sum of FN of all laboratories under the first controller with the number of FN by that controller. However, this validation can be misleading in the following situations:

- when positive smears were added to the random samples, resulting in a higher SPR in the samples rechecked as compared with the routine. In this case the controllers are at a disadvantage and must be allowed more FN. (When the workbook is used, the computer analysis will take into account differences in SPR from routine performance data).

- in situations when fading happens rapidly due to climate conditions or poor staining, and only discordants were stained prior to second controls. The first controllers will not have been able to detect some of the true positives in smears that had faded.

ANALYSIS OF RECHECKING RESULTS

For each sample series rechecked, errors will be counted by type and by center. Final analysis can only be completed after the full year’s sample has been processed. However, interim results with general comments should be reported as feedback as early as possible to motivate laboratory staff to continue collection and storage of slides and make any needed improvements in their performance. Sometimes the first series of a sample will show problems that need to be investigated as soon as possible.

It is inappropriate to count all results together, ending with a single score of agreement or discordance, particularly when this includes qualitative aspects such as smear or stain quality. This approach would cause the loss of valuable detailed information and the aim of EQA is not to score laboratories but to detect specific performance problems.

- Although error frequencies can be calculated, they will not be accurate (wide confidence limits) due to the small sample size, except when considering all centers from a large area together. Frequencies for the laboratories can be calculated by type, using as the denominator all slides in the sample with this result as reported by laboratories. For instance, in the case of HFP, all reportedly positive or scanty slides rechecked will be the denominator. Similar for FN (LFN and HFN), all reported negative slides will be used as the denominator.

- The same calculations will be made for the first controllers, but their denominators may be slightly different. For instance, smears reported positive by the laboratories, but found negative by first controller will be included in the denominator for calculating false-negatives, and not for the false-positives. For validation of rechecking results, comparing FN errors made by the sum of laboratories with those made by their controller should always be done as a check on the controller’s reliability.

- INTERPRETATION OF RECHECKING RESULTS

- With zero tolerance parameter (d=0) for FN in the sample, not a single false negative should  have been found in the annual sample. FP are never allowable. In principle, finding just one error means that the laboratory cannot be assumed to work at the minimum acceptable quality level. This does not necessarily mean that the quality is bad, only that further checks are needed to exclude poor quality. However, at this point more refinements are possible, and some interpretation is indicated.

In practice it is highly unlikely that there be no errors at all in large samples. If no discordant smears are found in an annual sample of laboratories covered by a first controller and thus the results by the controller are similar to those of the laboratories, an inquiry should be made to determine if the rechecking was really blinded.

- Interpretation of the results should consider the limitations of the system.  Most of the centers with slightly too many errors may not have a real problem. Since the sample has to be small, chance findings are expected. For this reason, priority for corrective action must be given to centers clearly surpassing the limits: more than one HFP or HFN in a small sample, or many LFN or LFP. Otherwise, further routine rechecking will gradually make it clear if there is a real problem in the laboratory. In case of doubt as to whether there is a real problem or just chance errors, it is best not to criticize the staff. Instead, show them the results and encourage them to improve on qualitative deficiencies.  Subsequent series will clarify if the error was a chance finding or not.  Since LFP may be due to a limitation of the control, they can usually be ignored unless they occur at high frequency compared to other centers and controller.  In this case, it is best to check a sample of several recent scanty smears on the spot, to determine if there is a systematic error.

The NTP management, in collaboration with staff of the national reference laboratory, must decide when action should be taken in the case of HFP and FN results, and the level of false results, which action should be taken and by whom. These decisions should consider the local capacity for taking corrective action, including the availability of TB laboratory professionals.

EXAMPLES OF ANALYSIS AND INTERPRETATION OF RECHECKING RESULTS

1. Annual report of 7 laboratories

Below is an example of an annual report of laboratories under one first controller in a division. In the national TB control program, the acceptance number had been set at 0 and 60 randomly selected smears are rechecked per center annually.

FP: 2 centers (EF and FG) have HFP.  Center EF has 1 HFP and no LFP, the one HFP may be an administrative error, and therefore action can await further rechecking results.  Center FG has 4 HFP and 2 LFP, this center appears to have a serious problem, requiring immediate investigation.

FN: Five centers (BC, CD, DE, EF and GH) have HFN; all, except center BC, also with LFN. Priority investigation is required for center DE with 13 FN, and center GH having 7 FN.

Note: serious problems observed in the quarterly reports should be investigated as soon as they are detected, do not wait until the end of the year when the entire sample has been rechecked.

2.  Example of a manual analysis

Project: TTLCP								Year: 1999

| Center | Totals checked |  |  | Number of errors detected |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Pos | Scanty <4/100 | Neg | False positives HFP LFP |  | False negatives HFN LFN |  | QE |
| Madhupur | 66 | 19 | 183 | 0 | 3 | 1 | 1 | 0 |
| Gopalpur | 75 | 7 | 181 | 0 | 1 | 1 | 2 | 0 |
| Ghatail | 72 | 6 | 180 | 0 | 0 | 1 | 3 | 0 |
| Kalihati | 81 | 8 | 180 | 0 | 0 | 0 | 0 | 0 |
| Tangail S | 77 | 4 | 179 | 1 | 0 | 0 | 2 | 0 |
| Delduar | 64 | 3 | 164 | 0 | 0 | 0 | 1 | 0 |
| Mirzapur | 137 | 19 | 357 | 0 | 3 | 3 | 2 | 0 |
| Bhuapur | 75 | 5 | 180 | 0 | 0 | 2 | 6 | 0 |
| Nagarpur | 71 | 3 | 180 | 1 | 0 | 2 | 1 | 0 |
| Sakhipur | 72 | 12 | 179 | 0 | 0 | 1 | 1 | 1 |
| Basail | 70 | 8 | 180 | 0 | 0 | 0 | 1 | 0 |
| Sarishabari | 66 | 13 | 178 | 0 | 3 | 1 | 4 | 1 |
| Bokshigonj | 68 | 7 | 180 | 0 | 3 | 0 | 1 | 0 |
| Dewangonj | 76 | 9 | 180 | 0 | 1 | 2 | 1 | 1 |
| Melandaha | 76 | 7 | 178 | 0 | 0 | 1 | 4 | 0 |
| Madargonj | 65 | 9 | 180 | 0 | 2 | 1 | 1 | 0 |
| Jamalpur | 67 | 7 | 182 | 0 | 3 | 1 | 0 | 0 |
| Islampur | 73 | 14 | 181 | 0 | 0 | 5 | 3 | 0 |
| Totals | 1351 | 160 | 3402 | 2 | 19 | 22 | 34 | 3 |
|  |  |  |  | HFP | LFP | HFN | LFN | QE |
| Errors made by peripheral centers |  |  |  | 2 | 19 | 22 | 34 | 3 |
| Errors rates by center |  |  |  | 0.15% | 1.26% | 0.65% | 1.00% | 0.20% |
| Errors made by first controller |  |  |  | 1 | 22 | 18 | 41 | 1 |
| Error rates by controller |  |  |  | 0.07% | 1.45% | 0.53% | 1.21% | 0.07% |

The table above provides an example based on a sample size that allows a maximum of 4 FN, with over-sampling of positives. One HFP was found in each of two centers. As this finding is in a high number rechecked, it is unlikely that it is due to a problem that needs immediate investigation. LFP occur in eight centers, but the first controller has the same percentage of these errors, to they can be ignored. Adding the columns for HFN and LFN, five centers have more than 4 FN.  Mirzapur has 5, including 3 HFN, but the sample includes double the number of slides compared to other centers, so it may actually fall within normal limits. Of the other 4, only 2 centers, Bhuapur and Islampur, have considerably more than 4 FN, including more than 1 HFN. Priority should be given to investigating these centers for problems causing FN errors.  It is expected that a serious problem will be found in Islampur, considering the number of HFN in this center. Investigating problems in the other two, Sarishabari and Melandaha, may wait for further rechecking results as there was just one excess FN and most are LFN, it is unlikely that an investigation would discover a serious problem. First controllers have a slightly higher percentage of total FN errors than the total microscopy centers, 1.74% versus 1.65%. However, this is not due to their poorer performance, but to over-sampling of positives, leading to higher SPR and more FN allowed for the controllers. If relative sensitivities were calculated based on above numbers, it would show that the controller’s work was more sensitive. This is clearly one of the advantages of using the workbook.

FEEDBACK TO THE LABORATORIES

Regular feedback to the laboratories is very important for improving or maintaining motivation and improving performance, if applicable. Staff should not be criticized unless the problem is clearly caused by negligence or lack of will to do good work. In general, caution and exclusion of all possible sources of error is needed before jumping to such conclusions. Laboratory staff and controllers should be shown the results of rechecking and be given a copy of the sampling list with results of controllers and summary of errors.  For effective feedback, it is absolutely necessary to return slides with serious errors to the microscopy technician to allow them an opportunity to prove their results to a supervisor.  On-site supervision visits should be used to investigate all possible reasons for errors, and recommend corrective measures such as replacing a bad microscope, or improving smearing and staining techniques. Qualitative remarks on thickness and staining of smears from the controller will be helpful to guide the investigation. Feedback will also serve as continuous education, especially when slides with serious errors are returned.  A visit by a laboratory professional should be planned for centers with too many serious errors. These on-site visits are the best method for improving compliance and maintaining high motivation. Sending reports by mail or other means and requesting staff to improve will have little, if any, effect.

Appendix 1-1

Appendix 1-2

Appendix 2-1

Appendix 2-2

Appendix 2-3

Center = ABF microscopy lab, may be a peripheral or intermediate laboratory