RECHECKING – HIGH AND INTERMEDIATE  LEVEL

__________________________________________________

PRINCIPLES OF RECHECKING AND SAMPLE SIZE DETERMINATION _______________________________________________________________________

| Purpose: | To provide participants with an understanding of the principles of rechecking and sample size determination |
| --- | --- |
| Learning Objectives: | At the end of the module, the participant will be able to: - Explain the background of rechecking - Explain the prerequisites for rechecking - Describe the sample concepts - Describe the sample size parameters - Explain the influence of the prevalence of positive slides on the test performance - Explain determination of the sample size - List conditions for stratified sampling |
| Content Outline: | - Background of rechecking - LOT Quality Assurance Sampling (LQAS) and sample size determination |
| Exercises: | - Sample size determination |
| Appendix: | 1. Task analysis for the different levels of the microscopy network |

BACKGROUND OF RECHECKING

The purpose of rechecking is to screen for possible problems in AFB microscopy based on a sample of smears examined in routine work. The sample should be representative and random. Ideally it should allow for evaluating the work of the individual technician over a defined period, and the sample should be sufficiently large to attain statistical significance of the results. However, in practice this may not be possible, because of the danger of overburdening the recheckers (controllers). Therefore sampling is usually done per center over a longer period, i.e. one year, and using a statistical method that will allow for identification of centers that may be performing below the minimum level targeted by the NTP and thus in need of further action. Rechecking is not meant to confirm the diagnosis of patients and is also not a substitute for internal quality control and regular supervision. If properly executed, rechecking is capable of gradually improving the quality of the AFB microscopy network.

We always have to keep in mind that a high quality AFB microscopy service is very important for the management of patients. If the diagnosis is missed because the smears are incorrectly reported as negative, the so called false-negative (FN), the patient will not receive anti-TB treatment. This patient may die or be diagnosed later with advanced disease which, even if treated, may seriously affect his/her quality of life. If a patient is incorrectly diagnosed as smear-positive, the so called false-positive (FP), they will unnecessarily be treated with drugs that may cause side-effects. In addition, there may be social and economic consequences for the patient.

- A rechecking system is considered to be the best method for evaluating performance because it reveals the reality of the daily performance of the laboratory. Experience shows that blinded rechecking can be highly motivating for laboratory staff, and can be seen as a method of continuous education. Properly executed, rechecking can gradually bring the quality of the network to a high level.

However, this EQA method is the most resource intense and the most expensive to implement, and constitutes a great deal of work for intermediate and central level. The workload is proportional to the number of centers to be controlled, and inversely proportional to the prevalence of positive smears in the laboratories. Thus the more centers a controller has to recheck and the lower the TB prevalence, the higher the workload. A highly decentralized AFB microscopy service or a very low TB prevalence will make rechecking impossible (see below, under stratified sampling).  The same is true when statistical accuracy is over-valued, requiring a high number of slides for rechecking.

The principle of blinded rechecking is rather simple: a random and representative number of slides from every laboratory is rechecked at a higher level by a rechecker/controller who does not know the original results of the laboratory. Discordant smears, (positive at the laboratory and negative by the controller, or the reverse) are rechecked by a second controller who serves as the gold standard. This will result in errors assigned to the laboratory as well as to the first controller, so that controllers can also be evaluated and controls are validated. Analysis is meant to identify centers that may be performing below the standard. The sample of smears rechecked is relatively small, so when poor performance is detected it must be validated by other checks, preferably an on-site visit. The on-site supervision visit should identify possible sources of errors which require remedial action.

Prerequisites for rechecking are:

- a well functioning TB control program with regular supervision,

- a highly functional TB laboratory network,

- the potential to implement corrective measures when unsatisfactory laboratory performance is identified, and

- available resources to support the work, including personnel, materials and funds.

Resources needed:

Personnel:

- Supervisors for sample collection and routine feedback. These do not need to be laboratory technicians; they are usually the regular non-laboratory NTP (district) supervisors.

- Rechecking coordinators at the intermediate levels. This can be the non-laboratory district or regional TB coordinator.

- First level controllers, who can be junior but motivated laboratory staff. The required number depends on the number of centers, the sample of slides for rechecking per center/per year, and whether or not the controllers are full time or part time assigned to rechecking.

- Second level controllers, can be staff at central or at the (same) intermediate laboratories. Controllers should be experienced and motivated staff with adequate time to perform the work, but they do not need to be high level supervisors. They need plenty of time per slide to ensure they provide the most correct result, so they may only be able to process 5-10 discordant slides (if part-time) or 15 discordant slides (if full-time) per day. Based on experiences in several countries, the number of discordant slides is usually higher during the initial phase of EQA implementation and initially requires more second level controller time per unit of population.

Materials and funds:

- For the laboratories:

  - Sufficient slide boxes to keep all slides for at least two quarters

  - Tools for permanent identification of slides: either pencils for marking frosted slides or diamond markers

- For the supervisors collecting the slides:

  - Small slide boxes (20 slots) for slide transport; 1-2 per laboratory supervised

  - Forms for sampling

  - Transport

- For the controllers

  - Xylene to clean slides

  - Staining materials, equipment and facilities

  - Transport

- For the TB coordinators

  - Forms for listing discordant results

  - Forms for sending routine feedback back to the microscopy center (if different from the sampling form)

  - Forms for reporting upward to TB Program

  - Tools for analysis: either a standard spreadsheet format or computer forms if data management is computerized

  - Transport

The task analysis for the different levels of the microscopy network is presented in Appendix 1.

LOT QUALITY ASSURANCE SAMPLING (LQAS) AND SAMPLE SIZE DETERMINATION

Sample concepts

The sample should be representative and random. If the sample is to be truly representative, it must be possible to select any examined slide, whatever the result and whatever the type (suspect first spot, subsequent suspect specimens, follow-up specimen).

However, it is not practical or feasible to recheck sufficient slides to arrive at precise estimates of error rates for every center in the laboratory network. This would require large numbers of slides to be rechecked and controllers would be severely overloaded, so that their results would become unreliable. It is much better to select the smallest possible sample size and to execute rechecking technically as correctly as possible.

The Lot Quality Assurance Sampling (LQAS) is a statistically based sampling method used for rechecking. It is used for smears judged to be negative by the laboratory. The LQAS system allows for the smallest possible sample to be used to gain the information most needed to assure the quality of the majority of the laboratories. It will tell us that the centers without excess errors in the sample did not make more false-negative errors than the maximum allowed. For those centers with excess errors in the sample, this method tells us that there may be problems, but further evaluation is needed to determine if the errors actually represent a performance problem.  Although error rates could be calculated, these will not be accurate for individual centers, but only for groups of centers, such as those of a region, province or country. LQAS is a one sided test with wide confidence limits, except for large areas. For example, if 3% false negatives are found in a sample of 90 slides per microscopy center, the confidence limits are 0.8%-9.1% if used for one center, 2.1%-4.2% if used as the average of 12 centers and 2.9% to 3.1% if used as the average of 850 centers.

LQAS or other statistical method is not applied to calculate the sample size of the positive smears because clear-cut false-positives are never allowed. Therefore, every false-positive must be taken at face value without regard to statistical significance. High false-positives (HFP) are not accepted and need a further check of possible reasons for the error. Scanties (1-9 AFB’s per 100 fields), have a special status. Even for the most experienced technician they are easy to miss and difficult to confirm. Thus we cannot be absolutely sure which result is correct if a controller declares that no AFB were found in smears reported to have a scanty result by the center. Therefore, false-positives of the scanty level (low false–positives, or LFP) will be ignored as long as they occur only occasionally and in proportions similar to those of other centers and first-level controllers.

Similarly, there will be low false-negatives (LFN) and high false-negatives (HFN). These are counted together to arrive at the total number of false-negatives (FN). However, the type of FN error, LFN or HFN, will be considered for final evaluation and interpretation.

Quantification errors (QE) applies to a difference of more than 1 step in positivity result, (scanty versus 2+ and 3+, 1+ versus 3+), are not a high priority. Initially, quantification errors may be ignored in order to concentrate on resolving causes of major errors. However, they can sometimes help to understand the underlying cause of FN errors.

Sample size parameters

The LQAS system comes with several tables containing sample sizes required for certain values of the parameters used, which have been calculated by statisticians. The parameters needed to find the required sample size are:

1. The LOT (N) = the total number of negative slides examined during a specified period, usually one year.

2.  The acceptance number (d) = the maximum number of FN errors allowed in the sample before action is taken.

3. The critical value of FN = the maximum accepted % FN. This is replaced by the two parameters on which it depends:

- The slide positivity rate (SPR) = the percentage of positive slides out of the total number of slides (diagnostic plus follow-up slides) examined during a specified period, usually one year, and

- The sensitivity of smear microscopy= the ability of the technician to detect AFB relative to the controllers; a critical minimal value will be set, and the corresponding sample size will be sufficient to identify labs performing below this level.

The critical value of FN is not a straightforward parameter. It varies with prevalence of smear positives. If there are more positives, there will also be more FN for the equivalent and even better quality of work. This is shown in Appendix D2 of the guide: “External Quality Assessment of AFB Smear Microscopy” (EQA guide). This shows a table of calculated critical values of % FN for increasing sensitivity and prevalence of smear positive. Examples to show how critical values are calculated are also provided.

An example of the influence of the smear positive prevalence on FN is shown below. This shows that with equal sensitivity and specificity applied, there will be more FN when the prevalence of positives is 10% as compared with a prevalence of 5 %.

The sample size tables in the EQA guide, Appendix D, show total sample sizes, including positive and scanties. Positives and scanties have been added to the sample of negatives proportional to their occurrence in the laboratory register, thus based on the SPR. This process usually suffices because it usually doesn’t require many smear positives to detect FP errors that are due to an underlying problem that needs to be resolved. Generally these are systematic errors and not rare at all.

Other parameters used in determining the sample size are specificity and confidence level. The specificity is set at 100% because any false positive should trigger action.

Sample sizes have been developed to determine if the microscopy center has met the expected sensitivity within 95% confidence level.

The LOT and SPR should be calculated from the laboratory registers.  The selection of the acceptance number and the sensitivity to be used should be made by the chief of the national TB laboratory in consultation with NTP management.

The LOT and smear positivity rate (SPR)

To determine the LOT and the SPR:

- Prepare a list of all laboratories in the country. This list includes the name of the centers, the total number of slides examined during a year, the number of positive slides and the number of negative slides for each center. These lists are easy to make if laboratories report their performance regularly.  If performance reports from the centers are not available, the information can be obtained by the supervisors during their last supervision visit of the year, counting the slides examined during the first three quarters of that year plus those examined during the last quarter of the previous year. At the national level, the lists made by all supervisors is compiled, and this is used to calculate:

  - The average annual SPR for the country.  Avg SPR = the total number of positive slides/total number of all slides examined x 100%. This number should be rounded off to the nearest %.

  - The average annual number of negative slides per center.  Avg Neg = the total number of slides reported negative in all laboratories/total number of laboratories. This should be rounded off to the closest 100, for averages below 1000 and to the closest 1000, for averages above 1000.

Use of the recommended performance reporting formats provides for quarterly collection and annual compilation of numbers of positive, scanty and negative smears examined, by type of sputum (AFB Laboratory Performance, Quarterly/Annual report form).

The acceptance number and the sensitivity

The acceptance number is usually set at d=0, meaning no FN is allowed. This does not mean that there will be no FN in the LOT. When a sample of the required sample size is checked and no more than the specified number of FN errors is found, we can be statistically 95% assured that there are no more than the critical percentage of FN in the total number of slides. This percentage depends on set critical sensitivity and SPR, as explained.  If a program has been well established and sufficient resources are available, the acceptance number could be set at 1 or higher. However the required sample size will almost double with each unit increment of d, with only a modest improvement of statistical accuracy.

Sensitivity relates to the ability of the technicians to detect positive smears of any grade (scanty to 3+), in reality this may be close to 100% for high positives, but below 50% for low scanty results.  To select the sensitivity that will be used, the NTP should do one of the following:

- Decide first on an acceptable total workload and the sample size per laboratory, and then find the corresponding sensitivity in the tables provided.

- If workload is not a problem, decide first on the sensitivity, then find the corresponding sample size in the tables provided.

It is strongly recommended to start with a sensitivity target of not more than 75%. Laboratories working below this level have serious problems that need to be solved. These centers will be identified right away, together with some laboratories whose quality of work is marginal (borderline acceptability). Although sensitivity targets could be increased later on to detect centers with minor problems, particularly those where too little time is taken for reading smears, this is not recommended. It will increase sample sizes making detection of these errors technically less uncertain. It is better to ensure continuity and consistency of rechecking, in this way laboratories with less critical problems will be detected over time.

Sample size determination

After the SPR and the average number of negative slides examined annually per center have been calculated, and the acceptance number and the sensitivity have been determined, the required sample size can be taken from the tables in Appendix D of the EQA guide. When a sample of at least the required sample size is checked, and not more than the specified number of FN errors is found, we are statistically confident that the laboratory has met the sensitivity goal established by the program.

Using the sample size tables (EQA Guide Appendix D4)

The heading of the tables shows the sensitivity relative to the controller. The first column gives the average numbers of negatives examined annually per center, the second column the acceptance number (d) and the remaining columns the smear positivity rates.

For example, if the sensitivity is set at 80%, an average of 1000 negatives are examined per center annually, d is set at 0, and the positivity rate is 10%, the total annual sample for rechecking is 96 slides per center per year. If the positivity rate observed in the program is not listed in the table, simply take the sample size of the positivity rate closest to the observed rate, since calculation of an exact sample size is not essential.

Adapting the sample size to variations between laboratory workload or prevalence of smear positives should be done only in the case of major differences in SPR among the centers being controlled. Preferably not more than 2-3 different sample sizes should be set, or, in case of extreme variation in SPR, a small table of sample sizes should be constructed at national level. Variations in negative slides examined annually influence sample size so little that they can be disregarded; the average number can be applied to all centers.

In comparison with the rechecking method that was recommended in the past (rechecking of 100% of the positive slides and 10% of the negative slides), the sample size using LQAS. This is shown in the following example:

| Annual volume of slides | Positive slides | Negative slides | Old rechecking method | New rechecking method |
| --- | --- | --- | --- | --- |
| 520 | 74 | 456 | 120 | 62 |
| 2556 | 182 | 2374 | 459 | 135 |
| 6822 | 355 | 6467 | 1001 | 208 |

After the annual sample size per center has been determined, a convenient interval has to be chosen for visiting the centers to select the slides. The usual interval is quarterly. If 96 slides have to be selected annually, this means that 96/4= 24 are selected quarterly. If the sample size is 77 annually, each quarter 77/4= 19.25 slides, rounded off to 20 slides, should be selected or three times 19 slides and one time 20 slides. In cases where visits for selection of slides are made three times annually, the sample size has to be divided by three. If all of the visits for selection of slides cannot be done as frequently as originally determined, then the number of slides to be selected by the supervisor during the last visit of the year or the first visit of the next year should be the annual sample size minus the number of slides already collected. The slides should be selected from those examined since the last visit until the end of the year.

Stratified sampling

Stratified sampling is a potential option in area of moderate TB prevalence where the standard LQAS sample could overload controllers. Stratified sampling is applicable to areas where the smear positivity rates in the laboratory registers are low, around 5%.  However, it cannot be used when almost no positives are detected.

The rationale is that follow-up smears from initially smear-positive patients contain, on average 10% of the positive, including scanty, results. Most positive smears are found at the end of the intensive phase of the treatment. The percentage will vary with prevalence of MDR-TB, with relative frequency of intensive phase smears among total follow-up smears (i.e. done monthly in some countries), and possibly other factors, such as HIV prevalence. These are relatively difficult smears, comprised mainly of the whole range of scanty smears and 1+ smears, with lower reproducibility than the diagnostic smears, therefore the minimal acceptable sensitivity in this group can be set lower than that of all smears. The recommended sensitivity is 65%.  If d=0 (no false negative errors allowed), this gives a required annual sample size of only 40 slides from among the follow-up smears examined annually.

Using stratified sampling, 50-60 slides will be rechecked per laboratory per year, all results (positive, scanty, negative) together. These include 40 randomly selected follow-up smears per center per year. The 10-20 remaining smears should be taken from among suspect smears, randomly and with any result. In practice, the supervisor collecting the sample will start with the random identification of follow-up slides (for instance 10 for one quarter), applying the rules for sample collection described in the module “rechecking procedures” and only counting follow-up smears in the register. Once this is done, the same procedure will be repeated counting only the suspect smears, until their required number (for instance 5) has been reached. In case there were too few follow-up smears, the number of suspect smears should be inflated to accommodate for this deficit, as long as it is only a few slides.

In centers with less than 40 follow-up smears annually and less than 3% SPR, it is better not to attempt rechecking. Instead, it is recommended to concentrate on panel testing and intensified supervision.

There are no fixed rules. One may adjust the parameters, i.e. the sensitivity. Most important is to keep the number of slides for rechecking to a workable minimum and justify what has been decided and implemented.

Records must distinguish between the type of smears (suspect, smear-positive case, smear-negative case, follow-up).This information must then also be added to the sampling form.

Objective and comprehensive analysis is possible using the principles explained in the module: “reporting and analysis of AFB-smear laboratories performance, interpretation and feedback”.

Appendix 1

TASK ANALYSIS FOR DIFFERENT LEVELS OF THE MICROSCOPY NETWORK

National level

- Training of laboratory staff in AFB microscopy and EQA

- Culture and sensitivity testing

- EQA, rechecking second control level (in smaller countries with good communication)

- Final analysis and report of EQA results

- Supervision (on-site visits) for specific quality problems

- Requests for ordering of laboratory supplies to NTP management, distribution of supplies to intermediate levels

Intermediate level

- Culture and sensitivity testing, according to NTP policy

- Selection of slides samples in peripheral centers for rechecking

- EQA, rechecking first control level (and eventually also second level in larger countries)

- Analysis of EQA results, forwarding to national level

- Supervision of peripheral laboratories

- Requests for laboratory supplies to national level, distribution of laboratory supplies to peripheral levels

Peripheral level

- AFB microscopy

- Internal quality control of laboratory

- Keeping all slides

- Ordering of supplies

RECHECKING – HIGH AND INTERMEDIATE  LEVEL

__________________________________________________

RECHECKING PROCEDURES

_______________________________________________________________________

| Purpose: | To provide participants with knowledge of the procedures for proper storage, collection and reading of slides. |
| --- | --- |
| Learning Objectives: | At the end of the module, the participant will be able to: - Describe how to keep the slides in the laboratory - Explain the process for collection of slides - List factors causing fading of fuchsin stain - Describe restaining policies and factors for making the most appropriate choice - Describe the restaining technique - Explain blinded rechecking - Describe the process of blinded rechecking - Assess the smear quality |
| Content Outline: | - Keeping the slides in the laboratory - Collecting the slides - Restaining the slides - Reading the slides - Evaluating the smear quality |
| Exercises: | - Sample collection - Quality of ZN stained smears |

KEEPING THE SLIDES

In order to perform rechecking, all slides examined at the center have to be kept in slide boxes until the sample of slides to be rechecked has been collected. To keep the slides, the laboratory staff should take the following steps:

- Keep the slides from all patients, irrespective their result, until the sample has been taken.

- Label the slides with the laboratory register number, including the result column number 1, 2 or 3 for the diagnostic slides; the result should not be written on the slide.

- After reading the slides, let the oil soak into absorbent paper until the end of the day or next morning, putting the slides smear downwards on a strip of toilet paper or newspaper. Do not try to remove remaining oil by xylene or any other solvent and never rub the smear.

- Arrange the slides in the boxes by number, leaving empty slots for those expected to arrive over the next few days (2nd and 3rd diagnostic smear). If there is a shortage of boxes, roll slides in toilet paper taking care that they do not touch each other and store them in their original carton packs.

- Start with a new box at the beginning of each sampling period (quarterly, every four months, etc.).

- Each center should have sufficient boxes to store slides examined during 2 sampling periods (i.e., 2 x 3 months if samples are taken quarterly). For example, if the quarterly visit of a supervisor takes place on 28 November, during that visit the supervisor will select the sample from the slides examined during the period 1 July-30 September (3rd quarter of the year). The laboratory has stored the slides for that period in boxes, and, starting with new boxes, also for the period 1 October-28 November. After the sample has been taken, the remaining slides for the quarter 1 July-30 September can be discarded and these boxes will be filled with the slides for the period 1 January-30 March of the following year.

COLLECTING THE SLIDES

Slides for the rechecking sample should be selected by a visiting supervisor. This does not need to be a laboratory specialist; usually it will be the supervisor who also supervises all TB related activities in a facility. During the visit to collect the sample, the supervisor should provide feedback on the results of the previous slides rechecked. This feedback will help to motivate laboratory staff to keep all slides and improve performance if necessary. If sampling is not possible during a supervisory visit, all slides examined since the previous visit, with a copy of the laboratory register, should be forwarded and sampling is then done by the rechecking coordinator. To be certain that a proper, representative sample is collected, the sample should only be selected by a trained supervisor or rechecking coordinator, the laboratory staff of the peripheral center should never be asked to select and send the sample.

When collecting the sample, follow these steps:

- During each visit, collect the fraction of the annual sample covering the period since last visit (i.e., one quarter every 3 months, or, one half year if the last quarterly visit did not take place).

- Mark the period in the laboratory register: 1 January - 31 March, 1 April - 30 September, etc, and draw a line under the last entry for that period.

- Count the total number of smears examined during this period. It may be very helpful if the laboratory staff counts the number of smears routinely and enters these in the laboratory register at the end of each quarter (or month).

- Calculate the sampling step (interval) by dividing the total number of smears by the sample size required to be selected, rounding down to the next integer. For example: 675 smears examined and 24 slides to be selected gives a sampling step of 675/24= 28, meaning every 28th slide should be collected.

- Enter the required information on the “blinded rechecking form of sputum smear examinations for acid-fast bacilli”, including identification of the laboratory, period covered by the sample, and date of sample. If possible use carbon paper to complete a second form at the same time, placing it only under the identification but not under the results column, to obtain one copy with results and one without. If no carbon paper is available, fill out two forms, one with and one without result.

- Use a random number (such as last digit of a bank note) to select the first slide of the sample, counting from the first result in the laboratory register and going left to right over the columns, then downwards over the lines. If the random number is 5, the first slide of the sample will be the 5th slide in the laboratory register, the second slide the 33rd slide (5+28) and so on until the required number of slides has been reached. Note the identification numbers of the slides on both forms and the results on only one form.

- Ask the laboratory staff to retrieve the selected slides from the boxes. Check these against the list. Mark any missing slide on the form and replace the slide with the next slide in the laboratory register. Ensure that each selected slide is labeled with the laboratory code.

- If no slides with a positive or scanty AFB result are selected, add the first two positive or scanty slides found in the laboratory register searching from the start of the period. These are additional slides and should not replace already selected slides.

- Ensure that the remaining slides not selected are discarded.

- Arrange the slides in a transport box in the same order as listed on the form.  The slides should be delivered as soon as possible to the first controller, together with the copy of the form without the results of the microscopy center. The form with the results of the microscopy center should be given to the rechecking coordinator.

RESTAINING THE SLIDES

- Do we need to restain?

- Restaining smears for rechecking may not seem to be logical when first considered, since it might hide possible sources of error such as poor staining, and will add to the workload. Restaining may also remove rare AFB, and make evident problems with contamination with saprophytic AFB (atypical shape, very high % of false negatives). However, experience has shown that poor staining may not be obvious, especially in the worst cases, when AFB will be missed by both the laboratory and the controller.  This is even more likely to occur when fuchsin stain fades.  This usually happens slowly, but when AFB are only weakly stained, they may lose all color and quickly become invisible.  High temperature and humidity will accelerate this process, so that even well stained AFB may fade within a few weeks.

This graph shows the speed of carbol fuchsin fading at 35 degrees C and different relative humidity (RH.

- Also as a result of fading, positives reported by the laboratories cannot be confirmed by the controllers, resulting in false accusations of FP.  Fading can have even greater negative impact on the ability of controllers to detect FN in low positives that are missed by the microscopy centers.

- It is not possible to predict where and when fading will occur to such an extent that it would render the results of rechecking unreliable or even totally wrong. Other factors beside temperature, humidity, time, quality of staining may cause fading, and variables such as the number of weeks are only indicative. We know that fading occurs at temperatures and humidity typical during Asian monsoon season, or in African rainforest and coastal regions. But it has also been seen elsewhere, possibly because of poor staining.

- The following result from Burundi, a country with a mild climate, shows the profound effect that restaining can have on rechecking results. False positives disappear while false negatives become apparent, and therefore required action is totally different

- Restaining policy

Ziehl Neelsen (ZN) stained slides

The recommended approach is to restain all discordant slides at the second controller’s level. However, restaining all slides at the first control level is also recommended when it is observed that the first controllers miss too many scanty and 1+ results, which are found positive at the second control level, or when quantifications are systematically lower for the controllers without restaining. If restaining causes a lot of false negatives due to contamination with environmental mycobacteria, the slides should be kept in closed boxes or absorbent paper could be kept on top of the slides.

To deal with any potential fading, and ensure that errors due to gross deficiencies of staining solutions or staining are not missed, all slides selected for rechecking may be restained at the first controller’s level. Restaining of all slides is justified whenever staining solution quality is not well controlled (i.e. no control on quality of chemicals procured; preparation of solutions decentralized to periphery). In addition, good internal controls on staining solution should be in place to ensure high quality of stains, while proper staining techniques should be applied.

To avoid another restaining at the second control level, which adds the risk of removing rare AFB, restained slides should be kept cool in a closed box until first and second controller results are finalized and no further readings are required.

Based on the conditions in the country and practical experience with rechecking, the head of the national TB reference laboratory, in consultation with NTP management, will decide whether and at what level the slides selected for rechecking will be restained. Restaining of fuchsin stained slides may be performed on:

- all discordant slides at the second control level, or

- all slides selected for rechecking at the first control level.

Restaining Auramine stained slides (Fluorescence Microscopy, FM)

All auramine stained slides should be restained with auramine-phenol stain at the first control level. This recommendation is justified based on country experiences, anecdotal evidence, and theoretical considerations. To make unbiased, blinded rechecking possible, auramine-stained fluorescence positives should be confirmed by changing to higher magnification when needed, but not after restaining with ZN.

Restaining technique

The slides should first be cleaned in xylene. There is no need to decolorize. After cleaning, the slides should be restained with ZN or auramine, using the standard staining procedure.

READING THE SLIDES

Blinded rechecking

First level rechecking should be completely blinded, meaning the first controller should have no access to the results of the centers. The first controller should receive only the slides and a copy of the rechecking form with the identification numbers without the results.

Process of blinded rechecking

First controller:

- Clean all oil-immersion examined slides with xylene.

- Inspect smears with naked eye, record an assessment of size, thickness, evenness, and background color on the form in the column “Remarks”. If slides are not restained, a record of AFB color and artifacts in the background will be added after microscopic examination.  If all slides are routinely restained at the first controller’s level, make a note on the form that restaining was done.

- Restain all auramine stained slides, and all ZN stained slides only if this is NTP policy.

- Read the smears using the appropriate microscopes, bright field for ZN, fluorescence for auramine, and standard magnification.

- Check the routine number of fields, usually 100 fields or one length for a negative smear. Do not check more fields, unless there is a question about the quantification after reading one length.

- Record the first control results on the copy of the rechecking form in the column “First Controls Result”.

- Add an overall appreciation of the quality of smearing (restained smears) or smearing and staining (non-restained smears) at the bottom of the form.

- Do not recheck slides with unclear identification or severely damaged smears; report these as “excluded/ID” or “excluded/damaged”

- Let the oil soak onto absorbent paper (toilet paper, newspaper) as for routine work; replace the slides in the original box in the proper order.

- Keep the box in a refrigerator if the climate is likely to cause fading.

- Send the forms to the rechecking coordinator.

EVALUATING THE SMEAR QUALITY

Smear quality:

The material should have been spread equally over an area of 2 cm by 3 cm or 2 cm by 1 cm. The quality of sputum is good if mucus threads with white blood cells are present; the quality of the smear is poor if the smear is thin with epithelial cells. However, follow–up smears may be thin and not contain white blood cells. Therefore, remarks on poor quality of sputum and smears should only be made if too many thin smears without white blood cells are seen.

Staining quality: A well stained ZN smear shows strong red AFB against weak blue background. A well stained auramine smear shows yellowish fluorescence rods against dark background. Moreover, the background should have very little remaining red in ZN, or fluorescence in auramine stained smears. There should not be any red artifacts in ZN stained smears, and only few fluorescing artifacts in auramine stained smears.

RECHECKING – HIGH AND INTERMEDIATE  LEVEL

__________________________________________________

IDENTIFICATION AND MANAGEMENT OF DISCORDANT SLIDES

______________________________________________________________

| Purpose | To provide participants with an understanding of the procedures for identification and management of discordant slides, including validation of the results and identification and allocation of errors. |
| --- | --- |
| Learning Objectives | At the end of the module, the participant will be able to: - Define discordant slides - Describe the principle of a second control level - Describe the process for managing discordant results - Explain the need for validation of rechecking results - Describe the process for validation of rechecking results - Identify and allocate errors |
| Content outline | - Identification of discordants - Management of discordants - Validation of rechecking results - Identification and allocation of errors |
| Exercises | - Discordants and error identification and validation |

IDENTIFICATION OF DISCORDANTS

To resolve discrepancies, discordant results between the first controller and the laboratory should be counterchecked by a second controller. First controllers can make mistakes, so results have to be validated to ensure the overall reliability of rechecking. Discordant slides are slides called positive (scanty, 1+, 2+ or 3+) by the laboratory and negative by the first controller or the reverse. Quantification errors (QE) are a difference in positive grading of more than one step: scanty versus 2+ or 3+ and 1+ versus 3+.

MANAGEMENT OF DISCORDANTS

Discordant slides (LFN, HFN, LFP, HFP and QE), must be counterchecked by a second controller. This person must be an experienced technician with sufficient time and sense of commitment to yield the most correct result possible. The second controller can be stationed at a higher level laboratory, provided that communication with the first controlling level is easy and fast. Otherwise, it may be more efficient if second controller is stationed at the same level as the first controller, but this obviously will require more discipline to guarantee objective results. Usually this requires a third person, the rechecking coordinator.

The second controller will know both results, but not which result is from the laboratory and which result is from the first controller. Assuming that AFB is present, second level rechecking may require a prolonged search to find the bacilli. Both reported results are used to guide the number of fields to be checked.

- If the discrepancy is negative versus clear positive result (1+, 2+, 3+), two lengths (200 fields) of the slide will be sufficient before concluding that the negative result is correct.

- For scanty versus negative results, 5 lengths of the slide are to be examined before declaring that the negative result is correct.

- For quantification differences examine as many fields as needed to feel confident in making the final judgment.

Positive smears should also be quantified as to the number of AFB seen, e.g. 98/200, 20/500.

Process for managing discordants

Rechecking coordinator

- Copy the results of the first controller to the original rechecking form with the results of the laboratory, and the results from the laboratory onto the copy from the first controller.

- Identify discordant results between the laboratory and the first controller. Using the discordant rechecking form, list all of the discordants with the name of the laboratory.  Use column “Result 1” for some entries of the result of the laboratory and others for entries of the result of the first controller and reversing for column “Result 2” so that the second controller does not know where each result was produced.

- Request the discordant slides from all laboratories on the discordant form from the first controller to be packed in a separate box.

- Send (give) this box, together with the discordant form, to the second controller.

- At this stage, do not give any feedback to the laboratories with discrepant slides, since errors have not yet been verified. Feedback can be given to a laboratory if no discrepant results were identified in the entire sample.

Second controller

- Restain all discordant ZN and auramine stained slides, if not done previously or if not protected against fading since restaining.

- Recheck the smears, allowing for sufficient time to carefully verify results.

- Record the quantified results on the discordant form in the column “Result” for the second controller; add qualitative remarks about smear quality (using same criteria as specified for the first controller).

- Make one copy for the file; send the original, together with the slides, back to the rechecking coordinator.

Rechecking coordinator

- Copy the results of the second controller to the column “Second Controls Result” on both copies of the laboratory rechecking form.

- Using these results as the gold standard, determine a) whether the error was made by the laboratory or the first controller and b) the type of error, according to the table on classification and definition of errors.

- Complete the cross-tables on the bottom of the rechecking form with numbers of slides rechecked and errors identified, both for the microscopy centers and the first controllers.

VALIDATION OF RECHECKING RESULTS

- It is important to validate the results of rechecking prior to sending reports to higher level and to centers. Final analysis and interpretation of rechecking results depends on proper validation, otherwise conclusions may be completely wrong.  Occasionally, controllers may not perform very well. This can be detected in several ways:

- There may not be any discordants identified. This is unexpected and highly unlikely when controller rechecks hundreds of smears originating from centers where positives are regularly encountered. This suggests that controllers somehow knew the result from the peripheral laboratory (results were not blinded).

- The first controller may have more false-negatives than all his/her centers combined together from one same series, i.e. one quarter.  In extreme cases, this means that the controller's examination has been so superficial that it must be assumed that false-negatives of the microscopy centers were not detected.  This would invalidate the results. However, this validation can only made if all smears were restained prior to reading by the first controller, and when the proportion of positive/scanty smears in the total sample was the same as the average SPR in routine.

- Together with few or no false-negatives, superficial controls will result in false declaration of false-positives committed by laboratories. If second controller counterchecks are done well, this will be recognized as FN of the first controller and corrected. If the second controller is also superficial, this may not be corrected and then centers would show many FP, while FN by laboratories as well as first controllers may be rare.

- If FP are detected in many centers, this should always raise suspicion about the quality of the controls. However, this may also be the result of rechecking without restaining. HFP, in particular, should seldom be a problem with trained people and reasonably good microscopes.

- Under some conditions (rapid fading, staining problems), validation can only be done fairly when all smears were restained prior to reading by the first controller. Otherwise, the first controller will have made many false negative errors, simply because the AFB were not visible.

- Validation of the rechecking results may also include feedback from microscopists and supervisors. When smears with serious errors are returned to the center, microscopists will be able to show supervisors the positives missed by both controllers.

- For objective comparison of first controllers and the average of the laboratories, the proportion of positives/scanty smears in the total of the samples must be the same as the average SPR in routine work. If this is not the case, a fair comparison will require a more elaborate analysis by means of the workbook, as discussed in the module: Reporting and analysis of AFB-smear laboratories performance, interpretation of results and feedback”.

IDENTIFICATION AND ALLOCATION OF ERRORS

- After validation of the rechecking results, the errors will be defined according to type, as specified in the table on classification and definition of errors. The result being controlled are the slides with discordance between the laboratory and first controller, and the second controllers’ result will then serve as the final result or gold standard, to decide who made an error and which type of error was made. Classification and definition of errors is specified in the table below.

- Identification and allocation of errors should be done after rechecking of each sample of slides (i.e. quarterly). This will allow for early investigation and corrective action, especially in the case of serious problems. It will also provide the opportunity to monitor whether centers with problems improve their performance with the next round of rechecking.

RECHECKING – HIGH AND INTERMEDIATE  LEVEL

__________________________________________________

REPORTING AND ANALYSIS OF AFB-SMEAR LABORATORIES PERFORMANCE, INTERPRETATION OF RESULTS AND FEEDBACK

________________________________________________________________

| Purpose: | To provide participants with an understanding of the process for reporting and analyzing AFB-laboratories performance, interpreting results and providing feedback. |
| --- | --- |
| Learning Objectives: | At the end of the module, the participant will be able to: - Report results of AFB-smear laboratories performance - Analyze rechecking results - Explain validation of rechecking results - Interpret rechecking results - Describe the purposes and methods of feedback |
| Content Outline: | - Reporting of AFB-smear laboratories performance - Validation of rechecking results - Analysis and interpretation of rechecking results - Examples of analysis and interpretation of rechecking reports - Feedback to the laboratories |
| Exercises: | - Discordants and Error Identification And Validation - Workbook analysis |
| Appendices: | Reporting and analysis documents |

REPORTING OF AFB-SMEAR LABORATORIES PERFORMANCE

Reporting and analysis of rechecking results can be done using a manual system, or together with a refined analysis by means of a workbook which requires computers for recording the data. Regardless of which system is used, data should be compiled and analyzed on a quarterly and annual basis. In addition, reports of laboratory performance should be included that provide essential information on workload and other parameters related to the quality of work. These reports are needed independent of rechecking reports. Analysis of rechecking results using the workbook is only possible with the supplemental data provided in the quarterly/annual report on AFB laboratory performance.

The following documents are used for developing reports on the performance of AFB-smear laboratories:

Quarterly Report “AFB laboratory performance and stock of consumables”: for reporting by the peripheral laboratories (Appendix 1-1).

Quarterly Report “AFB laboratory performance”: for the intermediary level coordinator, to report consolidated performance of the peripheral laboratories upward to higher/national level (Appendix 1-2).

Form 1 “Blinded Rechecking of Sputum Smear Examinations for Acid-Fast Bacilli”: used by the visiting supervisor to register slides sampled for rechecking, and used to provide rechecking feedback to the microscopy center (Appendix 2-1).

Form 2 “Rechecking of Sputum Smears for AFB, List of Discordants”: for use by the rechecking coordinator at the intermediary level to list discordant smears and corresponding results, this form is sent to the second controller to record second controller’s results (Appendix 2-2).

Form 3 “Rechecking of Sputum Smears for AFB, Quarterly / Annual Report Form”: for use by the rechecking coordinator at the intermediary level, to manually compile the rechecking data and forward these upward to higher/national level (Appendix 2-3).

Workbook “Evaluation of AFB-microscopy laboratories” (Appendix 3-1 to 3-11). There are two versions of this book: a smaller one for use at the national level in smaller countries, or at intermediary level in larger countries. It can hold the names of up to 250 laboratories and 31 administrative divisions, one level below the one used for the full workbook (i.e. the smaller country). The larger one is laid out in the same way, but it can be used to hold up to 1500 laboratories and 3 administrative divisions below country level (with max. 50 names for the first level, 100 for the second and 500 for the third). These workbooks are used to compile and analyze data on performance of the labs and rechecking results (from Form 3 and from the “AFB laboratory performance” reports combined).

Use of the laboratory performance documents

The Quarterly Report “AFB laboratory performance and stock of consumables” should be completed quarterly by each peripheral laboratory and sent to the district/region/ province. It can also be completed by a visiting supervisor with help from the laboratory staff. Laboratory staff should, at minimum, count and record the numbers of smears by type of result (positives, scanty, negative) each month, and check and record inventory of the remaining stock quarterly.

The information on smears performed will be used to compile a consolidated performance report for the district/region/province… that can be sent to the higher level. For this purpose the individual reports are either summarized on the quarterly report form “AFB laboratory performance”, either directly or in the Workbook. The following information needs to be recorded:

- name and sequence number of the laboratory,

- technique used: Ziehl-Neelsen or Fluorescence Microscopy,

- totals of AFB smears examined during the quarter/year: from the performance reports of the laboratories, totals of positive, scanty 1-9 and negative smears reported, by type of examination (diagnostic or follow-up),

- stocks remaining for the essential consumables.

The stock supply inventory is not reported to higher up, but used to estimate the needs of essential supplies at the intermediary level, for automatic supply to the laboratories. For this purpose a simple spreadsheet should be used. These stock reports are NOT used to calculate the supplies that need to be ordered from higher up, instead the number of smears examined and the number of functional laboratories will be used for this purpose.

Use of the rechecking documents

The forms

Form 1 is used by the supervisor collecting the sample to list the selected slides and corresponding results on one copy to be kept by the EQA rechecking coordinator; only slide numbers are listed on another copy which goes to the first controller together with the slides.  Upon receipt of the copy with first controller results, the coordinator will fill in the column of missing results on both copies, and identify discordant results.

Form 2 is used by the rechecking coordinator to list slides with discordant results, together with both results (peripheral and first controller).  Discordants from several laboratories can be entered on the same form, as long as the laboratory is also identified on the form. To ensure that the second controller does not know where each result was produced (laboratory or controller), it is imperative that the columns used for the results from the laboratory and the first controller are switched regularly and randomly. For example, enter laboratory results for a few smears using column 1, and column 2 for the first controller results, and listing laboratory results of other smears in column 2 and first controller results in column 1.

Form 2 with the corresponding discordant slides goes to the second controller, who will add her/his results in the third column, and return the form to the coordinator.

Form 1- the coordinator adds the second controller results to both copies of Form 1, and completes the boxes at the bottom of the form, entering:

  - the numbers of slides checked by result, for both laboratory and first controller: these numbers are obtained by counts made in the respective columns (“Peripheral lab result” and “Result first control”); the number will regularly be slightly different between laboratory and first controller, but their totals should be the same.

  - numbers of errors made by both laboratory and first controller in the boxes below the totals rechecked, classified by error type (HFP, HFN, etc.)

Form 1 is then completed by adding recommendations at the bottom, depending on the findings (errors; qualitative remarks on smear and staining quality). One copy of Form 1 is used to give feedback to the laboratory; the second one is filed by the coordinator and will be used to compile the quarterly report (Form 3).

Form 3 is used by the coordinator to compile rechecking results for all the laboratories controlled in a specified area (district, region, province). Only the summary counts are listed, not individual slides. Alternately Form 3 can be omitted if information from Form 1 is entered directly into the workbook (computer). The workbook is then used to report to the higher level. The following information is entered on Form 3:

  - names of the peripheral laboratories and a sequential number (left-hand columns);

  - peripheral labs: total numbers of positive, scanty 1-9 and negative smears rechecked (from Form 1, totals by result for the peripheral lab) and errors found (from Form 1, summary of errors detected for the peripheral lab);

  - names or codes of the first controllers, by laboratory rechecked;

  - first controller(s) summary counts per laboratory rechecked: total numbers of positive, scanty 1-9 and negative results reported for smears rechecked and errors committed by the first controller (from Form 1, as for the peripheral labs);

  - totals for each column are calculated and entered at the bottom.

Form 3 is then sent to the higher level quarterly (an annual report is optional, only required if the compilation/analysis workbook is not used).

The workbook

The workbook is initially for use at the national level, eventually it may also be used at the intermediate level. It will automatically make the additions quarterly and annually, by laboratory, by one or more higher administrative division(s), and by first controller. It will also automatically analyze the results, provided that all information needed is entered, including the performance of each laboratory (as numbers of smears examined). Its use is described in detail in a special module.

VALIDATION OF RECHECKING RESULTS

To check the reliability of the first controller, compare the sum of FN of all laboratories under the first controller with the number of FN by that controller. However, this validation can be misleading in the following situations:

- when positive smears were added to the random samples, resulting in a higher SPR in the samples rechecked as compared with the routine. In this case the controllers are at a disadvantage and must be allowed more FN. (When the workbook is used, the computer analysis will take into account differences in SPR from routine performance data).

- in situations when fading happens rapidly due to climate conditions or poor staining, and only discordants were stained prior to second controls. The first controllers will not have been able to detect some of the true positives in smears that had faded.

ANALYSIS AND INTERPRETATION OF RECHECKING RESULTS

With the use of forms 1-3

For each sample series rechecked, errors will be counted by type and by center. Final analysis can only be completed after the full year’s sample has been processed. However, interim results with general comments should be reported as feedback as early as possible to motivate laboratory staff to continue collection and storage of slides and make any needed improvements in their performance. Sometimes the first series of a sample will show problems that need to be investigated as soon as possible.

It is inappropriate to count all results together, ending with a single score of agreement or discordance, particularly when this includes qualitative aspects such as smear or stain quality. This approach would cause the loss of valuable detailed information and the aim of EQA is not to score laboratories but to detect specific performance problems.

- Although error frequencies can be calculated, they will not be accurate (wide confidence limits) due to the small sample size, except when considering all centers from a large area together. Frequencies for the laboratories can be calculated by type, using as the denominator all slides in the sample with this result as reported by laboratories. For instance, in the case of HFP, all reportedly positive or scanty slides rechecked will be the denominator. Similar for FN (LFN and HFN), all reported negative slides will be used as the denominator. The same calculations will be made for the first controllers, but their denominators may be slightly different. For instance, smears reported positive by the laboratories, but found negative by first controller will be included in the denominator for calculating false-negatives, and not for the false-positives. For validation of rechecking results, comparing FN errors made by the sum of laboratories with those made by their controller should always be done as a check on the controller’s reliability.

- With zero tolerance parameter (d=0) for FN in the sample, not a single false negative should  have been found in the annual sample. FP are never allowable. In principle, finding just one error means that the laboratory cannot be assumed to work at the minimum acceptable quality level. This does not necessarily mean that the quality is bad, only that further checks are needed to exclude poor quality. However, at this point more refinements are possible, and some interpretation is indicated.

In practice it is highly unlikely that there be no errors at all in large samples. If no discordant smears are found in an annual sample of laboratories covered by a first controller and thus the results by the controller are similar to those of the laboratories, an inquiry should be made to determine if the rechecking was really blinded.

- Interpretation of the results should consider the limitations of the system.  Most of the centers with slightly too many errors may not have a real problem. Since the sample has to be small, chance findings are expected. For this reason, priority for corrective action must be given to centers clearly surpassing the limits: more than one HFP or HFN in a small sample, or many LFN or LFP. Otherwise, further routine rechecking will gradually make it clear if there is a real problem in the laboratory. In case of doubt as to whether there is a real problem or just chance errors, it is best not to criticize the staff. Instead, show them the results and encourage them to improve on qualitative deficiencies.  Subsequent series will clarify if the error was a chance finding or not.  Since LFP may be due to a limitation of the control, they can usually be ignored unless they occur at high frequency compared to other centers and controller.  In this case, it is best to check a sample of several recent scanty smears on the spot, to determine if there is a systematic error.

The NTP management, in collaboration with staff of the national reference laboratory, must decide when action should be taken in the case of HFP and FN results, and the level of false results, which action should be taken and by whom. These decisions should consider the local capacity for taking corrective action, including the availability of TB laboratory professionals.

With the use of the workbook

The analysis of the performance calculates total workload, prevalence of positives among diagnostic versus follow-up smears, and the total prevalence for positives or scanty positive smears. These are all indicators which should be checked when probing for problems. Using the workbook, the rechecking results analysis takes into account differences in prevalence of positives between laboratories, and between the samples presented to the first controllers and the routine slides.

The “percentage of true positive/scanty overall detected by the controllers” is a check on the performance of the controllers. This is done by calculating the percentage of true positive/scanty detected by the controllers (= the relative sensitivity of their work), overall and by controller. The formula for this indicator is “((True Positive/(True Positive + False Negative))” x 100%. The value should be (close to) 100%.

The quality of the work of each laboratory is compared to that of the controllers by calculating a ratio: “detection proportional to controllers”, which shows the detection by the laboratory compared to the average for the controller. It is a ratio that converts FN errors by the microscopy center to a relative sensitivity compared to the controller. If this value is 1.00, the laboratory and the controllers detect positives equally; a value below 1.00 means that the laboratory is not as good at detecting true positives, while above 1.00 means the laboratory is better than the controllers. If the value is e.g. 0.65, the center detects only 65% of the positive smears compared to the controller.

The HFP detected are taken into account by showing the “percent true positives out of all positives”, representing the accuracy of a positive result. Scanty results are not included here, since they cannot always be confirmed by rechecking. This indicator is also called positive predictive value (PPV). HFP converted to PPV provides better discrimination between the microscopy centers than the absolute number of HFP. However, it will immediately drop to poor values with very low numbers of positives rechecked, even for a single HFP; therefore the absolute number of HFP should also be taken into consideration. The formula for this indicator is “True Positive/(True Positive + False Positive)”. A value of 100% indicates complete reliability of a positive result. The lower the value, the worse the performance of the laboratory.

The laboratories with inadequate performance can be listed after setting the criteria in the annual analysis sheet, i.e. defining them as those with a detection proportional to the controllers below 0.75 (or less than 75% detection), or those with less than 95% true positives. For quarterly analysis the following thresholds for inadequate performance are recommended: laboratories with (almost) all 1+ to 3+ results false positive, provided that more than just one was rechecked; and centers with a detection ratio below 0.6.

- Although these indicators may not be entirely accurate, they provide a more objective comparison of centers, and allow for setting priorities for intervention, particularly when a fixed sample size is used, since they eliminate the bias caused by differences in prevalence of smear positives in the various centers.

- EXAMPLES OF ANALYSIS AND INTERPRETATION OF RECHECKING REPORTS

1. Annual report of 7 laboratories

Below is an example of an annual report of laboratories under two first controllers in a division. First the report according to the manual system is given, followed by the reports according to the workbook. In the national TB control program the acceptance number had been set at 0 and 60 randomly selected smears are rechecked per center annually.

a. Report according to form 3

FP: 2 centers (EF and FG) have HFP. Center EF has 1 HFP and no LFP, the one HFP may be an administrative error; action can await further rechecking results. Center FG has 4 HFP and 2 LFP, this center appears to have a serious problem, requiring immediate investigation.

FN: Five centers (BC, CD, DE, EF and GH) have HFN; all, except center BC, also with LFN. Priority investigation is required for center DE with 13 FN, and center GH having 7 FN.

Note: serious problems observed in the quarterly reports should be investigated as soon as they are detected, do not wait until the end of the year when the entire sample has been rechecked.

b. Reports according to the workbook

NB: Quarter 4 in above is actually the annual report, covering the 4 quarters of the year.

The workbook provides a more detailed analysis of the rechecking results than the manual system. It takes into account differences in prevalence of positive smears between the laboratories and between the samples and the routine slides. In the annual compilation sheet-evaluation of the laboratories, and in the annual analysis of the performance of the laboratories, unsatisfactory results are marked in color: orange in case of a problem, red in case of a serious problem. In this example the percentage true positives/scanties overall detected by the controller is 100% and the percentage true positives out of all positives is 90%. Of major concern is the overall detection of 0.54 relative to the controller.

The summary sheet for results of rechecking, if used for all laboratories in a large area or the country, is particularly useful for annual reports.

2.  Example of a manual analysis

Project: TTLCP								Year: 1999

| Center | Totals checked |  |  | Number of errors detected |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Pos | Scanty <4/100 | Neg | False positives HFP LFP |  | False negatives HFN LFN |  | QE |
| Madhupur | 66 | 19 | 183 | 0 | 3 | 1 | 1 | 0 |
| Gopalpur | 75 | 7 | 181 | 0 | 1 | 1 | 2 | 0 |
| Ghatail | 72 | 6 | 180 | 0 | 0 | 1 | 3 | 0 |
| Kalihati | 81 | 8 | 180 | 0 | 0 | 0 | 0 | 0 |
| Tangail S | 77 | 4 | 179 | 1 | 0 | 0 | 2 | 0 |
| Delduar | 64 | 3 | 164 | 0 | 0 | 0 | 1 | 0 |
| Mirzapur | 137 | 19 | 357 | 0 | 3 | 3 | 2 | 0 |
| Bhuapur | 75 | 5 | 180 | 0 | 0 | 2 | 6 | 0 |
| Nagarpur | 71 | 3 | 180 | 1 | 0 | 2 | 1 | 0 |
| Sakhipur | 72 | 12 | 179 | 0 | 0 | 1 | 1 | 1 |
| Basail | 70 | 8 | 180 | 0 | 0 | 0 | 1 | 0 |
| Sarishabari | 66 | 13 | 178 | 0 | 3 | 1 | 4 | 1 |
| Bokshigonj | 68 | 7 | 180 | 0 | 3 | 0 | 1 | 0 |
| Dewangonj | 76 | 9 | 180 | 0 | 1 | 2 | 1 | 1 |
| Melandaha | 76 | 7 | 178 | 0 | 0 | 1 | 4 | 0 |
| Madargonj | 65 | 9 | 180 | 0 | 2 | 1 | 1 | 0 |
| Jamalpur | 67 | 7 | 182 | 0 | 3 | 1 | 0 | 0 |
| Islampur | 73 | 14 | 181 | 0 | 0 | 5 | 3 | 0 |
| Totals | 1351 | 160 | 3402 | 2 | 19 | 22 | 34 | 3 |
|  |  |  |  | HFP | LFP | HFN | LFN | QE |
| Errors made by peripheral centers |  |  |  | 2 | 19 | 22 | 34 | 3 |
| Errors rates |  |  |  | 0.15% | 1.26% | 0.65% | 1.00% | 0.20% |
| Errors made y first controller |  |  |  | 1 | 22 | 18 | 41 | 1 |
| Error rates |  |  |  | 0.07% | 1.45% | 0.53% | 1.21% | 0.07% |

The table above provides an example based on a sample size that allows a maximum of 4 FN, with over-sampling of positives. One HFP was found in each of two centers. As this finding is in a high number rechecked, it is unlikely that it is due to a problem that needs immediate investigation. LFP occur in eight centers, but the first controller has the same percentage of these errors, to they can be ignored. Adding the columns for HFN and LFN, five centers have more than 4 FN.  Mirzapur has 5, including 3 HFN, but the sample includes double the number of slides compared to other centers, so it may actually fall within normal limits. Of the other 4, only 2 centers, Bhuapur and Islampur, have considerably more than 4 FN, including more than 1 HFN. Priority should be given to investigating these centers for problems causing FN errors.  It is expected that a serious problem will be found in Islampur, considering the number of HFN in this center. Investigating problems in the other two, Sarishabari and Melandaha, may wait for further rechecking results as there was just one excess FN and most are LFN, it is unlikely that an investigation would discover a serious problem. First controllers have a slightly higher percentage of total FN errors than the total microscopy centers, 1.74% versus 1.65%. However, this is not due to their poorer performance, but to over-sampling of positives, leading to higher SPR and more FN allowed for the controllers. If relative sensitivities were calculated based on above numbers, it would show that the controller’s work was more sensitive. This is clearly one of the advantages of using the workbook.

FEEDBACK TO THE LABORATORIES

Regular feedback to the laboratories is very important for improving or maintaining motivation and improving performance, if applicable. Staff should not be criticized unless the problem is clearly caused by negligence or lack of will to do good work. In general, caution and exclusion of all possible sources of error is needed before jumping to such conclusions. Laboratory staff and controllers should be shown the results of rechecking and be given a copy of the sampling list with results of controllers and summary of errors.  For effective feedback, it is absolutely necessary to return slides with serious errors to the microscopy technician to allow them an opportunity to prove their results to a supervisor.  On-site supervision visits should be used to investigate all possible reasons for errors, and recommend corrective measures such as replacing a bad microscope, or improving smearing and staining techniques. Qualitative remarks on thickness and staining of smears from the controller will be helpful to guide the investigation. Feedback will also serve as continuous education, especially when slides with serious errors are returned.  A visit by a laboratory professional should be planned for centers with too many serious errors. These on-site visits are the best method for improving compliance and maintaining high motivation. Sending reports by mail or other means and requesting staff to improve will have little, if any, effect.

Appendix 1-1

Appendix 1-2

Appendix 2-1

Appendix 2-2

Appendix 2-3

Appendix 3-1

Appendix 3-2

Appendix 3-3

Appendix 3-4

Appendix 3-5

Appendix 3-6

Appendix 3-7

Appendix 3-8

Appendix 3-9

Appendix 3-10

Center = ABF microscopy lab, may be a peripheral or intermediate laboratory