## Slide 1: Group exercise: ON-SITE OBSERVATIONS
Module ON-SITE SUPERVISION OVERVIEW

## Slide 2: Tasks for the group exercise:
review the photos
tell what deficiencies can be noticed
and suggest corrective actions

## Slide 3: Slide #1
Notes: Bars on window – the window can not be properly open for ventilation

## Slide 4: Slide #2
Notes: Inappropriate storage of smears

## Slide 5: Slide #3
Notes: Safety issues: food in the laboratory

## Slide 6: Slide #4
Notes: Safety issues: gloves

## Slide 7: Slide #5
Notes: Improper labeling of specimens

## Slide 8: Slide #6
Notes: Improper application of immersion oil - contamination

## Slide 9: Slide #7
Notes: A technician holds microscope incorrectly

## Slide 10: Slide #8
Notes: Exposition of smears to direct sun light

## Slide 11: Slide #9
Notes: No labeling on bottles with reagents

## Slide 12: Slide #10
Notes: Danger of mislabeling (containers on the photo demonstrate no labeling); non-observation of safety precautions (several sputum containers already opened).

## Slide 13: Slide #11
Laboratory Register
Notes: There is no such result 10/100; there should be 1+ instead
Repeated uniformity of results (1+ and 1+; 2+ and 2+; 5/100 and 5/100) – falsification possible???

## Slide 14: Slide #12
Microscopic observation of smears
#1
#2
Notes: #1 - Under-decolorized
#2 - Dirt with crystals

## Slide 15: Slide #13
#1
#2
#3
#4
Macroscopic observation of smears
Notes: #1 – Too thick
#2 – Too big / uneven
#3 – Under decolorized
#4 – Too thin / uneven