ALGORITHM FOR AFB SMEAR MICROSCOPY FALSE POSITIVE ERRORS INVESTIGATION

INVESTIGATION STEPS

POSSIBLE REASONS

Single HFP

Check the listed results: rechecking form vs. lab. register

If number / result is not the same

If result / identification number is uncertain

Is the lab. register up-to-date?

Are sputum examination forms used?

Is labeling on sputum pots done consistently?

Do results on forms / lists correspond to those in the register?

If yes:

Correct all records

Exclude the rechecking result from analysis

Look for inconsistency in smear results:

Are there more than a few percent isolated true positives and/or scanties?

May point to frequent misidentification of slides

FALSE POSITIVES

More HFP, with or without LFP

Administrative error;

Failure of the controllers

Sloppy administrative procedures;

Confusion regarding recognition of AFB;

Fading of staining but no restaining is used;

Not examining all smears but copying results;

Rarely contaminated carbolfuchsin stain.

Only LFP

Limitation of the controls: poor reproducibility in this range; damaged smears….;

Hazy microscope

Confusion regarding recognition of AFB.

Make the technician show the AFB in the FP slides with his/her microscopes:

Check if the administration is precise and punctual:

Many LFP?- Too many LFP?

If view is not clear

If no AFB but artifacts

If there are clear AFB and true HP

If no AFB shown but technician insists they were there

If yes:

If no:

NOTE: Investigation is not needed for LFP occurring at low frequency, similar to other labs and the first controller. Investigation is also not a priority for a single FP.

Adjust / repair / replace the faulty microscope

Educate the technician

Improve rechecking process

Restain and reexamine HFP:

Do positives re-appear?

Fading is the problem

Probably damaged smear

Check technicians’ quantification skills

Check as many recent LFP as possible on the spot, without restaining

Ask the technician to show what he / she considers to be AFB and correct if needed