## Slide 1: Module PROBLEM-ORIENTED SUPERVISION AND PROBLEM SOLVING
Notes: AG: Again no mention of fm

## Slide 2: Content Overview
Scope of problem-oriented supervision
Problem-oriented supervision as the aim of the EQA program
Problem identification and solving
  Many (or almost all) HFP as well as HFN
  False positives
  False negatives
  Quantification errors
Laboratory indicators

## Slide 3: Scope of Problem-Oriented Supervision
Checking each and every item and procedure is very time consuming:
  the need to observe the routine work for most of a day
Alternative: screening for a few important points based on earlier findings
Opposite to screening the whole process
  only specific areas of the laboratory work are targeted for checks
Effective time management
Most of it can be done by general TB supervisors

## Slide 4: Laboratory Register Checks
The register should be complete and up to date, results look plausible
Check on the last pages if all sputa arrived have been entered
Check if results of the day before were registered
Check with the recent past results if, according to the WHO revised policies:
  most (but not all) suspects had two smears examined
  detected cases had at least one positive result
Notes: STATE that the WHO revised definition of a new sputum-smear positive pulmonary TB case is based on the presence of at least one AFB(+) in at least one sputum sample in countries with a well functioning EQA system. WHO also recommends that in places where a well-functioning EQA system exists, where the workload is very high and human resources are limited the number of specimens to be examined for screening of TB cases can be reduced from three to two.
NOTE that in principle all columns in the laboratory register should be filled, but this will of course depend also on information supplied on request forms, so 100% should never be expected nor demanded.

## Slide 5: Laboratory Register Checks (cont.)
Cross-check the TB lab register against the TB patient register:
  All cases recently detected should be registered for treatment
  Promote writing a note in the lab register:
    the patient register number
    where the detected case was sent for treatment
Cross-check the TB patient register against the TB lab register:
  assess delay between a positive result and start of treatment
Make counts
  Estimate total workload
  Calculate indicators
Notes: EXPLAIN that all cases recently detected should be register for treatment, even if they did not yet (or never) start treatment, an exception could be cases
referred elsewhere

## Slide 6: Looking at Smears
Macroscopic inspection with the naked eye:
  detect poor smearing and identify deficiencies
  assess quality of staining, i.e. decolorization and counterstaining
  check identification / labeling of
    sputum containers
    freshly smeared or stained smears awaiting examination
Notes: EXPLAIN that during the on-site visit the supervisor should macroscopically inspect some recently prepared smears to check the overall quality of smearing, staining and labeling.

## Slide 7: Looking at Smears (cont.)
Microscopic observation of a few recent positives:
  assess quality of staining - if the red colour of AFB is deep and even
  inspect how the microscope functions
Note: smears should be strong positives and definitely recent ones to avoid a wrong impression because of possible fading

## Slide 8: Blinded Rechecking Efficiency and Reliability
The collection of slides kept should be inspected
  for probable completeness
  conditions of storage (out of sunlight, properly arranged by number)
Slides should NOT
  be separated as positives and negatives
  have the result written on them
Notes: NOTE that the suggested evaluation of blinded rechecking practices falls within the responsibility of a higher level national supervisor.

## Slide 9: Blinded Rechecking Efficiency and Reliability (cont.)
Inquire about
  regularity of sampling and how this is done
Check feedback preferably by requesting reports received
  If reports are present:
    check reports for plausibility and results
  In case serious errors were detected in rechecking
    inquire about actions taken and their effectiveness (“evolution of errors”)
Notes: EXPLAIN that the supervisor should check how and at which regularity rechecking and / or panel testing are done. The supervisor should check whether the sample is taken by the supervisor or sent by the laboratory itself and explain that the sample should be taken only by the supervisor, and not by laboratory technicians.
NOTE that absence of EQA feedback reports may indicate probable absence or negligent implementation of EQA blinded rechecking program. Blinded rechecking reports present should be reviewed for the plausibility of results and correct conduct (Were errors found in big series or not found at all? Was the second control done on discordances?).

## Slide 10: Important Issues
Blinded rechecking system and on-site supervision
  are complementary to each other
  will function less effectively without each other
Analysis of laboratory indicators and EQA data will guide supervisors to focus on those elements which can be possible sources of errors
Development and observance of algorithms while investigating of errors will help supervisors to
  find deficiencies
  correct deficiencies
  prevent appearance of errors in future

## Slide 11: Problem Identification and Solving: Investigation of Rechecking Errors

## Slide 12: Investigation of Errors
Consider:
  types and number of errors detected
  the controllers’ qualitative remarks
false positive, false negative or both?
false positives:
  only low, or also high?
  in which numbers?
false negatives:
  low, high, or both?
  in which numbers?
quantification errors?
  when interpreting consider if restaining was used or not

## Slide 13: Many (or almost all) HFP as well as HFN
Indicate a serious problem:
  total lack of training
  or unusable microscope
  or smears were not examined
Investigation:
  examine a 3+ smear with that microscope, to see if it works properly
  request the microscopist to examine a clear-cut positive and a negative smear with a good microscope, to see if he/she knows what is AFB
  if both are OK: conclude at total neglect, not examined

## Slide 14: False Positives
Just one HFP:
  administrative error
    poor slide numbering
    transcription error
  failure of the controllers to detect AFB
Only LFP
  limitation of the controls
    poor reproducibility in this range,
    damaged smears...
  hazy microscope
  confusion regarding recognition of AFB

## Slide 15: False Positives (cont.)
More HFP, with or without LFP
  sloppy administrative procedures
  confusion regarding recognition of AFB
  no restaining in situations with fading of staining, not even of discordants
  not examining all smears but copying positive results to the second and/or third smear of a series
  rarely contaminated carbolfuchsin stain
Notes: EXPLAIN that contaminated carbolfuchsin stain may leave non-fixed AFB on the smear, which can easily be lost again (during cleaning, transport, restaining..) and then are no longer found by the controllers.

## Slide 16: Investigation of False Positives
Not needed for LFP occurring at low frequency, similar to other labs and the first controller
Not urgent for a single high false positive
Check the listed result on the rechecking form with that in the laboratory register:
  same number and result?
    if not: correct all records
    if result/identification is uncertain: exclude the rechecking result from analysis

## Slide 17: Investigation of False Positives (cont.)
Ask the technician to show the AFB in the false positive slide with his / her microscope:
  clearly visible?
    if NO: faulty microscope - adjust or repair or replace microscope
  true AFB?
    if NO AFB but artefacts: educate the technician
    if YES: report controller error to the rechecking coordinator
      rechecking process needs improving?
      recheckers to be replaced?
  if no AFB are seen: restain and re-examine HFP again
    If AFB appeared: fading; need to restain more?
    If NO: damaged smear?
Notes: NOTE that replacement of recheckers is needed only when they repeatedly erroneously inform about false positives – replacement should not be considered if it was an occasional mistake.
NOTE that restaining and reexamination of high false positive smears is done when laboratory technician insists that the initial result was strongly positive.

## Slide 18: Investigation of False Positives (cont.)
Look for signs of sloppy administration / identification:
  is laboratory register up-to-date?
  are sputum examination forms used?
  is labeling on sputum pots done consistently?
  do results on forms/lists correspond to those in the register?
  are isolated positive/scanty results in case series not so rare?
Many LFP:
  check if quantification scale is understood
  if too many LFP:
    check as many recent LFP as possible on the spot, without restaining;
    ask the technician to show what he/she considers to be AFB

## Slide 19: False Negatives
Single LFN
  NO investigation
  Monitor evolution
Single HFN
  an administrative error; or copying results from other smears
  some smears are not examined: days of overload

## Slide 20: False Negatives (cont.)
Several HFN, and/or LFN
  very thick smears
  bad stain or poor staining / poor staining technique
  poorly lighted or hazy microscope (due to dirt, fungus or oil inside, other problems)
  contaminated methylene blue or rinsing water
  superficial or no examination
  very rarely: technician doesn’t know AFB; colour blindness
Notes: EXPLAIN that contaminated with AFB methylene blue or rinsing water is an exceptional situation found in some areas but not everywhere. It can be revealed only if rechecked smears were restained, otherwise these contaminants remain invisible. If contamination is confirmed, the complete sample should be excluded from evaluation

## Slide 21: Investigation of False Negatives
Look at a full box of smears:
    if many too thick, AFB hard to see: poor smearing / less light
    if many too red background: destaining faulty
    if many too dark blue: smearing / counterstaining faulty
Look at carbolfuchsin stain poured on a new slide on the staining bridge:
  can you see the bridge easily? is colour too light?
    If YES: bad carbolfuchsin staining solution, replace

## Slide 22: Investigation of False Negatives (cont.)
Examine positive smears with their microscope
  is light in the microscope less bright?
  If YES:
    check correct position of a condenser and diaphragm
    remove filters (if used)
    if mirror is used: find better place for examination
    check for heavy fungus growth
Notes: EXPLAIN that fungus can grow on and etch the lenses, the eye-piece tube and prisms causing the microscope image to become hazy and unclear. This problem may be correlated with tropical or near-tropical environmental conditions.
EXPLAIN that to check for fungus turn the microscope on:
rotate the 10X objective into grow the light part
Take out both eyepieces and look in their tube holes from a distance: are fungus spots or threads seen?
  If YES: tube prisms must be cleaned (trained person).
  If NO: try replacing the objective and check again
        if still NOT OK: also replace condenser
        if still NOT OK: replace microscope
NOTE that to prevent fungal growth:
The microscope should be kept in a warm cupboard (with tightly fitting door, heated by a light bulb (maximum 40W), located at the top of the cupboard near to the microscope head).
Where power is unreliable, the microscope should be stored in its box (or tight fitting cover) with silica gel, salt or rice (about 100 grams) placed in an open pot on the microscope stage. Ensure silica gel is timely regenerated, rice/salt is renewed. Silica gel changes color from blue to pink when it is unable to absorb any more moisture – dry silica gel by placing in a hot air oven or heating in a saucepan until the blue color appears.
If proper storage not available, keep the microscope in the shade and with good air circulation

## Slide 23: Investigation of False Negatives (cont.)
Examine positive smears with their microscope (cont.)
  hazy view, or complete blur? If YES:
    try to clean eyepieces, condenser, light source and objectives and check again; replace obviously damaged parts
  check red staining of AFB in a few recent non-restained strongly positive smear:
    strong or weak red?
    solidly stained or thin or granular AFB?
Notes: EXPLAIN: A few recent non-restained strongly positive smear should be checked for the staining of AFB: if AFB are strong or weak red? If AFB are solidly stained or thin or granular?
If the staining of AFB is weak red but solid, and staining solution is good:
    suspect poor staining procedure;
    ask to stain few smears and observe if there is sufficient heating and time of carbolfuchsin contact
if also or only too thin or granular AFB: the problem can be the same, but it may happen also be due to poor quality basic fuchsin dye

## Slide 24: Investigation of False Negatives (cont.)
Exclude contamination of methylene blue (only if rechecked smears were restained)
  repeatedly stain known negative smears, check if (atypical) AFB appear
  make ZN smears from taps and containers / glassware used for preparation of solution and staining: AFB found?
    If contamination is confirmed, exclude the laboratory’s sample from evaluation
Check the register for daily workload:
  less than 25 smears on average per technician involved in ZN microscopy?
Notes: COUNTRY CUSTOMIZATION: the problem of contaminated AFB in methylene blue or rinsing water can be revealed only if countries use restaining of smears sampled for rechecking (otherwise contaminants remain invisible).

## Slide 25: Investigation of False Negatives (cont.)
No reason found after all previous checks:
  assume superficial reading to be the cause if false negatives not so rare and more are LFN
  suspect administrative mistake if exceptional (more chance to be HFN)
Notes: NOTE that it is necessary to check administration issues:
  is the laboratory register up-to-date?
  are sputum examination forms used?
  is labeling on sputum pots done consistently?
  do results on forms/lists correspond to those in the register?

## Slide 26: Quantification Errors
May be caused by:
  lack of quantification skills or motivation
  poor microscope
  poor staining solutions or staining procedure
When interpreting, consider:
  whether the rechecking procedure was used with or without restaining?
Notes: STRESS that when quantification error interpretation changes completely depending whether or not rechecking was done with or without restaining.

## Slide 27: Investigation of Quantification Errors
Lack of quantification skills:
  use a panel for confirmation
Smears for rechecking were restained, and quantifications consistently higher for recheckers (and possibly accompanied by false negatives)
  check staining solutions (and staining procedure if solutions are good)
  check the microscope (light, sharpness)
  (check positive smears without restaining to differentiate)

## Slide 28: Investigation of Quantification Errors (cont.)
Smears for rechecking were not restained and quantifications consistently higher for recheckers (and possibly accompanied by false negatives)
  check the microscope (light, sharpness)
No problem detected: lack of efforts

## Slide 29: Laboratory Indicators

## Slide 30: Laboratory Indicators for Sputum Smear Microscopy: Important Issues
Are not well known
May not be easily used:
  chance variations / highly seasonal variation of case-finding
  accurate record keeping is needed
  small counts in small laboratories
Can be used as a part of internal quality control practices
Experts’ assistance is needed to determine the best indicators for the country
Recording and reporting forms
Notes: STRESS the importance of properly developed recording and reporting forms. It will be impossible to calculate the desired indicators if recording and reporting forms are not well designed and/or filled in to collect the needed data.

## Slide 31: Positives Prevalence of Suspects Smears
Indicates lab quality, but also
  accessibility (financial, geographical)
  selection of suspects
Normal prevalence: 10% ?
  varies ~ NTP: accessibility, other diseases (HIV !)
  varies ~ level of service
  seasonal variations ...
Extract lab quality ?
  compare units at same level
  only gross lab deficiencies may show
Notes: <number>
<footer>
EXPLAIN that this indicator is calculated as positive + scanty smears / total suspect smears.
STRESS that this is more often an indicator of accessibility and suspects’ selection than of AFB-smear quality. If clinicians wait too long before requesting a sputum examination, the proportion of positives will be high, and inversely, when they suspect TB for cough of short duration it will be very low. Only comparison of laboratories in the same setting (area, level of service) will allow suspicion about too low positivity rate because of poor quality in some laboratories.

## Slide 32
Monitoring Percent of Positive Suspects Smears (I)
Notes: <number>
<footer>
EXPLAIN that the plotting of the indicator values can be used for internal quality control: the graph presented shows that the values of the prevalence of positives of suspect smears are almost stable during the year.

## Slide 33
Number of TB suspects examined by sputum smear microscopy x 100
Total number of TB suspects identified
Monitoring Percent of Positive Suspects Smears (II)
Notes: <number>
<footer>
NOTE that this is the same example, but with severe fluctuations: a bad staining solution was used early in the year, soon leading to very low detection. It was replaced in July, then cases missed earlier were retrieved (peak in August), thereafter the normal rate installs.

## Slide 34
Sensitive indicator of lab quality
Normal prevalence may be around 10%
  ~ population: early treatment? MDR-TB? HIV?
  ~ guidelines: no. of smears first 2-3 months?
  ~ correct registration: targets!!
Interpretation of very low prevalence
  superficial reading or bad microscope
  and/or poor staining
Positives Prevalence of Follow-up Smears
Notes: <number>
<footer>
EXPLAIN that this is the number of positives/scanty smears among follow-up smears as percent of all follow-up.
STRESS that this is the best indicator for laboratory quality, provided that positive as well as scanty results are faithfully recorded (i.e. sometimes all turned into negative because of wrong emphasis on conversion and outcome targets). Some population difference may also influence: less positives are expected with early detection (and maybe HIV also); more – with high leves of MDR TB.
NOTE that half of these positives tend to be scanty: they will be missed in case of reading problems as well as with bad staining.

## Slide 35: Positives Prevalence of Follow-Up Smears : Example of Data Analysis
Notes: EXPLAIN that this example demonstrates absence (in 16 laboratories) or very low percentage of positive follow up smears in the majority of laboratories. Poor microscopes, ineffective quality control procedures and high workload were named the reasons for that.

## Slide 36
Positives Prevalence of Follow-Up Smears: Example of Data Analysis (cont.)
Notes: EXPLAIN that this chart demonstrates a totally different, maybe optimal situation, with good microscopes and stains, effective QC and EQA practices. All labs detect low-positive follow-up with a mean around 10%.

## Slide 37: How to Use These Indicators?
Internal monitoring, charting
  problem: fluctuations++
    --> only larger units & selected indicators ?
During supervision visits
  can be used by TB supervisors as well
  one value, i.e. full quarter or year
  problem: time constraint in big units
    selected indicators only, i.e. low positives
    scan rather than count
Calculate from lab reports
  ideal: analysis by computer
  problem: few NTPs have lab reports!
Notes: <number>
<footer>
EXPLAIN that indicators can be used in various ways, depending on the situation:
The easiest way is when they are reported and calculated using a computer spreadsheet, also allowing easy comparisons
During supervision visits those that are not reported can be counted and calculated. Easy if counts were already done as part of internalmonitoring. Otherwise time may be insufficient: then flipping through the pages may still give an impression
It is good practice also to compare them with rechecking findings

## Slide 38: Key Messages
Problem-oriented supervision focuses on solutions or strategies for managing particular problems that were identified based on earlier findings
To ensure high quality, besides observations also analysis of laboratory indicators, routine performance reports and data obtained from other EQA activities, especially blinded rechecking, should be used.

## Slide 39: Key Messages (cont.)
It is recommended to develop and follow algorithms when investigating errors, including rechecking errors, during on-site supervision
When developing indicators, NTP managers and NRL staff should properly develop relevant recording and reporting forms for data collection