On-site supervision short checklist exercise - ANSWERS

Laboratory:							 	LABORATORY A

District/Administrative Unit:					District D

Number of Microscopists/Technicians:		One laboratory technician / microscopist

Qualifications of current staff: 				basic training for technicians in 2000

Supervisor/Head of Laboratory: 				Dr. H.L.

Date of Visit: 								November 11, 2006

Visiting Supervisor:						Dr. V.S.

Item				Adequate/		Problems Identified

Acceptable

SOP		Y	N		SOPs are disorganized, and are not kept at the place of use

Separate area for TB work		Y	N		__________________________________________________

Separate tables for specimen		Y	N		__________________________________________________

receipt/smear preparation/microscopy

General order/ cleanliness		Y	N		cleanliness is poor, lab is disorderly and dirty

Power supply		Y	N		power often not available during working hours

Running water supply		Y	N		_______________________________________________

Waste containers with lid		Y	N		waste containers are without lids

Waste disposal by

autoclave/burning/buried		Y	N		waste is thrown out in the corner of the yard

Balance/		Y	N  		not applicable, stains are received from intermediate lab

measuring glassware/ pure water

(if stains are prepared)

Adequate Stock & Supply of:

Specimen containers 		Y	N		__________________________________________________

Slides					Y	N		__________________________________________________

Stains					Y	N		__________________________________________________

Slide boxes				Y	N		__________________________________________________

Microscopes					Y	N		an old monocular microscope with mirror is available

Laboratory register				Y	N		 positive results are not written in red

Laboratory forms				Y	N		__________________________________________________

Overall quality of smears		Y	N		many smears are too thick

Timely reporting of results 		Y	N		__________________________________________________

to clinicians

Personnel					Y	N		no NTP training in AFB smear microscopy

Is QC using positive and negative control slides performed as required by the NTP?		√Yes     No

____________________________________________________________________________________

Are all slides kept as required by the NTP EQA Program?					  Yes    √ No

all slides are kept but some of them have positive result written on them

Are slides properly stored in slide boxes?					  √Yes     No

____________________________________________________________________________________

Workload (last 3 months)

| Number of smears: |  | Number of suspects smears: |  | Number of follow up smears: |  |
| --- | --- | --- | --- | --- | --- |
| Total: 736 |  | Total: 655 |  | Total: 81 |  |
| # Pos 61 | # Neg 675 | # Pos 61 (incl. only few low positives) | # Neg 594 | # Pos 0 | # Neg 81 |

Average number of smears read by each technician per day?                about 10-12 smears per day

Proportion of positive (including scanty) suspects smears among all suspects smears examined?         9.3%

Proportion of positive (including scanty) follow-up smears among all follow-up smears examined?      0%

Rechecking and/or Panel Testing Results (refer to feedback form)

Have any performance problems (based on criteria set by NTP) been identified through rechecking or panel testing?  											  Yes  √  No

No errors were found by rechecking during the whole year.

______________________________________________________________________________________

Has corrective action been adequately implemented?		  Yes     No

_________________________________________________________________________________

Overall Remarks:

Too few low positive suspects smears and no positive follow-up smears, in addition to a lot of thick smears may indicate there would be many false negatives, but this was not found by rechecking.

Action required:

- the problem of unsafe waste disposal should be urgently solved: provide biohazard waste bins with lids / ensure burning or burial of waste

- ensure laboratory cleanliness and proper orderliness of the lab

- ensure proper organization and use of the SOP

- providing a binocular microscope is strongly recommended

- staff should be taught to prepare smears of appropriate thickness

- to ensure that rechecking is performed accurately, no results should be written on smears

- the laboratory technician needs to receive a refresher training in AFB smear microscopy