ALGORITHM FOR AFB SMEAR MICROSCOPY QUANTIFICATION ERRORS INVESTIGATION

INVESTIGATION STEPS

NOTE: Interpretation of quantification errors should be done with knowledge of the rechecking procedure used: was it with or without restaining?

Check whether the technician lacks AFB quantification skills; use a panel for confirmation.

QUANTIFICATION ERRORS

Quantification errors

Lack of quantification skills or motivation;

Poor microscope;

Poor staining solutions or staining procedure

Were smears restained and quantifications consistently higher for recheckers? (And were they possibly accompanied by false negatives?)

Were smears NOT restained and quantifications consistently higher for recheckers? (And were they possibly accompanied by false negatives?)

No reason was identified after all previous checks:

(ALMOST) ALL POSITIVES ARE FALSE; NUMEROUS FALSE NEGATIVES

NONSENSE RESULTS

Total lack of training;

Unusable microscope;

Smears were not examined.

Check the staining procedure (see algorithms for FN and FP errors)

Check the microscope (light, sharpness) - see algorithms for FN and FP errors); check positive smears without restaining to differentiate.

POSSIBLE REASONS

Check staining solutions (see algorithms for FN and FP errors)

If solutions are good:

Check the microscope (light, sharpness) - see algorithms for FN and FP errors)

ALGORITHM FOR AFB SMEAR MICROSCOPY MANY HFP AND HFN ERRORS INVESTIGATION

POSSIBLE REASONS

INVESTIGATION STEPS

Examine a 3+ smear with the same microscope, to see if it works properly

Does the technician understand the quantification scale?

If no:

Assume lack of effort by microscopy technician

Request the microscopist to examine a clear-cut positive and a negative smear with a good microscope, to see if he/she knows how to recognize AFB.

If previous checks revealed no problem:  conclude probable total neglect - smears were not examined

Explain