ALGORITHM FOR AFB SMEAR MICROSCOPY FALSE NEGATIVE ERRORS INVESTIGATION

INVESTIGATION STEPS

POSSIBLE REASONS

Look at full box of smears with naked eye; check some smears microscopically:

Single LFN

Assume superficial reading (if FN not so rare and many of them are LFN)

Suspect administrative mistake if rare (more chance to be HFN)

Look with naked eye at carbolfuchsin stain poured on a new slide on the staining bridge:

FALSE  NEGATIVES

Single HFN

Probably a chance finding; no investigation needed; monitor evolution

Administrative error

Copying results from other smears;

Some smears are not examined;

Days of high workload.

Several HFN and / or LFN

May indicate:

Very thick smears;

Poorly lighted or hazy microscope; other problems with microscopes;

Bad stain or poor staining technique;

Contaminated methylene blue or rinsing water;

Superficial reading;

No examination;

Technician doesn’t  know how to recognize AFB; color blindness (very rarely)

Examine positive smears with technicians’ microscope:

Exclude contamination of methylene blue (only if rechecked smears were restained):

Are there too many thick smears?

Are there too many smears with extensive red background?

Can you see the bridge easily? Is the color light?

If yes:

Is the view hazy or completely blurry?

Check staining of AFB in a few recent non-restained strongly positive smears:

Are there too many dark blue smears?

Is the light in the microscope faint?

If yes:

If yes:

If no:

If yes:

If AFB weak red but solid

If only too thin or granular AFB

If contamination is confirmed:

Check the laboratory register for daily workload:

Are there less than 25 smears on average per technician involved in Z-N m-py?

No reason was identified after all previous checks:

Bad carbolfuchsin solution

Are fungus spots or threads seen?

Replace the objective, try again

Replace

Check correct position of condenser and diaphragm; remove filters (if used)

If mirror is used: move to place with better light

Clean eyepieces, condenser, light source, objectives; replace damaged parts

Take out eyepieces, look in the tube holes from a distance with microscope light on and 10x objective

Tube prisms must be cleaned by a trained person

Take precautions against fungi

If not OK: replace condenser

If still not OK: replace microscope

Suspect poor staining procedure

Suspect poor quality basic fuchsin dye

Repeatedly stain known negative smears

Make Z-N smears from taps and containers / glassware used for preparation of solution and staining

Exclude the complete sample from evaluation