## Slide 1: ON-SITE SUPERVISION – CHECKLISTS

## Slide 2: Checklists
Serve as documentation of the visit and record of current conditions and actions needed
Help carrying out on-site supervision in a consistent and structured way
Need of standard definitions of what is acceptable for each checklist item
  criteria to be established by NTP
Standardized structure:
  Open, non-leading questions
  Results of on-site observations
Notes: STATE that in order to ensure that on-site supervisions are carried out in a consistent and in a structured way, NTP should develop standard check-lists to assist both laboratory and TB supervisors during their field visits. Checklists are the instrument that helps improve efficiency of on-site supervision and provide the means for reporting back to both laboratory staff and NTP.

## Slide 3: Checklists (continued)
Two types:
  Comprehensive (for lab supervisors)
    More effective if shortened version is used later on
  Short (for TB supervisors)
Consider practical use
  Focus on problems that are frequently identified or most likely occur
  Evolution with time
  NTP may refine objective criteria for acceptable practices
Notes: <number>
<footer>
STRESS that checklists is not a static document. Comprehensive checklists may be very useful at initial stages, but as experience is gained there will be recognition of frequent problem areas. The laboratory program should then focus on reviewing those areas where problems are frequently found, and condense the checklist into key questions and observations through country experience. So, the program should think about checklists as an instrument that can undergo evolutionary changes with time, when needed.

## Slide 4: Check Items
Basic facilities / safety
Laboratory equipment
Personnel
Stocks of supplies and consumables
Registration and transmission of results
Laboratory register counts: workload, indicators
Observation of sputum collection, smearing, staining, microscopic examination procedures
Quality control
EQA documentation

## Slide 5: Basic Facilities / Safety
Laboratory space and furniture
Ventilation
Electric power supply
Type of water supply: tap & drain
Sink or basin, buckets, bottles
Waste bucket with a lid; autoclave or pressure cooker; burning facility
Disinfection / overall cleanness
Notes: STATE that supervisors should assess the laboratory general set-up. It should include checking of space, ventilation, presence of running water, electric power supply, safety practices. The supervisor should check if the laboratory uses an NTP approved disinfectant active against TB, if an adequate supply of this disinfectant is available. The supervisor should make sure that the laboratory is clean and has proper waste disposal practices. Biohazard waste bins must have lids. Used containers should be burned or properly buried and never be reused. If slides are reused, they should be properly disinfected and cleaned and never reused for AFB microscopy. If the laboratory uses wire loops for smearing, a sand bucket with ethanol must be available to clean wire loops prior to flaming.

## Slide 6: Laboratory Equipment, Personnel
Types and number of microscopes
Microscopes: objectives, spare bulbs; mirror, evidence of protection
Only if stains are prepared:
  Is there a water distiller or filter? balance? measuring cylinder & glassware available?
Personnel trained for AFB microscopy
Workload
Notes: STATE that the microscope should be binocular, with good 100x immersion objective, good light and mechanical stage. Availability of a fluorescent microscope is justified if workload exceeds 25 smears daily.
NOTE that number of microscopes should be sufficient - too few microscopes may lead to unreliable results, the same time if there are too many microscopes they may spoil or disappear.
STATE that there should be at least 1 spare bulb and be protected against dust as well as against humidity if humidity is over 70% part of the year. The supervisor should check if silicagel against humidity is regenerated, rice/salt is renewed.
STATE that if stains are prepared at the laboratory, the supervisor should check availability of a measuring cylinder, glassware, balance. For stains preparation purified water is an absolute minimum.
STATE that the supervisor should check qualifications, training status of personnel. Laboratory technicians must have at least basic technician training plus working experience or NTP training. The supervisor should also check whether number of staff is sufficient to the existing workload.

## Slide 7: Stocks of Supplies and Consumables
Storage conditions
  tight sealed containers
  out of sunlight
  expiration date
Adequate quantities should be available for 3-6 months:
  check number of examinations over last 12 months; calculate monthly average
Notes: STATE that supervisors should make sure that all staining reagents are available and are within the expiration dates, that stains are protected from sunlight (stored in brown bottles), all bottles and fluids are well closed.
STATE that supplies and consumables quantity in relation to smears processed per month should be roughly estimated. The required quantity of sputum containers and slides is 1100 per 1000 smears examined. The number of slide boxes should be sufficient to keep all slides of a sampling period (quarter?), or 500 slides at a minimum. Adequate quantities should be available for 3-6 months. If properly prepared ZN staining solutions can be stored for 2 years (not more), or during 12 months at least. Auramine stock solution may be kept safely for at least 6 months, working solution for not more than one or two months, and always in the dark. There should be at least 3L of carbolfuchsin or methylene blue and 6L acid solution per 1000 smears. Immersion oil must have acceptable viscosity (neither too thick nor too thin), the required volume of immersion oil is 1 ml per 10-20 smears or 50-100 ml per 1000 smears. The required volume of the burning spirit is 1-2L per 1000 smears

## Slide 8: Registration and Transmission of Results
Check sputum pots and request forms for labeling and completeness of information
Check the laboratory register
  Is it updated daily?
  Is essential information filled in (address of suspects, new / follow-ups)?
  Are results correctly recorded and look plausible?
Cross-check the laboratory against the District Register
  Are patient records in the laboratory register consistent with the district register?
Notes: STATE that specimen containers and request forms must contain legible information needed for patient identification. Request forms must be properly filled in by clinicians.
STATE supervisors should make sure that the laboratory register is present and all columns are completed properly. The register should be updated daily.
NOTE that in principle, all columns in the laboratory register should be filled, but this will of course depend on information supplied on request forms, so 100% should never be expected nor demanded. 100% same results in repeated series (1+;1+; 5AFB/100; 5AFB/100) may look suspicious (falsification possible?). Explain inadmissibility of preparing all smears from one specimen to declare that all specimens were really examined, speak about the problem with clinicians.
STATE that supervisors should check if patient records in laboratory register are consistent with the district register. The supervisor should compare 10 patients from the laboratory register and determine if all 10 patients are listed in the district register. If more than 2% smear positive cases are lost in small, and more than 5% in big centers then think about poor registration and/or transmission results; or poor reception of patients.

## Slide 9: Check on Performance
Calculate workload
  At least 2-3 smears per day but not more than 20-25 smears / per day / per technician
Calculate indicators from lab register / internal monitoring chart
  positive suspects:5-15% ?
  follow-up (FU) positives: 5- 10%?
Notes: EXPLAIN that the positivity rate of suspect smears should be about 5%-15%. If it is abnormal the supervisor should discuss patients’ selection criteria, check if a lot of saliva specimens are submitted to the laboratory. Rarely can it also indicate gross technical errors (stain). The positivity rate of follow up smears should be about 5%-10%. Too low percentage may indicate careless microscopy or stain/staining problem. If the percentage is very high, it may indicate that technician doesn’t know AFB properly, it also can be due to contaminated carbolfuchsin stain.

## Slide 10: Smearing, staining, microscopic examination
Smear preparation:
  Always new slides for AFB smears?
  Cleaning prior to use if greasy?
  Properly labeled?
  Wire loop cleaned in sand and sterilized by flaming OR a new disposable stick used every time?
  Smears completely air dried before fixing?
  Properly heat fixed?
Notes: STATE that supervisors should make sure that technicians verify that containers are properly labeled with patient identification on the side of the container, not on the lid.
The supervisor should make sure that new slides are used for AFB microcopy and they are cleaned prior to use if greasy, labeled with laboratory code, serial number and sequence identifier.
STRESS that improper (confusing) labeling of smears may cause administrative errors and make accurate rechecking difficult.
STATE that the supervisor should make sure that when preparing smears the wire loop is sterilized by flaming after every smear or a new wooden stick is used to prepare each smear.
Prepared smears should be completely air dried prior to fixing and then heat fixed by passing 3-5 times through flame.

## Slide 11: Smearing, Staining, Microscopic Examination (cont.)
Check smears macroscopically
  Thickness, size, color?
Check smears microscopically
  Bright image?
  AFB clearly seen; strong red colour?
How often are stains filtered?
Is the staining done according to NTP guidelines?
  Staining with carbolfuchsin or auramine
  Decolorization
  Counterstaining with methylene blue or permanganate
How many fields are examined to report positive / negative results?
Notes: STATE that supervisors should check smears macroscopically to ensure that their thickness is appropriate (not many too thick or too thin smears); that the size of smears is correct (not many very large or very small smears). The color of ZN smears should be only light blue and should not demonstrate much red color.
STATE that supervisors should examine recent positive smears, using the centre’s microscope to ensure that the microscope is properly functioning, that the image is bright and AFB are clearly seen. ZN stained AFB should have strong red color and completely stained. Smears should have only few or none red dots, crystals.
NOTE that the supervisor should make sure that carbolfuchsin or auramine stain is always filtered before use.
During the staining procedure the supervisor should observe that:
Slides are stained with hot, steaming carbol fuchsin for 10 minutes, more is OK but 5 is the minimum. For auramine, staining is cold during 15 to 30 minutes, not more or less.
Stain is not permitted to dry on the slide.
Slides are decolorized for 3 minutes, repeat decolorization is performed only when needed, slides are not over-decolorized.
Slides are counterstained with methylene blue or permanganate for 1 minute maximum, less is OK
During smear examination the supervisor should observe that microscopist takes at least 5 minutes (and examines 100 fields) to report a negative smear for ZN smears, or 1-2 minutes for auraine smears in FM. The technician should wipe the oil immersion objective at end of every day. The supervisor should make sure that the results are consistent with NTP recommendations for grading and reporting.

## Slide 12: Conduct of Quality Control
Check records and bottles for stain preparation quality control (if stains are prepared)
  Are bottles numbered and dated?
  Are records on controls kept?
  Are positive as well as negative controls tested with each batch?
  Are negatives read after repeat staining?
  Are control positive and control negative smears available?
Notes: EXPLAIN that the supervisor should check records and bottles for stain preparation (if stains are prepared). If no evidence (labels, records) can be provided, probably there are no controls done. Positive as well as negative controls should be tested with each batch; negatives should be repeatedly stained: positive control should give strong red AFB after 1 cycle; negatives should stay without AFB after repeated cycles (it means no contaminants in methylene blue or other stains).
EXPLAIN that the supervisor should also observe availability of sufficient quantity of control slides – absence of unstained control smears may indicate absence of QC practices.

## Slide 13: Storage of Slides and EQA Documentation
Randomly select slides (according to lab register):
  Can >90% of them be found in boxes?
  Is numbering of slides unique / results not written on them?
Examine internal documentation of the rechecking program
  Are records with feedback present?
  Is rechecking regularly done?
  When was it done last time?
  Is sample representative?
  Are results plausible?
  Is interpretation possible?
Notes: STATE that the supervisor should make sure that ALL slides are kept in slide boxes for rechecking – the supervisor was able to retrieve all slides identified from the laboratory register for EQA. Slides should not be separated as positives and negatives. Slides should not have the results written on them. Slides should be cleaned before storage and be stored in boxes so that oil can drain without touching or contaminating other slides.
STATE the supervisor should also examine the internal documentation of EQA programs. Absence of EQA feedback reports may probably indicate absence of EQA blinded rechecking and/or panel testing activities. It is necessary to note that rechecking is not so effective without feedback and regularity. The supervisor may then check whether results are available at the controlling laboratory. The supervisor should examine EQA results for their plausibility. The supervisor should make sure that sample of smears are not taken by technicians themselves, that rechecking was really blind, that the second control was done on discordances. The supervisor should also pay attention to the number and type of errors made by controllers. If the rechecking was not blinded, or if controllers made too many errors results can be disregarded.